Title:

High-Frequency Ultrasound Imaging

AUTHORS:

CORRESPONDING AUTHOR:

KEYWORDS:

SHORT ABSTRACT:

LONG ABSTRACT:

Amelia R. Adelsperger, Evan H. Phillips, and Craig J. Goergen, Weldon School of Biomedical Engineering, Purdue University, West Lafayette, Indiana

High-frequency ultrasound systems are used to acquire high resolution images. A state-of-the-art system, which has high temporal and spatial resolution, will be demonstrated to image the morphology and hemodynamics of small pulsatile arteries and veins found in mice and rats. Ultrasound is a relatively inexpensive, portable, and versatile method for noninvasive assessment of vessels in humans as well as large and small animals. These can be some key advantages as compared to other techniques, such as computed tomography (CT), magnetic resonance imaging (MRI), and near-infrared fluorescence tomography (NIRF). CT involves ionizing radiation and MRI can be prohibitively expensive or impractical in some scenarios. NIRF, on the other hand, is limited by the penetration depth of light required to excite fluorescent contrast agents.

Ultrasound has limitations in terms of imaging depth; however, this may be overcome by sacrificing resolution and using a lower frequency transducer. Abdominal gas and excess body weight can severely diminish image quality. In the first case, the propagation of sound waves is limited, while in the latter case, they are attenuated by overlying tissues, such as fat and connective tissue. As a result, no contrast or faint contrast may be attainable. Finally, ultrasound is a highly user-dependent technique, requiring the sonographer to be familiar with anatomy and to work around issues, such as the appearance of imaging artefacts or acoustic interference.

INTRODUCTION:

Ultrasound is a common clinical imaging modality. The fundamental principles of ultrasound involve the propagation of acoustic waves, their interaction with tissue, and the recording of reflected and scattered waves (i.e. echoes). Recently developed high-frequency transducers can emit acoustic waves between approximately 13 – 70 MHz. For example, a transducer with a frequency range between 22 and 55 MHz has a center frequency of 40 MHz. This range enables a spatial resolution on the order of 50 &#181;m in the direction of the ultrasound beam, making it suitable for imaging structures at the millimeter scale. In order to scan, a transducer first emits a beam of acoustic waves. Some waves are reflected back to the transducer when they collide with a boundary between two tissues, which have different acoustical impedance. The transit time for a wave (i.e. between emission and detection by the transducer) is used to determine individual horizontal lines in an image. Scattering of acoustic waves, that is, the deflection of waves in many directions when they interact with structures much smaller than the wavelength, is responsible for most of the ultrasound image information. Some of this acoustic wave scatter is recorded by the transducer, providing the fine details within an ultrasound image. The use of high-frequency acoustic waves also entails a lower depth penetration because of higher attenuation of sound in tissue. For this reason, high-frequency transducers are only practical for imaging up to a depth of 15-30 mm. In vascular imaging, ultrasound image contrast appears along vessel walls. Red blood cells and platelets also provide speckle contrast within blood. The velocity (v in cm/s) can be measured according to the principle of the Doppler effect:

where c = speed of sound in tissue (154 cm/s); ΔF = Doppler shift frequency (1/s); Fo = transmitted frequency (1/s); and, 𝛳 = angle between the ultrasound beam and the direction of blood flow. Doppler imaging of vessels is a useful means for assessing blood flow dynamics in healthy and diseased states.

PROTOCOL:

1.	Image setup

1.1	Turn the ultrasound system on using switch on the back. Turn the monitor on with the switch on the left side.

1.2	Plug in the physiological monitoring unit and turn on the Heart Rate and Temperature buttons. Turn on the gel warmer and ensure the light is on.

1.3	Check the isoflurane level in the anesthetic vaporizer and refill if necessary.

1.4	Turn on your O2 tank or filtered air source and adjust the air flow on the vaporizer to 1.5 L/min.

1.5	Attach either the mouse or rat stage and then plug the VGA cord into that stage. Then secure the corresponding mouse or rat nosecone into place, and check that the isoflurane (black) and waste gas (blue) tubes are properly connected to nosecone.

1.6	Choose a transducer to use for your procedure and plug it into the “active” port below the imaging system. Run the transducer cable through plastic holders above the probe mount and secure the transducer into the clamp. Note there is a small, raised line on one side of each transducer to determine image orientation.

1.7	Proceed with knocking down and preparing animal for imaging. Add ophthalmic ointment to the eyes, secure the paws to the stage electrodes, and remove the hair in the area of interest using a depilatory cream. Cover the area to be imaged with warmed ultrasonic transducing gel.

2.	Acquiring images

2.1	On the system, start a new study by selecting New and Study, or find a study you previously started and select New and Series. Once in a new series, select a user off of the menu and name your series appropriately.

2.2	Once your series has been created, select B-mode (brightness mode) from the keyboard. All imaging modality keys are on the bottom row of the black keyboard.

2.3	Roll the transducer down to the desired location making sure to look at the screen to avoid squishing the animal too hard. Check the respiratory rate (RR) that appears on the screen as well, because squishing the animal too hard may cause the RR to drop.

2.4	While looking at the B-mode image, adjust the placement of the transducer by gently turning the x- and y-axis knobs found on the stage until you find the desired location.

2.5	Once you have the desired location, wait for the white bar at the bottom of the image to fill before pressing Image Label to save the image. When labelling your image, the type of modality will be indicated next to the image label on the study management screen, so you do not need to worry about including this in the image label.

2.6	For M-mode images, select M-mode (motion mode) from the keyboard. Use SV gate to narrow or widen the yellow bars and the cursor to align the bars over the desired location. Once placed correctly, push M-mode again. The placement of the bars may be adjusted while in m-mode.

2.7	For EKV imaging, select B-mode and ensure you are in the proper location. Then select EKV, adjust the settings, and push select image. The EKV modality averages b-mode image over several cardiac cycles.

2.8	To use color doppler, select B-mode, check that you are in the proper location, and then select Color. Select Update and move the cursor up, down, left, or right to achieve the desired box size, and select Update to lock it. You may then use the cursor to move the box to the desired location. Turning the Velocity knob up increases the velocity threshold and may decrease background signal.

2.9	To use pulsed wave doppler for blood flow velocity, it is helpful to first be in color Doppler mode before pressing PW. Two yellow, angled lines will appear, and the user must align the shorter, dotted line to be parallel to the anterior and posterior vessel walls by adjusting Beam Angle and turning the PW Angle knob. The dotted, yellow line will turn blue if the angle is turned too far. Once this is aligned, press PW, and then adjust the Baseline, Velocity, and Doppler Gain controls to center and brighten the wave forms.

2.10	To use 3D mode, begin in B-mode and align the probe in the center of the structure you wish to image. Use respiratory gating and EKG-trigger for cardiovascular applications. Press 3D and set the desired scan distance and step size. Once the scan is completed, click Load into 3D to visualize your 3D data.

2.11	You may view your acquired images at any time during imaging by pressing Study Management from the top right of the two columns of buttons on the left side.

2.12	When you are finished acquiring images in a series, select Close Series from the study management screen. You may now open a new series within your study, if needed.

3.	Data export and clean-up

3.1	To export data for analysis, select the studies or individual series you wish to export from the study management screen.

3.2	Click Copy to from the top right corner on the study management screen. Select your desired file location, and press okay.

3.3	You may now remove and recover animal.

3.4	To clean the ultrasound set-up, spray a paper towel with T-spray to wipe off the heated stage and rectal probe. Never spray the disinfectant directly on the stage.

3.5	The transducer should be wiped off with 70% ethanol on a paper towel before being placed back in the holder.

3.6	To turn off the air, shut off the O2 tank or filtered air source. You should see the air flow bead slowly drop to 0 on the vaporizer.

3.7	Once you are done with the system, click the power button on the study management screen in the top right corner and allow the monitor to turn off completely.

3.8	Turn the power button on the back of system off only after the monitor is completely off. You should hear the fan stop once it has been properly shut down.

Representative Results:

This procedure allowed the anatomical and functional imaging of the abdominal aorta. Acquiring real-time images in short-axis and long-axis by B-mode, M-mode, and Doppler ultrasound takes at least thirty minutes and therefore requires careful monitoring of the anesthetized animal. Some data are readily analyzed on-the-fly, such as two-dimensional B-mode scans (Fig. 1). These data can provide aortic diameter or cross-sectional area measurements. Other data, such as three-dimensional B-mode (Fig. 2), M-mode (Fig. 3), Color Doppler (Fig. 4), and PW Doppler (Fig. 5), are usually analyzed offline in order to determine aortic volume, circumferential cyclic strain, and blood flow velocity. Together these datasets provide quantitative and qualitative information on the three-dimensional morphology as well as the hemodynamics and pulsatility of the abdominal aorta.

Discussion:

Recently developed high-frequency ultrasound transducers are well suited for visualizing small structures to a depth of up to 3 cm. Here the versatility of a small animal ultrasound system was demonstrated to acquire in vivo imaging data of the dynamics of the mouse aorta. This technique requires practice and recognition of common difficulties, such as abdominal shadows and Doppler scan alignment. Despite these limitations, it is a powerful and versatile technique for quickly obtaining non-invasive imaging data. Importantly, this technique lends itself well to serial imaging of the same animal for longitudinal studies of disease progression or treatment.

Small animal high frequency ultrasound can be used in a variety of cardiovascular applications. Vascular applications include screening for aortic disease (such as aortic aneurysms and dissections), detecting atherosclerotic plaque, and measuring blood flow in patients with peripheral arterial disease. Carotid arteries, iliac arteries, and the inferior vena cava, can be easily imaged with ultrasound. Cardiac imaging is also a major application of this technique and is used to be able to visualize the atria and ventricles of mice or rat hearts. Cardiac ultrasound imaging can give the user a lot of information about the heart including anatomical dimensions, contractility, stiffness, cardiac output, flow patterns, valve function, and/or thrombus formation, to name a few. Ultrasound may also be used for reproductive system imaging (such as the uterus and cervix) or on the bladder. Reproductive system imaging would be useful to look at the structures of and obtain dimensions for the uterus, cervix, and/or vagina. Pups could also be visualized and measured in a pregnant mouse or rat. Due to the advances in transducer technology and innovations in ultrasound technology, these applications work well in small animals and may also have applicability to superficial human imaging.

Disclosures:

ACKNOWLEDGEMENTS:

REFERENCES: