Title: Hot-wire anemometry

Overview:

# Source: Ricardo Mejia-Alvarez and Hussam Hikmat Jabbar, Department of Mechanical Engineering, Michigan State University, East Lansing, MI

Hot-wire anemometers have a very short time-response, which makes them ideal to measure rapidly fluctuating phenomena such as turbulent flows. The purpose of this experiment is to demonstrate the use of hot-wire anemometry.

Principles:

The hot-wire anemometer

A hot-wire anemometer is a device used to measure flow velocity based on the heat dissipated from a very thin electrically-heated wire. The heat generated by an electric wire, q, is given by the relationship:

| q= RwI2 |  |
| --- | --- |

Where Rw denotes the wire’s electric resistance and I the electric current flowing through the wire. The electric resistance depends on the wire’s temperature according to the following relationship:

| Rw=Ro1+αTw-Ro |  |
| --- | --- |

Where Ro is the wire resistance at the reference temperature To and α is a constant that depends on the material of the wire. While equation (1) expresses the heat generated by the electric current, the heat dissipated by the flow, follows King’s law [2]:

| q=a+b∙un |  |
| --- | --- |

Here, a, b, and n are calibration constants and u is the flow velocity around the wire. The value of n depends on the Reynolds number, and it has already been found that n=0.45 is satisfactory for the range of Reynolds numbers achievable in this particular experiment. To obtain a relationship between temperature, electrical current, and velocity, we combine equations (1) and (3):

| RwI2=a+b∙u0.45 |  |
| --- | --- |

Here, the temperature dependence enters through the electrical resistance (equation (2)). The measurement strategy that we will use in the current experiment is to maintain the temperature (and hence the resistance) of the wire constant. From equation (4), it is clear that if the electrical resistance is constant, the current needs to fluctuate to follow the tendency of the velocity. In other words, cooling rate changes with flow velocity, and that would change the wire’s temperature unless the current is changed to compensate. Obviously, it is necessary to have a quick responding electric system to measure a rapidly varying velocity signal. This is achieved with a Wheatstone bridge like the one shown in Figure 1(A). From the figure, the hot-wire is one of the four resistors in the circuit. Figure 1(B) shows its physical configuration, which is a very thin wire set between two prongs (a 5 m tungsten wire for the current experiment). The control resistor, R1, from Figure 1(A) is adjusted initially to produce a zero bridge-voltage, EB, for the desired baseline temperature (ergo electrical resistance) of the hot-wire. When in operation, the bridge-voltage is used as a feedback signal to increase or decrease the current to the wire in order to maintain the hot-wire at a constant temperature. On the other hand, EB, is amplified to achieve an easier-to-read voltage scale, E. This voltage is related to the current through Ohm’s law:

| I=ERw |  |
| --- | --- |

Hence, equation (4) can be expressed in terms of the voltage as:

| E2=A+B∙u0.45 |  |
| --- | --- |

With the calibrations constants now defined as: A=aRw and B=bRw. The main purpose of this experiment is to find the value of these calibration constants. To this end, the hot-wire probe will be set in a reference flow system. This flow system will be used to issue several flows with known velocities. Then, the calibration constants will be found using a least-squares regression.

As shown in the schematic of Figure 2, the reference flow to use herein is the vena contracta of a free jet. The mean velocity at the vena contracta is well characterized by the following equation [3, 4, 5]:

| Vvc=0.612Ppl-Patmρ |  |
| --- | --- |

Here, the constant 0.61 is the discharge coefficient of the jet, Ppl is the pressure inside the plenum, and Patm is the atmospheric pressure. The position of the vena contracta is well defined by the relationship:

| xVCW=1.5 |  |
| --- | --- |

Where xVC is the distance from the jet exit along its centerline and W is the width of the slit from where the jet is issued. This is the location where the hot-wire anemometer will be located for its calibration. Figures 3 and 4 show the flow system used herein. In this system, a fan pressurizes a plenum that has two exits, a slit to produce the jet, and a stack to divert the flow. As the flow through the stack is restricted with orifice plates (see Figure 4 for reference), the flow rate of the jet increases. This setup will help us producing a scatter plot of Vvc0.45 and the voltage E measured at the Wheatstone bridge.

Figure 1. Schematic of the planar jet showing: the vena contracta and the diagram of connections.

Figure 2. Experimental setup. (A): flow facility; the plenum is pressurized by means of a centrifugal fan. (B): slit for issuing the planar jet. (C): traversing system to change the position of the anemometer along the jet.

Procedure:

- Measure the width of the slit, W, and record this value in table 1.

- Verify that the data acquisition system follows the schematic in Figure 2.

- Connect the positive port of the pressure transducer (see Figure 2 for reference) to the plenum pressure tap (Ppl).

- Leave the negative port of the pressure transducer open to the atmosphere. Hence, the reading of this transducer will be directly Ppl-Patm; as required by equation (7).

- Start the program for hot-wire calibration. Set the sample rate at 100 Hz for a total of 1000 samples (i.e. 10s of data).

- Make sure that channel 0 in the data acquisition system corresponds to the voltage of the hot-wire anemometer.

- In the field corresponding to channel 0, select the value of the constant n to 0.45.

- Set the hot-wire anemometer at the position of the vena contracta (on the centerline, at x = 1.5 W).

- Make sure that channel 1 in the data acquisition system corresponds to the signal of the pressure transducer.

- Enter the values of local air density (typically 1.2 kg/m3 for average local conditions) and conversion constant from volts to pressure (76.75 Pa/V) in the fields corresponding to the pressure transducer. Record these values in table 1. With this, the data acquisition system will report the data directly in velocity in m/s according to equation (7).

- Cover the stack completely to establish the condition for maximum velocity at the jet.

- Turn on the flow facility.

- Acquire a dataset.

- Change the stack plate for one with a lower restriction (larger diameter)

- Acquire a dataset.

- Repeat steps 1.15 and 1.16 for a total of at least four times. Make sure that the last repetition is carried out with the stack fully unrestricted (lowest jet velocity).

- The data acquisition program will carry out the least-squares calculation and report the calibration constants automatically. Record these values on table 1.

Table 1. Basic parameters for experimental study.

| Parameter | Value |
| --- | --- |
| Slit width (W) | 19.05 (mm) |
| Air density (r) | 1.2 (kg/m3) |
| Transducer calibration constant (m_p) | 76.75 (Pa/V) |
| Calibration constant A | 5.40369 (V2) |
| Calibration constant B | 2.30234 (V2 s0.65m-0.65) |

Figure 3. Hot-wire anemometer circuitry. (A): Wheatstone bridge circuit to ensure constant temperature in the hot-wire. (B): detail of the structure of a hot-wire.

Figure 4. Flow control in the flow system. The stack on top of the plenum serves the purpose of diverting flow from the jet slit allowing to control the jet’s exit velocity.

Representative Results:

The experimental measurements are listed in Table 2 and shown in Figure 5. A linear regression of these data produced the following result for equation (6):

| E2=5.40369+2.30234∙u0.45 |  |
| --- | --- |

Which can be used to determine the velocity as a function of the voltage:

| u=E2-5.403692.3023410.45 |  |
| --- | --- |

Table 2. Representative results. Measurements of voltage square and velocity at the vena contracta to the 0.45 power.

| V_VC^0.45 | E^2 |
| --- | --- |
| 3.119 | 12.584 |
| 3.919 | 14.425 |
| 4.143 | 14.946 |
| 4.278 | 15.256 |
| 4.465 | 15.679 |

Figure 5. Hot-wore anemometer’s calibration curve.

Applications and Summary

Given that turbulence exhibits high frequency velocity fluctuations, hot-wire anemometers are suitable instruments for its characterization due to their high time-resolution. In the present experiment, we demonstrated the process of calibrating a hot-wire anemometer. To this end, we compared the voltage signal of the anemometer with known values of velocity at the vena contracta of a well-characterized jet. These measurements were used to determine the calibration constants for the linear response of the anemometer.

Hot-wire anemometry is extensively used in scientific studies of boundary layer flows in wind tunnels. Boundary layer is one of the oldest subjects of research in fluid mechanics due to its relevance to technological applications such as aerodynamic design, naval engineering, power generation, among others. To the detriment of all these fields, many effects involving the boundary layer are still incipiently understood: highly irregular roughness, density and viscosity gradients, and compressibility, to mention a few. With this in mind, hot-wire anemometry is used in laboratory settings to assess boundary layer flows relevant to the above-mentioned applications, using strategies similar to the one demonstrated in the current experiment.

References

- Chapra, S.C. and R.P. Canale. Numerical methods for engineers. Vol. 2. New York: McGraw-Hill, 1998.

- King, L.V. On the convection of heat from small cylinders in a stream of fluid: determination of the convection constants of small platinum wires with applications to hot-wire anemometry. Philosophical Transactions of the Royal Society of London. Series A, Containing Papers of a Mathematical or Physical Character 214 (1914): 373-432.

- White, F. M. Fluid Mechanics, 7th ed., McGraw-Hill, 2009.

- Munson, B.R., D.F. Young, T.H. Okiishi. Fundamentals of Fluid Mechanics. 5th ed., Wiley, 2006.

- Buckingham, E. Note on contraction coefficients of jets of gas. Journal of Research, 6:765-775, 1931.