TITLE 
Measuring Blood Pressure in Mice using Volume Pressure Recording, a tail-cuff method
Authors: 
Alan Daugherty
  Cardiovascular Research Center
  University of Kentucky
  Alan.Daugherty@uky.edu

Debra Rateri
  Cardiovascular Research Center
  University of Kentucky
  Debra.Rateri@uky.edu

Lu Hong
  Cardiovascular Research Center
  University of Kentucky
  hlu4@email.uky.edu
Anju Balakrishnan
  Cardiovascular Research Center
  University of Kentucky
  abala3@uky.edu
Corresponding author:  Alan Daugherty

Keywords: 
blood pressure, systolic, diastolic, tail-cuff, mouse, rat
Short Abstract: 
The CODA 8-Channel High Throughput Non-Invasive Blood Pressure system measures the blood pressure in up to 8 mice or rats simultaneously. This tail-cuff system uses Volume Pressure Recording (VPR) to measure the blood pressure by determining the tail blood volume.

Long Abstract:
The CODA 8-Channel High Throughput Non-Invasive Blood Pressure system measures the blood pressure in up to 8 mice or rats simultaneously. The CODA tail-cuff system uses Volume Pressure Recording (VPR) to measure the blood pressure by determining the tail blood volume. A specially designed differential pressure transducer and an occlusion tail-cuff measure the total blood volume in the tail without the need to obtain the individual pulse signal. Special attention is afforded to the length of the occlusion cuff in order to derive the most accurate blood pressure readings. VPR can easily obtain readings on dark-skinned rodents, such as C57BL6 mice and is MRI compatible. The CODA system provides you with measurements of six (6) different blood pressure parameters; systolic and diastolic blood pressure, heart rate, mean blood pressure, tail blood flow, and tail blood volume. Measurements can be made on either awake or anesthetized mice or rats. The CODA system includes a controller, laptop computer, software, cuffs, animal holders, infrared warming pads, and an infrared thermometer. There are seven different holder sizes for mice as small as 8 grams to rats as large as 900 grams. 
Text: 
1.0) Equipment Set-up
1.1)  Select an area in the lab where the room temperature is at or above 20(C. Avoid locations near air conditioning vents and/or loud noise. Avoid exposing the animal to odors that may cause irritation and/or stress, such as perfume, cologne, chemicals, and cleaners.

1.2) Turn on the CODA Controller.

1.3) Turn on the Warming Platforms and set to Level 3. If room temperature is above 25(C, start on Levels 1 or 2.

1.4) Open the CODA Software and select the CODA Device by clicking on it.

1.5) Click on “Test Selected Device”. Select the cuffs and channels to test and click “Test”.

1.6) After the test is complete, close the Device Test window.

2.0) Software Set-up

2.1) Select Tools > Manage Personnel and click on “Researcher 1” to change the name. Click on subsequent rows to add additional researchers.

2.2) Click on “Technicians Tab” to modify or add technicians.

2.3) Click on “Specimens Tab” to name the animals for easy identification and to select the animal type; rat or mouse.

2.4) Click on “Animals Tab” to alter the animal type and sensitivity.

2.5) Select File > New > Experiment and enter the Experiment Name.

2.6) Select the Key Researcher, a Begin Date and click “Next.”

2.7) Enter the Session Name.

2.8) Set “Acclimation Cycles” to 5. If you prefer not to include acclimation cycles, set to zero. To assist with acclimating the animals to the holders and cuffs, add additional cycles to your experiment and run the experiment for 5 consecutive days.

2.9) Set “Number of Sets” to 1.

2.10) Set “Time Between Sets” to 30 seconds for multiple sets

2.11) Set “Cycles per Set” to 20 cycles.

2.12) Set “Time Between Cycles” to 5 seconds and click “Next”.

2.13) Select the specimen from the specimen pool and assign it to the first available channel and click “Next”. 

2.14) Set “Maximum Occlusion Pressure” to 250mmHg.

2.15)  Set “Deflation Time” to 20 seconds for mice, 15 seconds for rats.


2.16) Set “Minimum Volume” to 15ul.

2.17)  Set “Display Style” to One Channel per Graph.

2.18) Add additional Researchers and Technicians and click “Next”.

2.19) Review the Session Script and click “Next”. Do not select Finish until you complete the Animal Preparation section.

3.0) Animal Preparation

3.1) To train and improve animal acclimation, add additional cycles to your experiment and run the experiment for five consecutive days. You also can assist with acclimation by placing the animal in the holder for 15 minutes for 3 consecutive days prior to the actual study.

3.2) Place each animal in a holder by picking up the animal by the tail and gently placing the animal into the rear of the holder which faces the open end of the nose cone. Note: the size of animal in relation to the size of the selected holder. The animal should be able to freely enter the holder.

3.3) Carefully secure the rear hatch to the holder by turning the red screw cap on the rear hatch. Care should be taken to avoid pinching the tail or any other body parts while securing the rear hatch.

3.4) Slide the nose cone toward the rear hatch, limiting the movement of the animal. The nose cone should be in a position to limit the animal from turning around while inside the holder. 

3.5) Place the holders onto the warming platforms in the designated position. If testing the same group of animals over a period of time, be sure to rotate the holder position on the warming platform to improve consistency and accuracy of your BP measurements. Allow the animals at least 5-minutes to acclimate to the holder. Do not touch or handle the animals while inside the holder. The increased contact could irritate the animals. Never leave the animals in the holders unattended.

3.6) Thread the tail through the “Occlusion Cuff.” Place this cuff as close to the base of the tail as possible without force. Secure the “Occlusion Cuff Tubing” in the notch on the top rear of the holder.

3.7) Thread the tail through the “VPR Cuff”, placing it within 2mm of the “Occlusion Cuff”. Secure the “VPR Cuff Tubing” in the notch on the top rear of the holder.

3.8) Attach the cuffs to the corresponding ports on the CODA Controller.

3.9) Allow the animals at least 5-minutes to thermoregulate. Record the temperature of the animal frequently with an infrared thermometer. The animal temperature should be 32 to 35(C.

4.0) Run Experiment

4.1) Click “Finish” and the experiment will begin immediately.

4.2) Return the animals to their cages after the “Simple Sessions Summary Report” is displayed.

4.3) After completion of the last cycle the accepted cycles will be automatically displayed in spreadsheet format within the CODA application, click on the “Export to Excel” button.

4.4) Select the folder to place the file in, name the file and click “Save”. Open the Excel application and click on “Open”. View all file types to display the “.csv file saved”, click on the file.

Representative Results: –
There are several ways to process the data within the Excel file. A common practice is to obtain the average and standard deviation. Delete measurements if the standard deviation is greater than 30. Additionally, rejected cycles may be viewed and the entire database can be easily exported to Excel. Consult the User’s Manual for additional information. When you are done processing the data you will have accurate and consistent blood pressure measurements. Volume Pressure Recording is a clinically validated tail-cuff method which is used for a wide variety of clinical research, including phenotyping and cardiovascular 
research 1,2.  
[image: image1.jpg]@ Simple Session Summary Report
) (o Torwn |

Description

Channel

Reqular Cycle

|Aooepted

Diastolic

Flow

0.81

770

2270

20.90
1997

18.17

15.72

19.90

257
2819

1144

2943

2456

JHHHEHBUEHOUHE

HEEEERERRER@DEOOOOO

KEEREEREORNERREREODODO

19.82

=y
o
w
W W W W W W W W W W W W W W W Ww W W W W w W w w w w ww w w
&

Mouse 3] O
Mouse v ] 3251
Mouse ] ™ 2359
Mouse | %] 30.88
Mouse ™ V] 2141
Mouse [v] (] 17.41
Mouse “ ]
Mouse [v] [v]
Mouse ] ]
Mouse [v] []
Mouse [v] [v] 3173
Mouse v ] 29.12
Mouse [ ] 2820
Mouse ] ™ 3195
Mouse | ] 2968




Figure 1, Example of a Simple Session Summary – this image will be provided once the experiment is complete.
Discussion: 
n/a
Acknowledgments: 
n/a

Disclosures: 
I have nothing to disclose

Table of specific reagents and equipment:

	Name of the product
	Company
	Catalogue number

	CODA
	Kent Scientific Corporation
	CODA


References: 
1.   Feng, M., Whitesall, S., Zhang, Y., Beibel, M., D’Alecy, L., DiPetrillo, K. American Journal of Hypertension. 2008 Dec; 21(12):1288-91.
2. Daugherty, A., Manning, M.W., Cassis, L.A. The Journal of Clinical Investigation. 2000 Jun; 105(11): 1605-1612.

Tables
n/a

Figure legends
Figure 1, Example of a Simple Session Summary
