JoVE Article Template:
OPERATION OF AN OXYGENATION INSERT FOR ADHERENT CELLULAR CULTURES 

Authors: Shawn C. Oppegard, Marie-Elena Brett, Elly Sinkala, Keith Gooch, and David T. Eddington
Author Contact Info: 
Shawn C. Oppegard
Department of Bioengineering
University of Illinois at Chicago
soppeg2@uic.edu
Marie-Elena Brett

Department of Bioengineering

University of Illinois at Chicago

mebrett76@gmail.com
Elly Sinkala

Department of Bioengineering

University of Illinois at Chicago

esinkala@gmail.com
Keith Gooch

Department of Biomedical Engineering

The Ohio  State University

gooch.20@osu.edu

David T. Eddington

Department of Bioengineering

University of Illinois at Chicago

dte@uic.edu

Corresponding author: David T. Eddington
Keywords: hypoxia, cell,culture, control, wound, healing 
Short Abstract: should include a general description of the article (10 word minimum, 50 word maximum)

A device has been designed and validated to provide precise control over the temporal and spatial oxygenation exposed to cells and tissues in a multiwell plate.  In this article, we will demonstrate utilizing the device to study wound healing and ex vivo tissue slices.
Long Abstract: (150 word minimum, 400 word maximum)

Oxygen is a key modulator of many cellular pathways, but current devices permitting in vitro oxygen modulation fail to meet the needs of biomedical research.  A microfabricated insert for multiwell plates has been developed to more effectively control the temporal and spatial oxygen concentration to better model physiological phenomena found in vivo.  The platform consists of a polydimethylsiloxane insert that nests into a standard multiwell plate and serves as a passive microfluidic gas network with a gas-permeable membrane aimed to modulate oxygen delivery to adherent cells.  Equilibration time is on the order of minutes and a wide variety of oxygen profiles can be attained based on the device design, such as the cyclic profile achieved in this study, and even oxygen gradients to mimic those found in vivo. The device’s applicability to studying the in vitro wound healing response and the effect of intermittent hypoxia on ex vivo vein function will be demonstrated.
1. Device function
1. The hypoxic insert device contains 6-pillars that nest into a standard 6-well plate.  Gas flows into the pillar, across a microfluidic network at the base of the pillar, and flows back out of the device.  At the base of the pillar forming the bottom wall of the microchannel is a 100 µm thick gas-permeable PDMS membrane that permits oxygen diffusion between the gas microchannel and the culture media.  Thus, the device operates by establishing a concentration gradient that drives the media oxygen concentration towards a desired value.
1. The device offers a number of advantages over the hypoxic chamber and other cell oxygenation systems: 1) minimizes the oxygen diffusion path permitting faster temporal control, 2) enhances the spatial control over oxygenation, as dependent on the pillar microchannel design, 3) possesses a smaller lab footprint and higher throughput to increase experimental efficiency, and 4) adapts to common cell culture tools (ie multiwell plate) without the requirement for specialized knowledge and equipment.
2. Device fabrication
2. First the pillar array is replica molded with PDMS in a previously machined Delran mold.
2. Next the gas microchannel at the bottom of each pillar is fabricated using standard SU-8 photolithography and PDMS replica molding.
2. The gas-permeable membrane is fabricated by defined spinning of PDMS on a silicon wafer to achieve the desired thickness.  In this example we use a 100 (m thick membrane which was fabricated by spinning 500 rpm for 10 seconds and then 900 rpm for 30 seconds.
2. All components are bonded together after oxygen plasma treatment with a handheld plasma device (model XYZ, company ABC).
· Scenes

· Show Delran mold and pouring PDMS into it
· Show spinner and wafer, and then photomask and UV system, then pouring PDMS on SU-8 master
· Show pouring PDMS on wafer and spinning, then curing, then peeling
· Show hand-held plasma treating surfaces, then bonding
3. Device setup 

3. Gently insert device into the plate, making sure to avoid bubbles.  Angling the device during insertion helps to expel bubbles out one side.  For actual cell-based experiments, this step must be done inside a sterile laminar flow hood.  The chances for contamination are greatly reduced once the device is inserted, so the assembly can be carried over to the incubator in a non-sterile environment.
3. Connect tubing from the source gas tank to the inlet and outlet ports of the device.  For cell-based experiments, the culture incubator should have a hole to permit tubing entry and the tubing should be connected after carrying the device and placing it into the incubator.  Be careful to avoid putting excess pressure on the device, which could deform the PDMS enough to crush the underlying cells.
3. Make sure the precision flow regulator is closed before opening the tank top regulator to avoid overflowing the device.  Start the flow of gas.  Gently open precision flow regulator to desired flow rate (50-100 mL/min).  Watch the value closely over the next hour and make adjustments as needed, since the pressure drop will change while the system equilibrates, altering the flow rate.  To avoid formation of bubbles in the media, reduce the flow rate to between 10-20 mL/min after an initial 15 minute equilibration duration with the higher flow rate.
3. After desired experimental duration, stop gas flow, remove the plate, and process cells accordingly (ie lyse, stain, count, etc.).
· Scenes

· Show setup in culture room as if there are cells.  Use culture media for liquid.
4. Device Validation
4. Calibration
4. Select the number and location of the positions on the fluorescent oxygen 
(FOXY) sensor slide (depending on the validation requirements) that will be used for oxygen measurement.  The slide contains a fluorescent ruthenium dye coating that is quenched by oxygen.
4. Expose slide directly to 0, 10, and 21% oxygen from gas tank and capture images after 5 minutes for proper equilibration.
4. Export the mean image intensity for each position and plot oxygenation concentration as dependent on fluorescent intensity.
4. Generate calibration curve by fitting linear curves to the 0-10% line and 10-21% line.
4. Heterogeneity
4. Establish multiple points spanning the width of the channel at a defined interval (ie every 1 mm).  Note that there may be overlap of images depending on the spacing.
4. Measure the oxygen concentration at the well surface through the device
4. Equilibration
4. Choose three points on the FOXY slide in which to measure the oxygen concentration.
4. Immediately after capturing a first image at ambient oxygen levels, open the precision flow regulator to start gas flow into the device.  Capture images at an appropriate interval to assess the duration and extent of oxygen concentration equilibration (ie every 10 seconds for 30 min).
· Scenes

· Show FOXY slide and explain why calibration at each point is necessary due to the heterogeneity of the coating
· Show calibration curve for a position and the curve fit equations
· Show cycling of oxygen concentration via the FOXY slide as I switch between source gas tanks
· Show cyclic profile device
5. Applications

5. Vein Experiment

5. Set up experiment using a 6 well plate. Pour enough media into plate to cover pillars.  Place vein into media.

5. A day before the experiment, soak the sterilized PDMS insert in serum-free medium to reduce the inhibit gas bubble formation in the well.  
5. Focus microscope and decide witch point images will be collected.  Go to XY tab in software and set points.

5. Set flow rate of different oxygen concentrations to pre-determined values desired for experiment. 

5. Place tubing into device. Insert device into 6 well plate.

5. Set time points and duration of experiment. Start experiment.

5. Wound Healing
5. A day before the experiment, soak the sterilized PDMS insert in serum-free medium to reduce the inhibit gas bubble formation in the well.  
5. Culture cells to 100% confluency in a 6-well plate.  
5. Create straight scratches in the monolayers with a p200 pipet tip to simulate wounds.
5. Aspirate the cell media, rinse with 5 mL media, and aspirate again. It is important to not disturb the monolayer of cells.  
5. Refill the wells with 4 mL serum-free medium to reduce cell proliferation.
5. Place insert into the wells and connect each well to a corresponding oxygen concentration.   
5. Place 6-well plate with insert on heated stage at 37◦C.
5. Capture timelapse images of the cells at the desired interval and total duration. MATLAB’s T_Scratch, a wound measuring algorithm, can be used to analyze the unhealed surface area.
Part 7: Representative Results: 
Device validation
The hypoxic insert device exhibits great improvements over the hypoxic chamber in terms of oxygen equilibration time and extent, requiring less than 2 minutes to stabilize to 0.5% oxygen.  The device membrane-well bottom gap size was the critical factor in determining equilibration efficiency, with larger gap sizes requiring more time to reach steady-state oxygen concentration values.  The device also permits a great deal of control over the spatial oxygenation in a single well, allowing the formation of multiple conditions, and even generating a cyclic oxygen profile across the surface of the well.  
Wound healing

Cell monolayers were exposed to <0.5%, 10%, or 21% oxygen and the surface area of the wound was analyzed over time.  Scratches exposed to 21% oxygen closed the slowest and 10% the fastest.  Figure 1 shows images of a scratch over the course of 17 hours.  The graphs in Figure 5 show the percent of open wound area for all three oxygen concentrations for the duration of the experiment.  (0% data will be added soon).
Tables and Figures:  Please make sure that text in all tables or figures is made in Arial font.  Figures should be submitted separately as layered .tiff or .psd files at 300 dpi.   All tables and figures should be given an appropriate title and should have a corresponding figure legend.    

[image: image1]

[image: image2]

[image: image3]

[image: image4]

[image: image5]
Discussion: Short discussion on critical steps, possible modifications, applications, significance, etc…

The device is fabricated by standard SU-8 photolithography, replica molding, and defined spinning and made entirely of polydimethylsiloxane.  Gas is introduced into the device to establish a concentration gradient between the pillar microchannel and the culture media, driving the system towards a desired equilibrium oxygen concentration.  The device has been shown to effectively modulate the temporal and spatial oxygenation inside a well, as well as modulate cellular behavior appropriately.  The spatial patterning of oxygenation is defined by the microchannel at the base of the pillar, so a variety of designs could be implemented in crafting the photomask.  Additionally, infusion of the desired gas into the gas-phase of the well is expected to improve equilibration time and extent of hypoxia.  A microfluidic mixing network could be adapted to the device to provide a means to produce novel gas mixtures from only a few stock gas tanks.  Finally, a mechanism for media exchange would eliminate the need for removal of the device from the multiwell plate, of which the cells may respond.

The device has applications in any in vitro or ex vivo experiment requiring control over oxygen concentration.  As oxygen is an important physiological variable affecting a vast majority of signaling pathways, the areas of research that would benefit is limited by the creativity of the researcher.  Some fields that would benefit from the enhanced temporal control of oxygen concentration include cancer metastasis, sleep apnea, and cardiac ischemia reperfusion injury, among many others.  For example, intermittent hypoxia has been correlated with more invasive cancers, upregulating a number of metastastis-associated genes relative to continuous hypoxia and normoxia.  Spatial control is also important, as oxygen gradients are critical in development, liver zonation, drug toxicity, and the stem cell niche. The device presented in this article will benefit a number of areas of research by providing a system with a smaller lab footprint, relatively simple operational requirements, and far greater control over oxygen exposure to cells.
Acknowledgments: This project was funded by the Illinois Department of Public Health and the National Science Foundation (DBI-0852416). 

.

Disclosures: I have nothing to disclose.
Table of specific reagents and equipment:

	Name of the reagent
	Company
	Catalogue number
	Comments (optional)

	PDMS-Sylgard 184
	Dow Corning
	
	

	Planar FOXY sensor
	Ocean Optics
	FOXY-SGS-M
	Coated microscope slide

	Gas regulator
	Omega
	FL-1472-G
	

	Gas
	Airgas
	Custom mixes
	All have 5% CO2

	SU-8 2150
	Microchem
	
	

	MDCK Growth Medium w/ L-Glutamine
	SAFC Biosciences
	M3803
	

	Fetal Bovine Serum 
	ATCC
	30-2020
	

	Trypsin-EDTA
	Sigma
	T4049
	

	L-Glutamine solution
	Sigma
	G7513
	

	
	
	
	

	
	
	
	

	
	
	
	


���


Figure 5.  Timelapsed images of wound closure 0, 7, and 17 hours after initial scratch.  Cells are delivered 21% oxygen throughout duration of experiment.  





�


Figure 2. Validation of the device with oxygen sensors. Oxygen tension within each well was characterized using a planar ruthenium oxygen sensor.  All oxygen mixtures contained balanced nitrogen and 5% CO2 for media buffering.  A) Plot illustrating the effect of post height, and thus oxygen diffusion distance between the membrane and cells, on the equilibration time and effectiveness.  Heights were established by cut-glass posts bound to the bottom of the device.  All three post sizes yield equilibration times much improved over the hypoxic chamber.  Note that time is on a log scale.  B) Plot depicting the rapid oxygen equilibration response time of the 0.2 mm gap device.  C) Multi-position linescans were also taken across the well under the microchannel to ensure homogeneity of the oxygen concentration introduced by the device.  Graph depicts the oxygen concentration measured after infusing 0%, 10%, and 21% oxygen for 10 min.    D)  Device effectively maintains 10% oxygen over 5 days.








�


Figure 3. Experimentation with more complex oxygen microchannel designs. A) Dual-condition microchannel setup yields a stable 0% and 21% oxygen profile over 14 days. B) An interdigitated and winding pattern of 500 μm width microchannels extending across the pillar results in a cyclic oxygen profile.  Note that the data only depicts one representative trial as microchannel alignment was difficult.








�


Figure 1. Schematic and diagrams illustrating device features.  The oxygen insert device is fabricated by conventional photolithography (microfluidic network), replica molding (microfluidic network and insert scaffold), and defined spinning of PDMS (gas-permeable membrane). A) The oxygen device nested into a 6-well plate.  B) Examples of 24 and 96-well pillar arrays.  C) A cross-sectional schematic of a pillar.  Oxygen flows into the device through the inlet and travels across a microfluidic network at the bottom of the pillar.  Oxygen can freely diffuse across the gas-permeable PDMS membrane at the bottom of the pillar and dissolve into the culture media.  D) A macroscope image showing the various features of a single-channel pillar from above, with bonded glass posts for the equilibration studies.





�


Figure 6. Effect of oxygen concentration on wound healing rate in a scratch assay.








[image: image6.png]


[image: image7.png]


[image: image8.png]


[image: image9.png]>

N
o

-
al

al

Oxygen concentration (%)
o

—e— 0.2 mm gap
—0— 0.5 mmgap
—»— 1.0 mmgap
—&— Hypoxic chamber

0.1

(o}

1 10 100

Time (min)

N
o

-
al

a

Oxygen concentration (%)
o

I e D S T

1 Mﬁ—n—o—a—o—n—o—c—o—a——uﬂ!—&o—o—o—mﬂ

—e— 0% Oxygen
—0— 10% Oxygen
—v— 21% Oxygen

5 & T 9 009 0 5 9 o506 sEsET

0 5 10 15 20

Position (mm)

% O 21% O

Start Start

s J 8 1 ¢ 1 1

Oxygen concentration (%)

O

Oxygen concentration (%)

20 ( Fr
15 |
10 |
5
0 >
0 5 10 15 20 25 30
Time (min)
20 A
15 4
—e— Omin —=— 24 hr
5 1 —0— 10 min —o— 72 hr
—v— 1hr —— 96 hr
—&— 6 hr —— 120 hr
0 A
0 5 10 15 20

Position (mm)




[image: image10.png]o

<

o (o] o (o] o
N

-— -~

(%) uonenuaosuod usbAxQ

(

20 {| —o— Day 2

n o 0 o
L =

%,) uolenuasouod usbAxQ

15 20

10
Position (mm)

15 20

10
Position (mm)



[image: image11.png]L Insert Pillar (PDMS)
— Well wall

— Culture Media

| Microchannel
Component (PDMS)
N Oxygen
: Microchannel
Cells Gas-permeable

Membrane (PDMS)

D
Well wall

AAAANAD A
AR A A AT

[ AN N
AR
e

\ -
28 A MY AL, 1
.1 '
W AN S

Zom==P

illar Edge.
-\
\> \\
Yo\ %

# Inlet/
Glass Posts Outlet



[image: image12.emf]Percentage of Open Wound Area for 0,10, and 21%

0

10

20

30

40

50

60

70

80

90

100

0 5 10 15 20

Hours

Percent 

10%

21%

