Title:  Determining cell number during cell culture using the Scepter™ Cell Counter 

Authors: First name, middle initial, last name  For example: Isaac S. Newton
Authors: institution(s)/affiliation(s) for each author:

· department

· Millipore Inc.

· e-mail

For example: 

Isaac S. Newton

Department of Mathematics

University of Cambridge

iLoveApples@gmail.com

Corresponding author: name
Keywords: Scepter, cell counting, cell culture, hemocytometer, Coulter, Impedance-based particle detection (Authors add any keywords you’d like)
Short Abstract: 
The Scepter Cell Counter is a handheld automated device that can be used to count cells, determine cell proliferation rates, monitor cell diameter and volume, and determine the effects of varying culture conditions on cell population distributions.  
Long Abstract:
Counting cells is often a necessary but tedious step for in vitro cell culture. Consistent cell concentrations ensure experimental reproducibility and accuracy. Cell counts are important for monitoring cell health and proliferation rate, assessing immortalization or transformation, seeding cells for subsequent experiments, transfection or infection, and preparing for cell-based assays. It is important that cell counts be accurate, consistent, and fast, particularly for quantitative measurements of cellular responses. 
Despite this need for speed and accuracy in cell counting, 71% of 400 researchers surveyed1 who count cells use a hemocytometer to do so. While hemocytometry is inexpensive, it is laborious and subject to user bias and misuse, which results in inaccurate counts. Hemocytometers are made of special optical glass on which cell suspensions are loaded in specified volumes and counted under a microscope. Sources of errors in hemocytometry include: uneven cell distribution in the sample, too many or too few cells in the sample, subjective decisions as to whether a given cell falls within the defined counting area, contamination of the hemocytometer, user-to-user variation, and variation of hemocytometer filling rate2. 
To alleviate the tedium associated with manual counting, 29% of researchers count cells using automated cell counting devices; these include vision-based counters, systems that detect cells using the Coulter principle, or flow cytometry1.  For most researchers, the main barrier to using an automated system is the price associated with these large benchtop instruments1.  

The Scepter cell counter is an automated handheld device, that offers the instrumentation and accuracy of Coulter counting at a relatively low cost. The system employs the Coulter principle of impedance-based particle detection3 in a miniaturized format combination of analog and digital hardware for sensing, signal processing, data storage, and graphical display. The disposable tip is engineered with a microfabricated, cell- sensing zone that enables discrimination by cell size and cell volume at sub-micron and sub-picoliter resolution. Enhanced with precision liquid-handling channels and electronics, the Scepter cell counter reports cell population statistics graphically displayed as a histogram. 
Protocol Text: 

2. Preparation of a single-cell suspension for counting

2.1. Dilute a single cell suspension of cells,  in phosphate buffered saline (PBS-  EmbryoMax® 1X PBS, from Millipore is recommended). A concentration of 10,000-500,000 cells/mL is optimal. (Authors, for the script, what type of cells will be used on the day of filming?)
2.2. Transfer 100 µl of diluted cells to a 1.5 mL microfuge tube.

3. Performing cell counts using the Scepter

3.1. Turn on the Scepter cell counter by depressing and holding the toggle on the back of the instrument and wait for on-screen instructions to appear.
3.2.  (Authors, please provide additional info on the tip). Scepter pipette tips consist of a precision molded sampling chamber, with an electronic sensing zone and integrated cell sensing electrodes.   The tips make it possible to discriminate cell size at sub-micron resolution, and cell volume at sub-picoliter resolution. 

3.3. When prompted, attach a tip to the end of the Sceptor unit with the  electrode sensing panel facing toward the front of the instrument.  The Scepter will then display detailed on-screen instructions for each step of the counting process as follows:

3.4. When the screen reads “Hold down the plunger to begin,” Depress the plunger.  The screen will then read “submerge the tip.”   Submerge the tip into the cell suspension solution and release the plunger to draw 50 µL of the cell suspension into the tip.  (Is this process slow?  Does the instrument regulate the speed at which fluid is drawn into the tip?)
3.5. The cell sensing zone of the tip will detect each cell as it is drawn into the tip, assessing cell size and cell volume.  Once the sample has been loaded, the instrument will beep, and “Sample loaded” will be displayed on the screen. 

3.6. The instrument will then calculate the cell concentration and display a histogram of cell size or diameter on its screen. 
3.7. After the count is complete, the instrument’s screen will read “Count Complete.  Please remove tip and discard.”  Following the prompt, remove the tip. (Are tips reusable or must they always be discarded?)
4. Data analysis

4.1. After the count is complete and the histogram is displayed on the instrument, apply gates by pushing the toggle button.  The initial gates used will are set either automatically based on the hisogram priofile, or are set to the same gates used in the previous count.
4.2.   The gates can be moved manually to fine-tune the data. 
4.3. Move the toggle to either side to switch the display between date, time, mean cell size, and mean cell volume.
4.4. Up to 72 histograms can be stored on the instrument itself.  To upload the data to a personal computer, attach the instrument via the USB cable pot on the display screen.  Then, use the Scepter Application software to upload the data.  (Please indicate any additional instructions for uploading the data here.) 
4.5. The Scepter application Software displays imported volume and size distribution date.  Downloaded files can then be exported to Microsoft Excel, with data in bins of specified cell diameters or cell volumes.
5. Representative Results: Linearity and precision of Scepter cell counting across multiple cell lines

5.1. 19 different cell lines were prepared for counting and theoretical starting concentrations were determined using a Coulter Counter ®.  These cell suspensions were then further diluted to obtain cell suspensions at theoretical cell concentrations.  The Scepter cell counter was then used to determine cell concentrations.   Figure 1 shows 5 dilutions that were measured for 5 representative cell lines along with the theoretical cell concentration.  The high degree of linearity as indicated by the R2 values, demonstrates that Scepter counting is a reliable method for the cell lines tested, across a wide linear operating range. 
[image: image1.png]™

atcal Call Concantratin (ot

o K62 r2-09982
U266 12209996
e Hela 1220999

- Mesenchymal SC 12209938

—a 520998
8 Theoretica cell concentration (celsimt)




5.2. Overall relative accuracy of Scepter counting is shown in Figure 2 by comparing counts obtained from scepter counting with counts obtained from Coulter Counting.  Plotting log of Scepter cell counts vs log of Coulter counts shows that across all cell lines tested, Scepter counts match Coulter counts.

[image: image2.png]Scepter Call Concentration

1000000

100000:
10000
.
1000
.
o
1000 ot o000

22 Coulter Cell Concentration

Tocooon




5.3. Counts of 19 cell lines were performed using a ZZ™ Coulter Counter, the Scepter cell counter, an automated vision-based counter such as Vi-Cell® or Countless ™ system, and a hemocytemeter.  Counts were performed according to the manufacturer’s instructions, using the same cell starting suspension and identical dilutions.   Results in Figure 3 show that cell concentrations measured by Scepter counting closely match theoretical cell concentrations with high linearity.
 [image: image3.png]Messurad Coll Concantraion (Clls/mb)

ThaoraticlCallConcantration Catu/mt)

o Cauter caunter o Scepter cytometer
o Hemocytameter
e Autamted vision-based counter

Figura 6. COS7 ce caunts using varios ystems. Daa shownare avrage ffour relcates Erorbars epresent stanéad deition COS) cets
epresent one cut of 19 tta el s ested





5.4.  The percent coefficient of variation (CV) was calculated for each dilution of each cell line for each counting system.  Shown are the average %CVs (over 19 cell lines) with respect to cell concentration and counting method.  In the operating range of the Scepter cell counter, it is more precise than vision based counting and hemocytometer,  displaying smaller standard deviations and smaller average coefficients of variation.  

[image: image4.png]Avrage %€V of Mesured

o Cauter counter —a— Scepter cytometer.

o Hemocytameter
= Autamated vision-based counter

Fipire 7. Th prcantcouficien of varition (V1 was caclated or eschdiuion of sch cel i for exch countingsysers Shown e th average
s lover il 19 cll ) with respect tocll coartrstion andcounting method.In e persing ange f th Sceper <o ccunter, 13 more

preis tan i b4 counting.nd hacstamery




5.5. As shown in Figure 3, Scepter counting is ten times faster than hemocytometry and faster than other automated counters too.  Here, SF9, MCF7, and HEK293 cells were counted using the methods described here.  The average time required to measure cell concentration was recorded.  Note that hemocytometry was performed only once per sample.

[image: image5.png]Time (seconds)

[ =] i

Scepter Flow Coulter Automated Hemocytometer
cytometer eytometer counter vision-based counter




Tables and Figures:  Please make sure that text in all tables or figures is made in Arial font.  Figures should be submitted separately as layered .tiff or .psd files at 300 dpi.   All tables and figures should be given an appropriate title and should have a corresponding figure legend.    Please use the following format for file names

####_Millipore_Figure1.tif

(#### = JoVE submission #)

Discussion: 

Comparing the performance of the Scepter cell counter to results from other counting methods, we conclude that this new handheld, automated cell counter delivers precise, fast, and reliable cell counts over a wide operating range. The functionality of Scepter counting is a result of the precision engineered technology embedded into the  tip  and the sophisticated counting instrumentation based upon the Coulter principle.  The performance quality suggests Scepter counting as method to increase the speed and improve the reproducibility of cell-based assays.  
Acknowledgments: List acknowledgements and funding sources.
Disclosures: describe potential conflicting interests or state “I have nothing to disclose”.  

Table of specific reagents and equipment:

	Name of the reagent
	Company
	Catalogue number
	Comments (optional)

	Scepter Handheld Automated Cell Counter
	Millipore
	PHCC00000
	

	Scepter Tips, 60 µm
	Millipore
	PHCC60050
	

	Universal Power Adapter
	Millipore
	PHCCPOWER
	

	Scepter O-Ring Kit
	Millipore
	PHCCOCLIP
	3 O-rings and one filter cover

	EmbryoMak® 1X DPBS
	Millipore
	BSS-1006-A
	1L


References:  Please give 10 or more references.  Please use the same format you would use for a Nature article.  For example:

1.    Johnson, A.B., Jackson, C. & Smith, C.D. Mechanism of wing development in flies. J. Biol. Chem. 282, 332-342 (2005). 

2.    Ying, E.F. & Yang, G.H. Gene X1 is important for mouse to walk. Curr. Biol. 9, 245-263, (1997).

References (additional references requested) 

1. Barghshoon, S. Cell Counting Survey. Millipore 2009 Feb. 

2.  Tucker KG, Chalder S, al-Rubeai M, Thomas CR. Enzyme Microb Technol 1994 Jan. 16(1):29-35. 

3.  Houwen, B. Fifty years of hematology innovation: the Coulter principle. Medical Laboratory Observer 2003 Nov.

