
Submission ID #: 2155
Editor Name: Christina Zelano
Videographer name: Ian McGuire
Film Date: 5/11
Authors and Affiliations: 
Prasad R Shirvalkar, Matthew L Shapiro

Prasad R Shirvalkar
Department of Neuroscience, Brain Institute

Mount Sinai School of Medicine

PShirvalkar@gmail.com

Matthew L Shapiro

Department of Neuroscience, Brain Institute

Mount Sinai School of Medicine

Matthew.Shapiro@mssm.edu
Title: 

Design and construction of a cost effective headstage for simultaneous neural stimulation and recording in the water maze
Authors, please fill out the brief questionnaire below.   
A. Will you require assistance with video microscopy, such as filming a complex dissection or microinjection technique (Y/N, please specify steps by number. Also, please list make and model of your microscope)?  ______Nikon SMZ645 Stereo microsope____
B.   Does your protocol include detailed, step-by-step, descriptions of software usage (Y/N, please specify steps by number)? __N_______

C.  Which steps of your protocol will viewers benefit most from having filmed? Please list 4-6 steps__Part (3) 4.1 - 4.6____
D.  What is the single most difficult aspect of this procedure?  

The single most difficult aspect is soldering the small electronic components on to the Printed Circuit Board.
1. Introduction (Schematic Overview and Interview)
A. Schematic Overview (read by voice talent at JoVE):

Authors, please select from “Procedural Narrative” or “Conceptual Narrative” and complete the statements below. Please do not add additional steps.  Then, attach your finished graphic overview.  See accompanying instructions for details and examples.  
Procedural Narrative:
The overall goal of this procedure is to _______(insert goal here)___________________ (Intro).
This is accomplished by first ________(insert 1st step)___________________________(P1).
The second step of the procedure is to _____(insert 2nd step)_____________________(P2).
The third step of the procedure is _______(insert 3rd step)_______________________ (P3).
The final step of the procedure is _____(insert 4th step)_________________________ (P4).

Ultimately results can be obtained that show _( insert type of results e.g. changes in protein localization)________ through _(insert method used to assay - e.g. immunofluorescence microscopy)______(P5).
Conceptual Narrative: 
The overall goal of the following experiment is to __construct a headstage - preamplifier system to simultaneously record and stimulate chronically implanted electrodes in awake behaving animals (Intro).
This is achieved by first designing a printed circuit board for a headstage with a targeted number of recordings and stimulating channels. (C1). 
As a second step, we solder the electronic components to the circuit board,which __amplify recording electrode and reference signals, ground the electronics and the animal, and convey stimulation signals through separate channels _________ (C2).   

Next,  we construct headstage and tether cable connectors in order to provide an interface between the headstage and the animal's implant on one side, and the wire tether cable on the other side, and to waterproof the system for recording in the water maze  (C3).
Results are obtained that show _reliable Local field potential recordings from the rat hippocampus (top) and poor quality recordings because of loose connections or contacts (bottom) _ based on  ___digitizing the sampled signal from the rat hippocampus and band-pass filtering the output from 0-100 Hz. __(C4).
Paste a copy of your graphic overview here.  The original file should be adobe illustrator (preferred) or powerpoint (see instructions) and should be uploaded through your online submission on the JoVE website.   

[image: image1.emf]
B.  Interview: (Said by you on camera. Don’t forget to smile!)  
Authors: Below are statements we would like you to complete that are complementary to the information contained within the schematic overview.    Only one statement should be chosen and completed per author who will be on camera demonstrating the protocol.    
1.1) The main advantage of this technique over existing methods, like _purchasing a commercially available headstage amplifier, is that _this method can be thousands of dollars cheaper and easily customized for specific stimulation and recording experiments. 

1.2) This method can help answer key questions in the Neuroscience/Neurophysiology_______ field, such as  those related to brain mechanisms of cognition and behavior, by helping to identify neural signals that encode, maintain, and decode information .  

1.3) The implications of this technique extend toward therapy (or diagnosis) of_cognitive disorders ______, because  recording brain activity can reveal key features of neural information coding and stimulation can be targeted to improve those coding mechanisms and thereby enhance brain function. _______.  

1.4) Though this method can provide insight into _recording and stimulating awake behaving animals it can also be applied to other systems such as __electromyogram (EMG) recording of motor signals, use in acute physiology to reduce noise, and can be modified for use with fiber optics and optogenetic manipulations.__________.
1.5) Generally, individuals new to this method will struggle, because ___precise soldering of small electronic components is difficult. ___________.
1.6)  I/We first had the idea for this method, when I/We wanted to record and stimulate brain activity simultaneously while rats swam in a water maze task and no commercially produced headstages sufficed. _________.
1.7) Visual demonstration of this method is critical, as the _PCB design and soldering steps are difficult to learn, and because _many biologists do not have prior experience with electrical engineering or circuit design.______________.   

Protocol (read by voice talent at JoVE):

Authors:  In order to ensure that your protocol can be filmed in a single day, the protocol text must be limited to 30 steps – each step being defined as 3 lines of 12 pt text in our formatting style below.  This amounts to 3 pages of protocol text.  The scope of the scripted protocol text should include only those aspects of the procedure that require visualization in order to be well understood.     

2. Part 1. Overview of the headstage pre-amplifier system
2.1. The low-cost headstage pre-amplifier system featured in this video allows simultaneous neural recording and stimulation, and can be made waterproof for use with swimming animals. 

2.2. A completed system consists of the following components:
2.2.1. A computer-designed integrated circuit board with four recording channels and two stimulation channels., The channel number is easily customized.

2.2.2. Electrodes that relay brain  signals to the headstage and stimulation signals to the brain via separate connectors on the animals head stage.

2.2.3. A tether cable and connector that relays the recording and stimulation channels through a slip-ring commutator to a computerized physiology workstation 
2.2.4. Two forms of waterproofing on the head stage and on the integrated circuit. Now lets take a look at the design and layout of the headstage.
3. Part 2. Design and Layout of the Headstage 

3.1. A printed integrated circuit board with 4 unity-gain source followers and 2 pass through stimulation channels is designed using a commercially available software CAD package (www.expresspcb.com).
3.2. First we create connector pads on both sides of the board for our pin connectors on the tether and animal connection.
3.3. A Quad-FET operational amplifier can accept upto 4 channels of +/- input (ie. 4 electrode wires). 
3.4. We will need two power lines, each which can be connected to ground via 0.1(F capacitors to prevent possible induced oscillations in the operational amplifiers.
3.5. To make a unity gain amplifier circuit, we route the signal into the positive input of the operational amplifier (x4) and then connect the output back to the inverting input.
3.6. Then, we add 1 MOhm resistors in series with the positive input to ground for the same reason as 3.4,
3.7. For the stimulation channels, we lay pass-through traces connecting outputs to inputs, using vias to guide the traces to the other layer of the board and back (in this example)..
3.8. Finally, we create a large ground plate on the back side layer to connect all vias that are meant for ground. The ground plate then should be connected to the ground pin out to the tether and animal side.
3.9. A single headstage is designed, and the design pattern is tiled across a larger sheet of circuit board to provide many headstages at a lower cost. 
3.10. The CAD file is transmitted to the printed circuit board (PCB) manufacturing service where double sided, two layer boards are printed using top and bottom copper layers with all holes plated through. 
3.11. The boards use tin/lead solder plated traces and pads that correspond to industry standard FR-4 laminate. 
3.12. Individual headstages are cut from the large tiled PCB sheet and electronic components are connected with solder. Now lets see how to connect the electronic components.
4. Part 3. Preparation and soldering of electronic components
4.1. Begin by organizing the components on the circuit board and prepare them for soldering. 

4.2. Place a small amount of solder on the tin/lead pad to be soldered, and bring the tip of the soldering iron close to melt and adhere the solder to the pad. 
4.3. Solder the capacitors first, resistors next, and then the operational amplifiers. 
4.4. Because the op-amps can be destroyed by over heating, use a variable heat soldering gun and the lowest heat necessary to melt the solder. 
4.5. It is best to minimize the heating time used to make a good solder joint. We found it convenient to create a small reservoir of solder that flows into the footprint pad by applying solder along the communication or power traces.
4.6. Now use an ohmmeter to test that the resistance across the resistors meets the expected value.
4.7. Inspect each pad closely to ensure that no solder is connecting or bridging adjacent pads. If there is a solder bridge, heat and melt the solder and use a tweezer to break the bridge. Now lets see how to assemble and fabricate tether interfaces.
5. Part 4. Assembly and fabrication of tether interfaces.
5.1. First insert a piece of 23 Ga hypodermic tubing into one of the unused ports of the 9 pin ABS socket for mechanical support.
5.2. Cut six 1 cm length pieces of insulated magnetic wire for connecting the female amphenol pins to the ABS socket.
5.3. Now using a small amount of solder in each pin, heat the pin from the tip to solder the female pins to the magnetic wire.
5.4. Solder the other end of each magnetic wire into the ABS sockets.
5.5. Position the male - female MillMax pins over the pads on the tether side of the headstage, and squeeze the pins over the PCB to ensure a pressure fit. Then, solder each pin to each pad in order.
5.6. Next, position the ring nut over the neural implant connector and push all the female pins into the desired slots on the plastic base.
5.7. Apply acrylic glue to the top of the ABS socket connector and secure it to the base of the headstage.  Finally, glue the 23Ga tube to the tether side of the headstage for support. Now lets see how to waterproof the headstage.
6. Part 5. Waterproofing the headstage

6.1. To waterproof the headstage, encase the entire integrated circuit and any exposed metal surfaces with 5 minute epoxy. Allow 5-10 minutes for initial curing, and 60 minutes for permanent curing.
6.2. Apply poster tack putty to the surface of the 9 pin ABS socket around the female pins protruding from the animal end.
6.3. For maximal waterproofing, the ABS socket and ABS plug should be mated and the ring nut should be tightened. 
6.4. Wrap a small strip of parafilm around the interface to prevent water from contacting the pin connectors.
7. Representative Results-(second to last section) 
7.1.  Now we’ll show some representative results. A good recording will follow an input source signal without cutting in and out during movement. To test the headstage, a function generator is used as an input to each recording pin, and the waveforms at the Mill Max header output end are evaluated. Physically moving the headstage should not influence the shape of recorded waveforms.

[Do you have a figure for this section that shows the waveforms that are evaluated? It is now included in the conceptual narrative.]
Authors, see instructions below

Representative Results Narration

(C4) 
INSTRUCTIONS FOR AUTHORS:

Please ensure that the representative results narration is appropriate and correctly describes your images, movies, or figures.  Our editors have ensured that the results are written in our format.   

We consider this section a critical aspect of the video, because here is where you provide validation for your experiments.  For example, if this is a cell culture preparation, this section is where the video will show your cells at various time points following culturing.  If this is an imaging prep, then this part is where you will show examples of your imaging experiments.  

Please limit the extent of narration to no more than 2-3 lines of text per image or movie file being described.   Figures with multiple panels submitted with the original protocol should be broken up so that each panel is a separate image.   Like the schematic, each image or movie file supplied in the results should be referenced by annotation in parenthesis, however for the results, the specific filename should be given in parenthesis.  

Below is an example of results text:

EXAMPLE REPRESENTATIVE RESULTS

5.  Evaluation of Morpholino Injection and Knockdown
5.1   Representative results of both morpholino injection and mRNA injection are shown here. The    

        uninjected control at 48 hours post fertilization looks normal, as expected 

        -LAB MEDIA: 0123_PIname_Figure1.tif  (Replace 0123 with your jove video #)

5.2   However, embryos injected with the morpholino heg_e3i3_egfr1, which knocks down Heg isoforms

                     containing the first of two EGF-like repeats, exhibit brain edema.


        -LAB MEDIA: 0123_PIname_Figure2.tif


5.3   Injection of heart of glass mRNA also produced an obvious phenotype. At 24 hours post fertilization, 

        the heads of the uninjected controls look normal 

        -LAB MEDIA: 0123_PIname_Figure3.tif

5.4   Conversely, some of the embryos injected with the mRNA exhibit cyclopia     

                     -LAB MEDIA: 0123_PIname_Figure4.jpg

Please visit the following URL to see an example of how the results will look when complete:

http://www.jove.com/index/Details.stp?ID=1597
5. Conclusion (said by authors on camera

Authors: Below are statements we would like you to complete that summarize and conclude the video. Only one statement should be chosen and completed per author who will be on camera demonstrating the protocol.    
5.1. Once mastered, this technique can be done in ___2 hours_________ (hours/min) if it is performed properly.
5.2. While attempting this procedure, it’s important to remember to _make multiple headstages at the same time to establish replicability and provide quick a backup for damaged headstages.__________.
5.3. Following this procedure, other methods like __targeted intracranial drug infusion___________ can be performed in order to answer additional questions like _what neurotransmitters or modulators are critical for specific mechanisms of behaviors ____________.
5.4. After its development, this technique paved the way for researchers in the field of __neurophysiology________ to explore __cross-frequency coupling__________ (subdivision of field, disease, natural phenomenon) in  the hippocampus of cognitively impaired and normal rats __________( model organism, patient demographic, organ system).
5.5. After watching this video, you should have a good understanding of how to _design a headstage with recording and stimulation capabilities for use with a chronic awake behaving experimental setup____________ (restate overall goal of the procedure mention specific steps).
5.6. Don't forget that working with __solder___________(reagent, pathogen, instrumentation) can be extremely hazardous and precautions such as __avoiding inhaling vapor residue or wearing safety goggles__________ should always be taken while performing this procedure.   
Provided Media

Authors, Please list all images, movie files, or 3-D rendered animations that can be included in the video per editor’s request.  The step in the script/video where these images will be inserted should be specified.   For example:

6.2 –  0123_PIname_Figure1.tif -  dual color imaging of tumor angiogenesis at 40X 

6.2 –  0123_PIname_Figure2.tif -  dual color imaging of tumor angiogenesis at 100X

Formats:  For static images we prefer .tiff files at dimensions of at least 720X480 pixels and 300 dpi.  The higher resolution, the better.  Likewise any exported movie files should have at minimum these dimensions and be rendered to .mov, .mp4, or .avi files.  

Insert your media filenames here.

1.1 - 2155_Shirvalkar_Figure1.ai
3.1 - 2155_Shirvalkar_Figure2.ai

3.2-3.9 - 2155_Shirvalkar_Jove_Fig3-8.pdf

General Preparation

It’s critical for a smooth and organized shoot that all reagents are accounted for, in advance.   

Any overnight or long incubation steps should be recognized and specimens/samples be prepared in advance so that prior steps can be recorded and shooting can continue with pre-prepared specimens/samples.  

All tubes/flasks should be pre-labeled neatly before we arrive.  

Ex. Luciferase assay done in 96 well plates should be labeled with negative/positive control wells and experimental samples are labeled accordingly.










( 2010, Journal of Visualized Experiments


_1333811056.pdf


Ch.1 Input 


Ch.2 Input 


. . . 


op-amp
Ch. 1 output


Ch. 2 output


x N recording channels


Pass through Stimulating paths


Implant (ABS plug)


Headstage
Amplifier


Tether


(C1)


(C2) Operational Amplifiers
1 MΩ Resistors
0.1 µF Caps 


(C3)


ABS Socket
Ring Nut


MillMax Header


PCB headstage


Time (sec)


1 2 3 4


N
or


m
al


iz
ed


 A
m


pl
itu


de
 


1


0


-1
1


0


-1
BAD


GOOD


(C4)






