Construction and testing of a 32 Channel Phased-Array for Magnetic Resonance Brain Imaging
Authors: 
Boris Keil, Christin Y.-M. Sander, Veneta Tountcheva, Jennifer A. McNab, Kyoko Fujimoto, Azma Mareyam, Christina Triantafyllou, Lawrence L. Wald

Authors: institution(s)/affiliation(s) for each author:

Boris Keil 

A.A. Martinos Center for Biomedical Imaging, Department of Radiology

MGH - Massachusetts General Hospital, Harvard Medical School

keil@nmr.mgh.harvard.edu
Christin Y.-M. Sander
A.A. Martinos Center for Biomedical Imaging, Department of Radiology

MGH - Massachusetts General Hospital, Harvard Medical School

Department of Electrical Engineering and Computer Science 

MIT – Massachusetts Institute of Technology 

csander@mit.edu
Veneta Tountcheva 

A.A. Martinos Center for Biomedical Imaging, Department of Radiology

MGH - Massachusetts General Hospital, Harvard Medical School

veneta@nmr.mgh.harvard.edu
Jennifer A. McNab

A.A. Martinos Center for Biomedical Imaging, Department of Radiology

MGH - Massachusetts General Hospital, Harvard Medical School

jmcnab@nmr.mgh.harvard.edu

Kyoko Fujimoto 

A.A. Martinos Center for Biomedical Imaging, Department of Radiology

MGH - Massachusetts General Hospital, Harvard Medical School

kyoncat@nmr.mgh.harvard.edu
Azma Mareyam
A.A. Martinos Center for Biomedical Imaging, Department of Radiology

MGH - Massachusetts General Hospital, Harvard Medical School

azma@nmr.mgh.harvard.edu
Christina Triantafyllou, 

A.A. Martinos Center for Biomedical Imaging, Department of Radiology

MGH - Massachusetts General Hospital, Harvard Medical School

McGovern Institute for Brain Research

MIT – Massachusetts Institute of Technology


" 

ctrianta@mit.edu



Lawrence L. Wald

A.A. Martinos Center for Biomedical Imaging, Department of Radiology

MGH - Massachusetts General Hospital, Harvard Medical School

Harvard-MIT Division of Health Sciences and Technology
MIT – Massachusetts Institute of Technology

wald@nmr.mgh.harvard.edu
Corresponding author: Boris Keil
Keywords: Magnetic Resonance Imaging, MRI, Coil, Phased-Array, 32-Channel, Preamplifier Decoupling
Short Abstract: (50 words maximum)
In this video we describe an optimized procedure for constructing a 32 channel array coil for brain MRI.  We describe the basic tools and procedures for tuning and decoupling the coil on the bench and an optimized work-flow for efficiently constructing high element count array coils.
Long Abstract: (150 words minimum, 400 words maximum)

In the last decade, parallel detection of the Magnetic Resonance Imaging (MRI) signal with multiple receive-only surface coils has proven valuable for increasing image sensitivity and acquisition speed. The success of parallel imaging methods has driven the design of MR receive arrays with an increasing number of elements. The high degree of parallelism requires systemized design, construction and testing in order to implement the large number of tuned receive circuits with minimal mutual interaction and has altered the workflow for how we construct receive arrays.

The main aim of this work is to inform MRI researchers about the basic procedures for phased-array construction and describe an optimized protocol for constructing, tuning and decoupling a highly parallel array coil. The goal is to provide a better understanding of the basic experimental RF tools and procedures to facilitate the efficient design and construction of highly parallel MRI receive-arrays. We demonstrate the protocol with the construction of a 32-channel brain array of overlapped circular surface coil elements for pediatric subjects, although the basic principles apply to any array of overlapping loop elements. 

The video will go through the following steps: i) the layout of the array elements, ii) creating and tuning a single loop element, iii) estimating the coil quality factor, iv) adjusting the single loop and match circuit to optimize preamplifier decoupling and v) active PIN diode detuning circuitry, v) placing the neighboring elements to allow them to be efficiently constructed and inductively decoupled, vi) assembling the array, vii) decoupling the array elements from one another, viii) tune and match the coil elements, and iix) performing final bench tests. The final array coil is tested for signal coverage, Signal-to-Noise Ratio (SNR), noise correlation, and accelerated imaging performance on a 3 Tesla MRI scanner using a pediatric-sized head phantom.

Protocol Text:

In this video, we describe the step-by-step procedure which we have developed for constructing and tuning highly parallel array coils. We will describe an array geometry with coil overlaps to reduce nearest-neighbor decoupling, although the procedure becomes simpler if a gapped design is used (in which case there is no geometric overlap to optimize).  If shared-capacitive decoupling is used, the procedure is also similar, with the adjustment of the shared capacitor replacing the geometric optimization. Additionally, we used a design, where the preamps are mounted adjacent to each coil element to minimize losses between the inductive loop and the first stage of amplification. The close proximity between loop and preamp eliminates the need for a long cable between loop and preamplifier. In addition to reducing losses, this approach opens a new degree of freedom to match the coil’s impedance so that the load impedance seen by the preamplifier differs from the traditional 50 Ohm. This load impedance must be matched to the load impedance ZNM required to operate the preamplifier at the lowest noise figure but can, in principle, be chosen based on preamplifier design considerations or to facilitate some aspect of the coil matching circuit. Choices other than the traditional 50 Ohm complicate certain measurements, where it is necessary to temporarily transform the coil impedance to 50 Ohm so that it can be connected to the port of a network analyzer (with 50 Ohm input impedance). Nonetheless, we describe the construction procedure for this general case. 

Part 1: Helmet Design

The helmet size was obtained from the surface contours of aligned 3D MRI scans from N~20 corresponding aged children. The helmet shape was taken by dilating the 95% contour to accommodate 3mm foam padding. Using a 3D CAD program, the helmet was designed as a deep posterior segment covering all but the face and forehead so the child can lie down into the posterior coil section. A separate “frontal paddle” is then positioned over the forehead. This configuration simplifies the need to mate the two segments and avoids a single helmet which comes down over the head. It maintains an open area over the eyes and face. 

The helmet model is shown in Fig. 1. The geometric structure of the coil element layout (see below) was engraved on the helmet as part of the CAD design (see Fig. 1). Additionally, standoffs for the preamplifiers were included into the 3D design which is shown after construction with a 3D printer in Figure 2. We also constructed a size-matched head-shaped phantom derived from the statistical head surfaces. We placed this phantom in the tightly fitting coil array and used it for loading the coil during adjustments at the bench as well as test imaging acquisitions on the MRI scanner. The final design of the helmet parts and phantom were printed in acrylonitrile butadiene styrene (ABS) plastic using a rapid prototyping 3D printer. 

Part2: Loop layout 
Given the geometrical constraints of the helmet and knowledge of the extent of anatomical coverage desired, the next step is to layout the pattern of overlapping circular coils which geometrically “tile” this 2D space. We determine the tiling pattern using hexagons and pentagons (where each pentagon or hexagon represents a circular loop element) either in the CAD layout (see Fig. 1) or by cutting the hexagons and pentagons from paper or cardboard and physically laying them out on the coil former1. The goal is to completely tile the curved surface using these elements. One tile choice is to place one hexagon surrounded by six other hexagons; a pattern useful for tiling locally flat sections. Alternatively, a “soccer-ball” tiling pattern can be used with hexagons and pentagons (one pentagon surrounded by 5 hexagons). This soccer-ball pattern tiles a spherical space. Thus the hexagonal pattern is useful for tiling flat sections of the body and the soccer-ball geometry is useful for adding local curvature. Given a helmet, we can then lay out a suitable combination of hexagon subsections and pentagon subsections adjusting the size appropriately to meet the desired number of channels. For the infant array under construction, the shape of the bottom part incorporated 25 hexagons and 3 pentagons. The anterior part (forehead paddle) completes the 32 channel with additional 4 overlapped elements. 

Part2: Loop Design 
The diameters of the loop coils corresponding to the pentagon/hexagon tiles on the helmet are determined from the size of the pentagon/hexagon tiles; the loop diameter is slightly larger than the diameter of the circle which inscribes the vertexes of the hexagon/pentagons. Next, we construct a pair of test loops to determine the values of the needed capacitors and fine-tune the final size for each coil element. We recommend that the researcher complete these test loops and validate them with imaging tests prior to completing the construction of the larger array.

We used 16-awg thick tin-plated copper wire to form the loops with bridges to allow the coil conductors to cross-over one another without touching. We divided each loop symmetrically by two gaps where we place the capacitors (see schematic circuit in Fig. 3, and photo in Figure 4). Thus, each half of the loop is joined at the top by the tuning capacitor and at the bottom by the output-circuit. These two sub-circuits are constructed on separate small circuit boards. The tuning capacitor circuit board contains a variable capacitor C3 to fine-tune the loop resonance to 123.25 MHz (3 Tesla for the Siemens system) and a series fuse F for passive protection against large currents potentially induced during transmit. Figure 4 shows the tuning of a single loop with a double probe S12 measure as well as the determination of the Q factor of the coil both alone and with the other elements in place (see Part 4). The output circuit-board (shown in detail in the photo of Fig. 5) contains a capacitive voltage divider (C1, C2) to match the elements output to the impedance, ZNM, desired by the preamplifier2 (Siemens Healthcare, Erlangen Germany) for an optimized noise match. The drive-point circuit board also incorporates an active detuning circuit across the match capacitor. The active detuning is achieved using a PIN diode D in series with a variable inductor L, which together with the match capacitor C1 resonates at the Larmor frequency. Thus, when the PIN diode is forward biased (transmit mode), the resonant parallel LC1 circuit inserts a high impedance in series with the coil loop, blocking current flow at the Larmor frequency during transmit. 

Part 3: Tuning the Active Decoupling

As described above, the trap LC1 together with the pin diode, detune the loop during transmit with the body RF coil. To test that the detuning is adequate, we measure first the tuning of the LC1 circuit using a single pickup loop (“sniffer probe” marked in Fig. 5 and Fig 12c) connected to a network analyzer S11 measurement as shown in Figure 5. This trap circuit tuning step is done without the main coil loop being tuned (variable capacitor was still missing) and with the PIN diode forward biased (conductive state). After initial tuning of the trap circuit with the small sniffer probe, we attached the variable capacitor C3 to close the circuit of the main loop and tuned the coil to 123.25 MHz using a double probe controlled via the S21 measure (Fig. 6). Holding the double probe on the same position, but actively switching the forward bias of the PIN diode on an off, we measure the loop response in the tuned and detuned state (Fig. 6 middle). We expect the active detuning to produce S12 >35dB isolation between these two states (see Fig. 6).
Part 4: Coil quality factor

After the determination of the loop sizes and needed components, we calculated the coil quality factor (Q), which describes how much the reactive impedance dominates over the resistive impedance in the loop coil. It informs us of the power dissipated in loss mechanisms relative to the energy stored in the tuned circuit. Under sample-loaded conditions, the resistive impedance is the sum of the coil and sample resistances. We measured the ratio of the coil quality factor in terms of an unloaded (Qunload) and a loaded coil (Qload) where the load is the 3D printed head shaped phantom filled with physiological saline and Gadolinium. We measure the ratio of Qunload and Qload for both a single loop element and a loop under test surrounded by its six non-resonant neighboring elements. The latter case informs us about the losses in the copper in nearest neighboring elements3,4. The procedure for measuring the Q of the loops using a loosely coupled double probe is shown in Fig. 4. Care is taken that the probes are not too close to the loop under test (and thus perturbing the measurement) and that conductive material (such as tools and solder) is removed from the area during the test.

Part 5: Preamplifier Decoupling 
Preamplifier decoupling is used to reduce coupling between next-nearest and further neighbors5. Optimization of the preamplifier decoupling is a critical step in constructing highly parallel arrays. The goal is to design the preamplifier/coil circuit so that the preamplifier performs a voltage measurement across the loop. Thus the output circuitry (match + coax + preamplifier) forms a series high impedance in the tuned loop reducing current flow and thus reducing inductive coupling with other loop elements.  

To transform the preamplifier input impedance to a high impedance in the loop, we first transformed it to a low impedance (short-circuit) across a parallel LC circuit (L and C1 in Figures 3, and 5) tuned to the Larmor frequency. This parallel LC circuit, in turn, introduces a high serial impedance in the coil loop. In this mode, minimal current flows in the loop and inductive coupling to other coils is minimized. We used the same L and C1 for the active detuning trap. We closed the parallel LC1 circuit by transforming the preamplifier input impedance to a short across the PIN diode shown in Fig. 3 and 5. This is done by carefully controlling the cable length (42 mm in this case) between preamplifier and coil terminal so that it transforms the input impedance of the preamplifier to a short across the diode.

We adjust the preamplifier decoupling viewing the S12 versus frequency for the coil in the tuned state (PIN diode reverse biased) with the preamp in place and powered on using the same decoupled double probe, but with reduced power output from the network analyzer (‑25dBm). It is important that the minimum of the S21 “dip” is at the Larmor frequency (Fig. 7) thus verifying that the correct impedance transformation at the Larmor frequency was performed. Low-loss circuitry between the preamp input and the loop insures a high degree of preamplifier decoupling which can be indirectly measured by assessing the depth of the “dip” in Figure 7. 

If a quantitative measure of the preamplifier decoupling is desired, the measurement can be repeated with the preamplifier replaced by a load with the conjugate complex impedance Z*NM of the preamplifier’s noise-match impedance ZNM. We constructed of a dummy preamplifier board with a conjugate complex input impedance Z*NM. Then a comparison of the S12 at the Larmor frequency with this board and with the preamp quantifies the preamplifier decoupling6.

Part 6: Determining the Critical Overlap
In this step the overlap between neighboring elements is adjusted to null their mutual inductance. Before performing this step, we must have completed the loop and active decoupling tuning of Part 3 for all loops on the helmet. We estimated the critical overlap ratio of a pair of loops by carefully moving those loops towards each other while measuring the S12 parameter between the tuned coils (with any other loops biased to the detuned state using their active decoupling circuit).We monitor the S12 interaction between the two neighboring coils using a test probe that plugs into the circuit board at the position of the preamplifier (see Fig. 8). In our case, the impedance of the circuit at this point is not 50 Ohm. Thus, we needed to temporarily transform this load impedance to the 50Ω input impedance of the network analyzer for this measurement. This transformation (from ZNM to the 50 Ohm expected by the network analyzer) is done on the test probe we plug into the preamplifier socket (see Fig. 8). If a preamplifier with 50Ohm noise match requirements is used, this step can be omitted.

Part 7: Array Assembly

After successfully testing the first loops, we assembled the whole array: First, all the drive points and capacitor solder pads, made out of FR4 circuit material, were mounted on the helmet. Second the 16-awg thick tin-plated copper wires are cut, bent and assembled over the whole array. Third, the preamplifier board including the 42 mm long coaxial cable were mounted to the designated holder (Fig. 2), as well as wiring up the preamplifier and bias tee circuitry to the plug cables. All preamps were carefully orientated along the z-direction to minimize Hall effect issues7. Fourth, the detuning traps on the drive were adjusted to be resonant at the Larmor frequency when the PIN diode was forward biased, as described in Part 3. 

Part 8: Geometrical Nearest-Neighbor Decoupling of the Array

The goal of the geometrical decoupling is to find a critical overlap of all adjacent coils in order to minimize the mutual decoupling between any pair of neighboring coils. To achieve this, we monitor the S12 interaction between two neighboring coils using the probe that plugs into the circuit board at the position of the preamplifier (see Fig. 8 and 9) as described in Part 6. This optimization of all next neighbors is the most time consuming adjustment procedure during phased array construction.

With a pair of coils hooked up to the network analyzer using the impedance transforming sockets, we measure S12 between each adjacent pair with all other coils detuned using their active detuning circuits. We minimized S12 by empirically bending the crossing bridges on the loops. After each bending, the impedance of the loop is often altered, and we need to readjust the test probe to keep the 50 Ohm impedance expected by the network analyzer. We also re-adjust the tuning of the loop (variable capacitor C3). After completing all combinations of elements, we repeated all these adjustments with a second go-thru.

Part 9: Tuning and Matching

After the geometrical decoupling, each element needs to be retuned and matched to the preamplifier’s noise matched condition Ohm. This was done using a direct S11 measurement with cables directly connected to the preamplifier sockets of the element under test (see Fig. 8 right photo). It is important that the electrical delay produced by the probe is accurately calibrated. When adjusting the tuning and matching all other unused elements of the phased-array are detuned. This procedure is done with the array loaded with a tightly fitted head-shaped phantom. 

Part 9: Final Bench Touch Up

The final step of phased-array construction is a careful check up of the loop tuning, active detuning, and preamplifier decoupling. This step was done with all of the preamplifiers present and the array plugged into a simulator that passes the PIN diode biases and DC power to each preamp from the scanner. Starting with all of the array elements in the detuned states (i.e. all PIN diode bias set to forward bias the diode in each element), the bias voltage for a given loop was toggled on and off while the detuning is monitored using a double inductive probe and the S12 measurement (see again Fig. 6). A fine adjusting of the detuning inductor is sometimes needed to achieve this. After this step, we verify the preamplifier decoupling by viewing the S12 versus frequency for the coil in the tuned state using the same decoupled double probe, but with reduced power output from the network analyzer (‑25dBm). It is important that the minimum of the S21 “dip” is at the Larmor frequency (Fig. 9). Note that this measurement requires that the preamplifier is powered on. Finally the cable traps on each plug were checked and adjusted via an S12 measurement using current probes (Fig. 10). 

Representative Results: 
The ratio of Qunload-to-Qload of a loop under test, surrounded by its six non-resonant neighboring elements, was Qunload / Qload = 231 / 44 = 5.25 (see Fig. 4) Thus, the noise is dominated by the sample for the constructed coil. The critical overlap ratio between the loop diameter and the center-to-center distance from two adjacent loops was found to be 0.76 Coupling between nearest neighbor elements ranged from ‑15 dB to ‑21 dB with an average of ‑17 dB (see Fig. 9.) This could be improved with an additional reduction of 23 dB via preamp decoupling (see Fig. 7). Furthermore, active PIN diode decoupling between tuned and detuned states resulted in 46 dB isolation (see Fig. 6). A 44 dB common mode current suppression could be achieved via a cable trap on each plug. The final array coil is shown in Fig. 11.

Tables and Figures: 

Figure 1: 3D model of the posterior and anterior coil segments. The geometrical structure of the whole coil element layout is engraved on the helmet model. The subject can easily lie down into the deep posterior segment. A smaller separate anterior coil section is then positioned over the forehead (orange part). Subject’s eyes and face are completely unobstructed to facilitate functional studies and provide a more comfortable environment for the young population. 

Figure 2: Posterior section of the 32-channel array. Left: 3D printed helmet with pentagon/hexagon layout and standoffs for preamplifier circuit boards. Right: Assembled array coil including wire coil elements, solder pads with fuses, capacitors, diodes and inductors. All preamplifier boards were carefully orientated in z-direction to minimize Hall effect issues. 

Figure 3: Circuit schematic for the coil element and preamp chain. The coil element comprises three capacitors: C3 to adjust tuning and C1 to provide impedance noise matched impedance. C2 and C1 have equal capacitance and provide a symmetrical circuit design. A detuning trap is formed around C1 using variable inductor L and a diode D. If the detuning trap fails, an RF fuse F is inserted in the circuitry. Bias T and preamplifier are located on a separate preamp circuit board. (C1=~15pF (trimmer), C1=C2=62pF, L=~24nH (trimmer), RFC=2.2μH). 

Figure 4: Measurement of the Ratio of Qunload to Qload: All the estimations of Q-ratios should be done for both; a single loop (left) and a loop under test surrounded by its six non-resonant neighboring elements (right). Usually, copper loading of adjacent elements reduces the unloaded Q. The coil quality factor is simply given by Q=ν0/(ν, where ν0 is the center frequency and (ν is ‑3dB bandwidth. It is important that the double probe location does not change during both of the above measurements. The graphic in the middle, shows Qunload (red S21 measure) and Qload (blue S21 measure) when the neighboring elements were present. (Qunload / Qload = 231 / 44 = 5.25). 

Figure 5: Pre-adjustment of the active detuning trap using a S11 measurement and a "sniffer" probe. The variable red inductor L and the match capacitor C1 forms a parallel resonance circuit when the PIN diode D is forward biased in transmit mode. This spoils the Q of the coil element by providing a high series impedance and consequently blocking current flow at the Larmor frequency. Final results of an active detuned element is shown in Fig. 6.

Figure 6: Measurement of active detuning. The S12 change was recorded for a given array element with and without forward bias current applied to the PIN diode trap. The difference of the two states determines the isolation between the receive array and the transmit coil. During the measurement of the tuning state (red trace), the preamplifier was replaced by a conjugate complex impedance load Z*NM (power matched). The constructed array shows an active detuning isolation of 40 dB.

Figure 7: Preamplifier decoupling was estimated as the change in S21 measurement in the case of power match (red trace), which has the conjugate complex impedance of the coil impedance under load without the presence of the preamp (left) and the case of noise match (blue trace), which has termination by low impedance preamplifier (right). The difference between the two states (in dB) determines the preamp decoupling (middle). For simple checks that the preamplifier decoupling is occurring at the proper frequency, the noise match trace (red trace) is not needed, and a simple check that the dip in the noise matched (blue) trace occurs at the Larmor frequency. Since the noise matched trace is obtained with the preamplifier present, it can be verified on a completed array (with all covers and cabling) by detuning all but the coil under test using a coil plug test rig.

Figure 8: Impedance matching of the coil elements. All coil elements are designed to provide a noise matched impedance for the preamplifier. However, network analyzers are calibrated to 50 Ohm. Thus, for direct S21 measurements a temporary 50 Ohm coil impedance is needed. 32 adjustable 50 Ohm probes were plugged into the preamplifier socket to transform the coil impedance to 50 Ohm (left, and blue impedance curve on smith chart). The final noise matched coil impedance was adjusted using a well calibrated cable via smith chart recording (right, red curve on smith chart).

Figure 9: Adjustment and recording of the geometrical decoupling between two adjacent coil elements. Decoupling measurements were done on a 50 Ohm basis using a direct S21 measurement with 50 Ohm transformers directly connected to the preamplifier sockets of the two elements under test. When measuring the decoupling between an adjacent pair, all other unused elements of the array were detuned. Minimized coupling was achieved by empirically bending the crossing bridges on the loops. Coupling between nearest neighbor elements ranged from ‑15 dB to ‑21 dB with an average of ‑17 dB.

Figure 10: Adjustment of a cable trap using current probes. All multi core cables for each plug comprising common mode cable traps to suppress common mode currents induced by the body coil. The cable is coiled up and bridged with capacitor on the outer shield. This forms a parallel LC circuit and presents a high impedance for common mode currents. Under S21 recording the trap LC circuit is adjusted to minimize current flow at the Larmor frequency.

Figure 11: Complete 32-channel brain array including a head shaped phantom and a mirror for visual stimulus. Five standard coil plugs were used to connect the constructed coil to the MR-scanner. 

Figure 12: Handcrafted RF probes used for coil construction: (a) double probes, (b) single magnetic field probes, (c) sniffer probe, (d) current probes.

Discussion: 
Thirty two channel brain arrays have become the standard type of adult head coil in clinical imaging. Higher element brain arrays are already used in research environment4. Pediatric neuroimaging8 is well suited to benefit from the SNR and acceleration capabilities potentially offered by highly parallel detection9,10. In this video we demonstrate a method to visualize the design, construction and characterization of a pediatric 32-channel receive-only coil for brain imaging on a 3 Tesla MRI scanner. We used a wire design to populate the whole array. This has the advantage that the array uses overall spatially sparser conductors, so that eddy currents and copper loading can be significant reduced in comparison to a circuit board array design3. In addition, bridges over the conductors of adjacent coils were formed by bending the wire rather than soldering an additional bridge, thus minimizing the number of solder joints. We also found that the ability to mechanically optimize the overlap between two loops by bending these wire bridges facilitated the element decoupling procedure.

Acknowledgment: 

We would like to thank Simon Sigalovsky for his great assistance and fruitful discussions. Thanks to Siemens Healthcare, Erlangen, Germany, for providing Tim Trio MRI scanner specific components. This project was supported by National Institutes of Health Grant (P41RR14075) and a grant from the Ellison Medical Foundation.

Disclosures: We have nothing to disclose
Materials:

Customized Bench Tools:

Handcrafted RF probes are commonly used in MRI coil construction. The most popular probe at the bench is the "double probe", which is used at almost every stage of array coil construction (Fig. 12a). The probe consists of two overlapped loops made out of semi-rigid coax cable with a gap in the shield placed symmetrically in the middle of each loop. The decoupling of both loops should be >70dB. Single magnetic field probes in different sizes are also useful for various applications (Fig. 12b). The small "sniffer" probe is useful to sniff out sources of resonances in circuits (e.g. isolated detuning traps). It consists of a small solenoid coil attached to the end of a coax cable (Fig. 12c). Current probes are used for adjustments of common mode current traps. They are made out of a ferrite core and a single turn semi-rigid coax cable (including shield gap) (Fig. 12d). For good accessibility in various measurement situations, it is helpful to have the possibility for opening up the current probe.

Commercial Materials:
	Item Description
	Model / Number
	Company

	Network Analyzer
	E5071C
	Agilent Technology Inc., Santa Clara, CA, USA

	3D Printer
	Dimension SST 1200es
	Dimension, Inc., Eden Prairie, MN, USA

	2D Circuit Router
	7000
	T-Tech, Inc, Norcross, GA, USA

	CAD Software
	Google-Sketchup7.1 
	Google, Mountain View, CA, USA

	Solder Station
	i-CON2 
	ERSA, Wertheim, Germany

	Non-Magnetic Chip Capacitors
	Series 11
	Voltronics, Corp., Danville, NJ, USA

	Non-Magnetic Variable Inductor
	165-02A06 
	Coilcraft, Inc., Cary, IL,USA

	Non-Magnetic Variable Capacitor
	BFC5 808 11339 
	Vishay Intertechnology, Inc., Malvern, USA

	PIN-Diodes
	MA4P4002B-402
	M/A-COM Technology Solutions Inc., Lowell, MA, USA

	Formable Coaxial Cable
	UT-85C-FROM,
	Micro-Coax, Pottstown, PA,USA


References:
1. 
Wiggins, G.C., Triantafyllou, C., Potthast, A., Reykowski, A., Nittka, M. & Wald, L.L. 32-channel 3 Tesla receive-only phased-array head coil with soccer-ball element geometry. Magn Reson Med. 56, 216-223 (2006).

2. 
Hergt, M., Oppelt, R., Vester, M., Reykowski, A., Huber, K., Jahns, K. & Fischer, H. Low noise preamplifier with integrated cable trap. In Proceedings of the 15th Annual Meeting of ISMRM, Berlin, Germany, p. 1037 (2007).

3. 
Kumar, A., Edelstein, W.A. & Bottomley, P.A. Noise figure limits for circular loop MR coils. Magn Reson. Med. 61, 1201-1209 (2009).

4. 
Wiggins, G.C., Polimeni, J.R., Potthast, A., Schmitt, M., Alagappan, V. & Wald, L.L. 96-Channel receive-only head coil for 3 Tesla: design optimization and evaluation. Magn Reson Med. 62, 754-762 (2009).

5. 
Roemer, P.B., Edelstein, W.A., Hayes, C.E., Souza, S.P. & Mueller, O.M. The NMR phased array. Magn Reson. Med. 16, 192-225 (1990).

6. 
Reykowski, A., Wright, S.M. & Porter, J.R. Design of matching networks for low noise preamplifiers. Magn Reson Med. 33, 848-852 (1995).

7. 
Possanzini, C. & Boutelje, M. Influence of magnetic field on preamplifiers using GaAs FET technology. In Proceedings of the 16th Annual Meeting of ISMRM, Toronto, Canada, p. 1123 (2008).

8. 
Raschle, N. M., Lee, M., Buechler, R., Christodoulou, J. A., Chang, M., Vakil, M., Stering, P. L. & Gaab, N. Making MR imaging child's play - pediatric neuroimaging protocol, guidelines and procedure. JoVE 29, (2009) http://www.jove.com/index/details.stp?id=1309.

9. 
Griswold, M.A., Jakob, P.M., Heidemann, R.M., Nittka, M., Jellus, V., Wang, J., Kiefer, B. & Haase, A. Generalized autocalibrating partially parallel acquisitions (GRAPPA). Magn Reson Med. 47, 1202-1210 (2002).

10. 
Pruessmann, K.P., Weiger, M., Scheidegger, M.B. & Boesiger, P. SENSE: sensitivity encoding for fast MRI. Magn Reson. Med. 42, 952-962 (1999).


 ADDIN 
