High resolution 3D Imaging of ex-vivo biological samples by micro CT

Amnon Sharir

Department of Molecular Genetics

Gregory Ramniceanu

Department of Biological Regulation

and

Vlad Brumfeld

Electron Microscopy Unit

Department of Chemical Infrastructure

The Weizmann Institute of Science, Rehovot, Israel

e-mail: vlad.brumfeld@weizmann.ac.il

Keywords: 3D imaging, tomography, x-ray, non invasive, ex-vivo

Abstract

Non-destructive volume visualization can be achieved only by tomographic techniques, of which the most efficient is the x-ray micro computerized tomography (µCT).

High resolution µCT is a very versatile yet accurate (1-2 microns of resolution) technique for 3D examination of ex-vivo biological samples1, 2. As opposed to electron tomography, the µCT allows the examination of up to 4 cm thick samples. This technique requires only few hours of measurement as compared to weeks in histology. In addition, µCT does not rely on 2D stereologic models, thus it may replace or complement histological methods3, 4, which are both time consuming and destructive.  Sample conditioning and positioning in µCT is straightforward and does not require high vacuum or low temperatures, which may adversely affect the structure. The sample is positioned and rotated 1800 or 3600between a microfocused x-ray source and a detector, which includes a scintillator and an accurate CCD camera, For each angle a 2D image is taken, and then the entire volume is reconstructed using one of the different available algorithms5-7. The 3D resolution increases with the decrease of the rotation step.  The present video protocol shows the main steps in preparation, immobilization and positioning of the sample followed by imaging at high resolution.

Materials

For image acquisition we have used a MICRO XCT-400 microfocussed X-ray tomographic system produced by Xradia, Concord, USA.

Images were processed and analyzed using ImageJ (NIH, USA), Avizo (VSG, France) and MicroView (General Electric, USA) software packages. Any available image analysis software can be used instead

Protocol

- Sample preparation

  - After extracting the tissue which is to be examined, mineralized tissues can be positioned in the instrument and imaged. To image the mouse femur one should follow the following steps:

    - Remove post mortem leg from a C57/Bl6 embryo 18.5 days postceutus (E18.5).

    - Seal the narrow end of a polystyrene pipette tip  (20-200 l) using epoxy resin or another glue, and fill the tip with the working buffer (PBS or other).

    - Fit tightly the leg into the tip and seal the other end with parafilm sheet.

    - Place the pipette tip into a suitable holder and follow the protocol from chapter 3. For visualizing the femur from the mouse leg embryo, the instrument is set 40 KV and 200 A.  For 8resolution 1000 projection images with 4x magnification have to be acquired.

  - Non-mineralized tissues have to be initially stained and fixed in order to increase the X-ray attenuation of the tissue of interest by using one of the many available protocols8,9. For the rat lungs and similar samples, the preparation protocol is:

    - Orthotopical implantation of non small cell lung carcinoma (NSCLC) NCI-H460 on  nude rat lungs

    - Lung cancer nodules start being detectable 4 weeks from implantation

    - Sacrifice the rat and immediately infuse it with saline solution mixed with heparin

    - Inject it with of a diluted solution of Microfil (Flowtech), (2 ml of compound solution, 3 ml of diluent and 0.3 ml of curing agent) into the left ventricle to stain the bronchial circulation

    - Extract the lungs and heart of the rat

    - Immobilize the sample (see chapter 2 of the protocol) by fitting it tightly into a 50 ml plastic test tube

    - Create a saturated ethanol atmosphere by placing on the bottom of the tube a cloth damped in ethanol

    - Glue or screw the tube in a holder of the instrument

    - Proceed with setting the imaging parameters (chapter 3).  For complete imaging of the rat lungs the source is set at 40KV and 100 A. In order to reach 16 resolution one has to acquire 2500 projection images at a magnification of 0.5x.

- Sample immobilization

At high resolution, it is important to avoid any change in the sample position during the measurement. For this, the sample is tightly fixed into a plastic recipient that fits its size. Polystyrene pipette tips, plastic Pasteur pipettes or specially built plastic holders are used in this respect. According to the experimental requirements, the sample can be examined in air or immersed in ethanol or buffer solutions. Typical immobilization and final positioning of the leg of the mouse embryo in the instrument is shown in Fig 1.

Figure 1 Final positioning of the embryonic mouse leg in the micro CT instrument

- Setting acquisition parameters: x-ray voltage and current, CCD exposure time

  - Placed into a holder, the sample is put in the rotation stage of the instrument

  - A first x-ray image is taken with the voltage and current arbitrarily set.

  - If the image is too dark, one should increase the number of photons first, so increase a little bit the current. If this is not enough, one should increase slightly the energy of the x-ray photons, i.e. the voltage on the x-ray tube.

  - If the image is too bright, one should first decrease the voltage, then the current.

  - The brightness of the image can be increased by binning. Binning of 1 takes into account the intensity of each pixel in the image, whereas binning of 2 takes the sum of each matrix of 2x2 pixels. The image will be about 4 times brighter than in the case of binning 1, but will have half the resolution.

  - After setting the optimum brightness, one has to optimize the exposure time of the camera to compromise between the best contrast on one side and a reasonable duration of the experiment on another side.

  - The contrast of images, especially of the low absorbing samples, can be improved by using filters, which reduce the photon flux, mainly that of the lower energy photons.

- Sample positioning

  - Choose the working magnification. Possible choices are 0.5x,4x,10x,20x and 40x. The field of view decreases with increasing magnification.

  - Get the optimal resolution and field of view by seting the distances between the x-ray source and the sample and between the sample and the detector. Increasing the source to sample distance decreases the field of view and increases the resolution. Sample to detector distance has the opposite effect.

The whole field to be viewed in 3D should be present in the projection image at all angles. One should check this by rotating the sample at different angles and by bringing the sample as close as possible to the rotation axis. For this, one should follow the following steps:

  - Take an image at 0 deg, and then rotate the sample at -20 deg. If the desired volume has shifted laterally, one should correct its position by repositioning the rotation axis.

  - After correction, the sample is to be rotated at another angle and the position corrected again, until the field of interest is inside the image at all angles from -90 to 90 deg.

- High resolution tomogram

  - During the measurement, the sample is rotated by a small angle at a time and at each angle one projection image is taken. The total number of images is always a compromise between the desired resolution on one side and the time of the measurement and the file size on another side. As shown in Fig.2, every single projection includes all slices in the sample superposed one over another, and therefore cannot reveal the 3D structure of the sample.

Figure 2 Projection images of the lung of rat at 00(A), 450 (B) and 900 (C) rotation angle

  - Only after taking projection images at least between -90 and 90 deg, one can proceed to the reconstruction of the sample volume. Reconstruction takes between 10 min and 2 hours, depending of the software used and of the number of projections. Again, the final quality of the 3D image is a compromise between the desired resolution and the time one wants to spend and the size of the resulting file.

- Image scale calibration

The pixel level (value) in a reconstructed image is unique for that image. In order to compare two different images, a unique intensity scale has to be imposed on each image. For this

  - Run a tomography with a standard phantom using the same experimental conditions as for the sample

  - Recalibrate the sample image using the values obtained for the phantom. The most common scale is the Hounsfield (or CT) scale. For 4x magnification the background value of 15,000 (for water or PBS) was replaced with 0 and the maximum value of 35,000 for the bone was replaced with the standard Hounsfield value of 3000. Other pixel values resulted from linear interpolation or extrapolation based on those limits.

- Image processing and analysis

After obtaining high resolution images, one has to extract relevant information by using image analysis software. The software package to be used has to be designed to work with very large files (up to 20Gb).

Representative Results

A representation of a femur from a C57/Bl6 mouse at embryonic day 18.5 (E18.5) - four days after the initiation of the mineralization process is shown in Fig 3.  The mineral layers are clearly visible (white), while soft tissues are not visible in this preparation. We took 1000 projection images with a linear magnification of 4x. The final resolution is 8 microns. A careful analysis of the volume rendering shown in Fig.1, shows that the bone volume fraction (the fraction of the bone volume that is occupied by mineralized tissue) is 0.18, and the bone mineral density is 723 mg/cm3.  Those values allow us to compare this structure with bones in other stages of development.

A.

B.

C.

Figure 3.   Different representations of a 3D image of a mouse femur embryo. The transversal (cross section) (A), the sagital (medio-lateral) section (B) and a snapshot of the volume rendering (C) are shown.

Figure 4 shows a 3D image of the lungs of a female nude rat (RNU), 12 weeks old, implanted orthotopically with non small cell lung carcinoma (NSCLC) NCI-H460. 2500 projection images were taken with a linear magnification of 0.5x, ensuring a final resolution of 16 microns.  The image shows the Microfil stained blood vessels (down to 20 micron diameter).The image analysis shows that 4 weeks after implantation, multiple cancer nodules are formed. They are covering a significant fraction of the lung volume (17%). Most of the pulmonary staining was found in the peripheral areas of the tumors. Significantly, as shown in the fig 2B, several blood vessels are present also inside the nodules, covering, according to preliminary analysis some 3% of their volume.

B.

A.

Figure 4. 3D image of cancer nodules growing in a rat lung. A snapshot of the volume rendering (A) and a section through the volume (B) are shown. The cancer nodules are marked with arrows.

As supplementary data, we show the volume rendering of the mouse femur in fig 1 (movie 1), as well as volume rendering of the rat lungs in Fig 2 (movie 2). We add also a movie including the serial sections through the lungs in which the nodules appear as grey areas in most of the slices (movie 3)

Discussion

C57/Bl6 mouse at embryonic day 18.5 (E18.5) is four days after the initiation of the mineralization process.  At this stage of development, the future bone is made of many layers of mineralized osteoids, clearly seen in Fig 3.

Results in Fig 4 show that in the orthotopic lung cancer animal models, human non small cell lung cancer can induce recruitment of blood vessels and neovascularization.  Unfortunately the presence of high leakages of contrast agent in the surroundings of the tumors prevented reliable quantification of the peripheral blood vessels. However, we could clearly conclude that at least for the bronchial circulation which was studied here, peripheral blood supply participates in tumor perfusion, with some perfusion present also inside the tumor.

References

- Schambach SJ, Bag S, Schilling L, Groden C, Brockmann MA. Application of micro-CT in small animal imaging. Methods. 50 2-13( 2010)

- Bauer JS, Link TM. Advances in osteoporosis imaging. Eur J Radiol.,71 440-449 (2009)

- Chappard D, Retailleau-Gaborit N, Legrand E, Basle MF, Audran M. Comparison insight bone measurements by histomorphometry and microCT. J Bone Miner Res. 20 1177–1184, (2005)

- Muller R, Van Campenhout H, Van Damme B, et al. Morphometric analysis of human bone biopsies: a quantitative structural comparison of histological sections and micro-computed tomography. Bone 23,59–66 (1998)

- K. Mueller, R. Yagel, and J.J. Wheller, "Anti-Aliased 3D Cone-Beam Reconstruction Of Low-Contrast Objects With Algebraic Methods," IEEE Transactions on Medical Imaging 18, 519-537 (1999).

- Kachelriess M, Schaller S, Kalender WA Advanced single-slice rebinning in cone-beam spiral CT, Med Phys. 27,754-772 (2000)

- Mori S. , Endo M., Komatsu S., Kandatsu S, YashiroT and Baba M  A combination-weighted Feldkamp-based reconstruction algorithm for cone-beam CT, Phys. Med. Biol. 51, 3953–3965 (2006)

- Marxen M, Thornton MM, Chiarot CB, Klement G, Koprivnikar J, Sled JG, Henkelman RM: MicroCT scanner performance and considerations for vascular specimen imaging. Med Phys 31,305-313 (2004).

- Plouraboué F, Cloetens P, Fonta C, Steyer A, Lauwers F, Marc-Vergnes JP: X-ray high-resolution vascular network imaging. J Microsc 215, 139-148 (2004)