Version:12/13/10
Transcranial Direct Current Stimulation Electrode Montage in Pain Research

Authors: 

Alexandre DaSilva1; Magdalena Sarah Volz2, 3; Marom Bikson4; Felipe Fregni2 

Authors institutions/affiliations:

(1) Headache & Orofacial Pain Effort (H.O.P.E.)TM, Biologic & Material Sciences, School of Dentistry, University of Michigan Ann Arbor, MI, USA
(2) Laboratory of Neuromodulation, Department of Physical Medicine & Rehabilitation, Spaulding Rehabilitation Hospital and Massachusetts General Hospital, Harvard Medical School, Boston, MA, USA
(3) Charité University Medicine Berlin, Berlin, Germany

(4) The City College of New York of CUNY, New York, NY, USA
Corresponding author:
Felipe Fregni, MD, PhD

Laboratory of Neuromodulation, Spaulding Rehabilitation Hospital, 125 Nashua Street #727, Boston, MA, USA 02114

Email: fregni.felipe@mgh.harvard.edu

Phone: 617–573-2195
Alexandre F. DaSilva, DDS, DMedSc

Headache & Orofacial Pain Effort (H.O.P.E.)TM, Biologic & Material Sciences, School of Dentistry, University of Michigan, 1011 N. University Ave., Room 1014A, Ann Arbor, MI, USA 48109-1078
Email: adasilva@umich.edu
Office: 734 615-3807 

Keywords: 

Transcranial direct current stimulation, pain, chronic pain, noninvasive brain stimulation, neuromodulation

Short Abstract: 
Transcranial direct current stimulation (tDCS) is an established technique to modulate cortical excitability1,2. It has been used as an investigative tool in neuroscience due to its effects on cortical plasticity, easy operation, flexible and safe profile. One area that tDCS has been showing encouraging results is for pain alleviation 
 ADDIN EN.CITE 
3-5
. Although tDCS is a relatively simple technique, one important factor to determine its success is the electrode montage. 

Long Abstract:
Transcranial direct current stimulation (tDCS) is a technique that has been intensively investigated in the past decade as this method offers a non-invasive and safe alternative to change cortical excitability2. The effects of one session of tDCS can last for several minutes, and its effects depend on polarity of stimulation, such as that cathodal stimulation induces a decrease in cortical excitability, and anodal stimulation induces an increase in cortical excitability that may last beyond the duration of stimulation6. These effects have been explored in cognitive neuroscience and also clinically in a variety of neuropsychiatric disorders – especially when applied over several consecutive sessions4. One area that has been attracting attention of neuroscientists and clinicians is the use of tDCS for modulation of pain-related neural networks3,5. Modulation of two main cortical areas in pain research has been explored: primary motor cortex and dorsolateral prefrontal cortex7. Due to the critical role of electrode montage, in this article, we show different alternatives for electrode placement for tDCS clinical trials on pain; discussing advantages and disadvantages of each method of stimulation. 

Protocol Text: 
0.) Materials

0.1.) Check if you have all materials needed (Table 1, Figure 1).

TDCS devices should be battery-driven and function as a constant current stimulator with a maximum output in the mA range. Constant voltage (voltage controlled) stimulators are not appropriate for tDCS.  Using electrical outlets to power the device is not appropriate as mal-functioning devices might deliver large intensities of electrical currents.
Electrodes used for tDCS generally consist of a metal or conductive-rubber electrodes enclosed in a perforated sponge pocket which is saturated with an electrolytes (fluid with salt). Prolonged passage of DC current across an electrode (where electrons from the stimulator are converted to ions carried through the body8) can produce undesired electrochemical products such as pH changes. The sponge pocket may act to physically separate, and thus buffer, the skin from electrochemical changes. For this reason, metal or rubber electrodes should never be placed on the skin during tDCS. Similarly during stimulation user should be vigilant against sponge dehydration and movement. A further related consideration is the durability and re-usability of tDCS electrodes. Our experience is that, especially when the polarity of electrodes is rotated, and proper stimulation conditions consistently maintained, rubber and metal electrodes can be re-used. The choice of electrolyte is discussed further below. From operating experience, it is recommended to use flat, and not too coarse perforated sponges, as they absorb electrolyte conduction solution the best8.

There is the possibility to apply topical anesthetics. Especially for short-lasting stimulation, when ramping is not possible, it might prevent somatosensory perception and uncomfortable sensation arising from tDC stimulation. Another reason for using topical application of local anesthetics is to create best comparability between sham and active tDCS conditions, since no subject would feel if current is flowing or not and optimal blinding situation would be guaranteed. This approach is particular vulnerable when planning to use larger intensities as blinding might be less effective in this situation7. Though sensation/pain and skin irritation are not always correlated, excessive use of topical anesthetics might mask serious adverse effects such as burning.

In this guide we illustrate the most typical tDCS set-up for pain management: using conductive rubber electrodes, pocket type perforated sponges, both placed on the head, with no topical aesthetic.
1.) Measurements
1.1.) Make sure the subject is seated comfortably. 
1.2.) The area of stimulation will be found through the measurement of the scalp. Usually the convention of the EEG 10/20 system is used7. The site of stimulation depends on your experimental approach.
1.3) Find localisation of the Vertex (Figure 2): 

Measure the distance of nasion to inion and mark halfway. 

Nasion – point between the forehead and nose, at the junction of the nasal bones (Figure 3). 
Inion – most prominent point of the occipital bone (Figure 3). 

Measure the distance between the pre-auricular points and mark halfway. Mark both halfway spots to find the Vertex.
1.4.a) Motor cortex (M1): Mark five cm lateral to the side of the Vertex to the direction of tragus (Figure 4) – this spot should correspond to C3/C4. Another possibility is to locate the primary motor cortex using transcranial magnetic stimulation (TMS) but this may not be necessary given the limited focality of tDCS.

1.4.b) Dorsolateral prefrontal cortex9 (DLPFC)9,10: Find the spot for motor cortex and measure five cm forward (Figure 5) – this spot should correspond to F3/F4.

2.) Position electrodes

2.1.) After finding the site of stimulation you should place one of the elastic head straps around the head circumference. The elastic head strap should be placed under the inion as to avoid movement during stimulation. The elastic straps should be made of non-conducting material (or they will function as electrodes) and non-absorbent material (to avoid the straps absorbing fluid from the sponges).
2.2.) Both sponges should be soaked with approximately 6 mL of saline solution (total of 12 mL per sponge). Be careful not to over soak the sponge (not excessively wet- there should be no water leaking; but also not dry as to have a good electrode contact). Avoid fluid leaking across the subject. You can use a syringe to add more solution if needed.
There is evidence that electrolyte solutions with lower NaCl concentrations(15 mM) are perceived as more comfortable during tDCS than those solutions with higher NaCl concentrations (220 mM)11. Since the ionic strength of deionised water is much less than that of all NaCl solutions, there is a significantly larger voltage required to carry current across the electrode and through the skin compared to NaCl solutions. Thus, it is recommend the use of solutions with relatively low NaCl concentration, in the range 15 mM to 140 mM, as tDCS at these concentrations is more likely to be perceived as comfortable, requires moderately  lower voltage while still allowing good conduction of current.11  The use of gels (adapted from applications such as EEG) has also been considered – a main limitation is the increased hassle of set-up clean-up following stimulation, without proven benefit regarding outcome when using perforated sponge electrodes.

2.3.) To optimize contact quality, move excessive hair from the site of stimulation to the extent practical and gently clean the surface of the skin to remove any signs of lotion, dirt, etc. and allow it to dry. Inspect the skin for any pre-exiting irritation, cuts, or lesions – avoid stimulating over damaged skin. Inspect the rubber insets and sponges for wear. If there is any evidence of deterioration, throw out the dirty components and use a new electrode.

2.4.) Insert the connector cord pin securely into the opening of the receptacle on the conductive rubber inset.
2.5.) Slide the conductive rubber inset into the sponge. The insulated portion of cable will protrude from the sponge-pocket opening. Ensure the entire conductive rubber inset is covered by the sponge and that there is no part of the connector cord pin is visible.
2.6.) The path from one device terminal, through one electrode, across the body, through the second electrode, and back to the second device terminal forms a circuit – the total resistance of which (sum of the electrodes and body resistance) can be measured.  If the overall resistance is abnormally high, this may indicate improper electrode set-up.  If your device measure resistance – this would be recommended – the indication field should display appropriate electrode contact. Ideally, one should aim to have the impedance under 5k Ohms. Some devices indicate the voltage across the path rather than resistance – in this case resistance can be calculated simply using ohms law (Resistance = Indicate Voltage / Current Applied ). Many devices continue to provide an indication of resistance during the course of stimulation, which provides a useful way to detect a potentially hazardous situation (such as a drying electrode). In some cases, device will automatically terminate stimulation or reduce stimulation intensity if resistance increases beyond a certain threshold.
2.7.) Connect the cables to the device. 

Ensure the connection polarity is correct as the effects of tDCS are highly polarity specific (typically, red indicates the Anode electrode, and black or blue indicates the Cathode electrode; this is the convention- but check with your device).  Note that in the context of tDCS (and more broadly electrical stimulation in general), “Anode” always indicated the relative positive terminal where positive current flows intro the body, while “Cathode” indicates the relative negative terminal where the positive current then exits the body.

2.8.) Place one sponge electrode below the elastic head strap (Table 2). Ensure that excessive fluid is not ejected from the sponge unto the scalp during this process as this will spread current flow across the scalp and deplete the sponge of fluid.
2.9.) Connect the second elastic head strap to the first elastic head strap according to the electrode montage you want to use (Table 2). Other elastic head straps can be used.  
2.10.) Place the second sponge electrode on the head under the second elastic head strap (Table 2). Make sure you place it onto the marked area you want to stimulate.
3.) Start tDCS
3.1.) The subject should be relaxed, comfortable and awake during the procedure.

Uncontrolled interference with current cortical activity during tDCS should be avoided. For stimulation to the motor cortex area, it has been demonstrated that intensive cognitive effort unrelated to the target area, as well as massive activation of the motor cortex by prolonged muscle contraction abolishes the effects of tDCS12. 
3.2) Adjust the settings on the tDCS stimulator to those you want to stimulate with (intensity, time and –if available- sham condition on/off).
3.3.) Turn on the tDCS device. To reduce/avoid adverse events start current flow with ramping up.  Many commercial devices include features to automatically ramp current on and off.

3.4) Some subject may experience heightened discomfort during the initial tDCS period – in such cases the current may be temporarily moderately decreased (e.g. by 50%) as the subject adjust and then gradually increased back up to the desired level.

At the beginning of stimulation, the majority of subjects will perceive a slight itching sensation, which then fades in most cases. Similarly, rapid changes of the stimulating circuit immediately might induce peripheral nerve firing. The subject can notice it as brief retinal phosphenes with electrodes near the eyes. These effects can be largely avoided by ramping the current up and down at the beginning and end of treatment. This might also prevent the dizziness or vertigo occasionally reported when current is suddenly increased or decreased.7
3.5) After the stimulation, current flow should be ramped off as well.
Note on High-Definition tDCS (HD-tDCS):  tDCS with electrodes smaller then approximately 2 cm2 is called HD-tDCS and often uses array of electrode (more than two) to guide current through the brain for specific applications13. This methods paper focused only on conventional tDCS (using larger sponge electrodes), and it is important to emphasize that HD-tDCS requires specific electrodes14, skin preparation, and stimulator hardware. It is not recommended to apply tDCS using 1-2 mA of current through small sponge electrodes
 ADDIN EN.CITE 
13,14
.
Representative Results: 

If you set up tDCS device properly, the device should display that either current is flowing during active tDCS situation, or the device should display sham mode when switched on for sham stimulation procedure. One important point here is that even though the device is showing that 1 or 2 mA is flowing through the system, it is possible that current is being shunted through the skin. In order to avoid this effect, it is recommended to have sufficient distance between electrodes (at least 8 cm) and avoid wetting of hair between electrodes. Furthermore, it is recommended to consult computer head models13 and neurophysiological studies. These additional steps would ensure that a specific montage is associated with significant changes in cortical excitability in the area that is being investigated. 

Figures and Tables:

Figure 1: Materials

Figure 2: Vertex Position. Cortical areas marked according to the 10/20 system. 

Figure 3: Nasion and Inion Position 

Figure 4: Motor cortex Position. Cortical areas marked according to the 10/20 system.
Figure 5: DLPFC Position. DLPFC= dorsolateral prefrontal cortex. Cortical areas marked according to the 10/20 system.

Figure 6: Change in cortical excitability due to current polarity and tDCS montage. Table: Induced effects of tDC stimulation on the size of the motor evoked potential (MEP), assessed by transcranial magnetic stimulation (TMS). MEP amplitudes after stimulation are given in percent of MEP without stimulation. Note that only the motor cortex (M1) - contralateral supra-orbital (Fp2) montage setting leads to a significant increase of MEP size after anodal and a decrease of MEP amplitude after cathodal stimulation. There are no significant effects on MEP amplitude in other tDCS montages. Figure: Electrode Placements6 (modified from Nitsche 2000).
Figure 7: Electrode Sizes

Figure 8: Decreasing the size of the electrode leads to a more focally effect of tDCS. Muscle-evoked potential (MEP) amplitude sizes of the abductor digiti minimi (ADM) and of the first dorsal interosseus muscle (FDI) during anodal or cathodal tDCS. Using the condition of a 35-cm2 electrode, anodal and cathodal tDCS influence the MEP amplitude size of the ADM and the FDI to a similar extent. At this montage, both hand muscle representation areas are located underneath the stimulating electrode. In the case of a smaller electrode, which is only placed over the representational area of the ADM, the effects of MEP amplitude changes of the cortical FDI representation are not reproducible (see yellow column)15 (modified from Nitsche 2007).
Figure 9: Tissue-depended current density. Current densities calculated in different tissues. Magnitude of current density is dependent on the conductivity of tissue. Note that approximately 10 % of current density reaches the Gray Matter16 (modified from Wagner 2007a).
Table 1: Materials
	Material
	

	
	TDCS device 

	
	Battery

	
	Two rubber head bands

	
	Two conductive rubber electrodes

	
	Two sponge electrodes (non-metallic)

	
	Cables

	
	NaCl solution

	
	Measurement Tape


Table 2: Electrode Positioning7


	Anode Electrode Positioning
	Cathode Electrode Positioning
	Observations
	Caveats

	Primary Motor cortex (M1)
	Supra-Orbital
	This is the most used montage. It has been proven that the cortical excitability can be changed up to 40%6 (Figure 6). Anodal stimulation results in neuronal depolarisation and increasing neuronal excitability while cathodal stimulation has opposite results6.
	Only one motor cortex is stimulated – might be a problem for bilateral pain syndromes. Also the confounding effect of the supra-orbital electrode needs to be considered.

	Primary Motor cortex (M1)
	Primary Motor cortex
	- Interesting approach when there is a bi-hemispheric imbalance between motor cortices (such as in stroke) 

- Can be used with two anodal stimulation electrodes (see sixth row), where cathodal electrode is placed in the supraorbital area for instance.
	Electrodes might be too close to each other- issue of shunting. 

A decrease of the area of the electrodes will increase the degree of shunting along the skin 16 
Therefore shunting might be related not only to electrode positioning but also to electrode size.
The relative resistance of the tissues is dependent upon the electrode position and size- the overall resistance on which the current flows is dependent upon the electrode properties16.

	Dorsolateral Prefrontal Cortex (DLPFC)
	Supra-Orbital
	Most used for DLPFC stimulation – positive results for treatment of depression17 and also chronic pain3.
	Only unilateral DLPFC stimulation situation is possible with this montage.

	Dorsolateral Prefrontal Cortex
	Dorsolateral Prefrontal Cortex
	- Interesting approach when there is a bi-hemispheric imbalance.

- Can be used for a two anodal stimulation situation (see sixth row), where cathodal electrode is placed in the supraorbital area for instance.
	Electrodes might be too close to each other- issue of shunting 16. (Please see second row, fourth column). 

	Occipital
	Vertex
	Interesting active control for chronic pain trials or for modulation of visual cortex.
	When used as active control, reference electrodes are placed in different locations- problem of comparability between intra- and inter- experimental approaches.

	Two anodal electrodes, e.g. both Motor cortices
	Supra-Orbital
	Simultaneous change in cortical excitability
	Transcallosal inhibition might add a confounding factor18

	One electrode over a cortical target, e.g. Primary Motor cortex (M1)
	Extra-Cranial
	Avoid the confounding effect of two electrodes with opposite polarities in the brain7.
	Depending on intended target, current distribution might not be optimal and therefore induce ineffective stimulation19


Discussion: 
Critical Steps:

Aspects to be checked prior to starting procedure:
· Check for any lesions on the scalp- additionally subjects should be specifically interviewed for the existence of skin diseases. If there is any present, tDCS procedure should be avoided, because it is reported that repeated daily tDCS causes clinically significant skin irritation under the electrodes in some patients7. So there is evidence of tDCS induced lesions according to skin integrity. For instance, it has been shown extensive redness and brown crusty intracutaneous changes with irregular round shapes due to tDC Stimulation at an intensity of 2 mA for a period of 2 weeks including five sessions each week20. Thus, the condition of the skin under the electrodes should be inspected before and after tDCS7.
· Check connectors of cables for electrolysis. Use another pair if evident. It is recommended to check the cables after approximately two months of use.
During both- active or sham- tDCS always ask whether subject still feels comfortable and is able to continue procedure.

Possible modifications:

· There are many varieties of electrode positioning7 (Table 2).
· There are many varieties of electrode sizes21 (Figure 7). For a given applied current, electrode size influences the current density15 and influences the focality of brain modulation (Figure 8). Clinical studies have suggest the smaller the size of the electrode the larger the current density21, however modelling studies suggest that the relationship between electrode size and area of brain modulation may be more complex22. Furthermore, the effects of small electrodes can differ qualitatively due to differential shunting of current in the scalp, and greater edge effect relative to the overall electrode area7. There were drastically superior levels of shunting for small electrode sizes reported than for the larger electrode schemes16. 

· High-Definition tDCS (HD-tDCS) is a technology that improves spatial focality but requires special hardware and procedural controls 14. 
· The electrode montage (electrode position and size) along with the applied current determine the generated electrical field strength in the brain which, in turn, determines the efficacy of tDCS. The use of just electrode current density, defined by the ratio between the current strength and the electrode size, has been proposed to normalize clinical outcomes – but modelling studies suggest this may only apply over a limited range and that overall electrode montage design determines outcome. Generally, increasing current intensity (or current density) for any given montage result in stronger effects. It is important to note that the current density at the skin surface is much higher than that in the brain16 (Figure 9).
· The position of the “return” (reference”) electrode may influence the overall current flow pattern through the brain, and thus even influence brain modulation under the active electrodes19. Thus the potion of both electrodes should be considered.

· The duration of the stimulation depends on the aim of the experimental approach. An increase of duration of the stimulation is associated with the occurrence and a longer duration of after-effects3,4.  However at least one study reported a reversal of effects directions when stimulation duration was increased, suggesting that more intensity does not necessarily translate into a more robust clinical outcome. Though tDCS within published parameters is considered safe and well tolerated, the potential for undesired side-effects increases with increasing intensity (time, duration, or repetition rate/number).
· Orientation of the electric field: defined by the electrodes positions and polarity. Cathodal stimulation typically decreases cortical excitability, whereas anodal stimulation typicallu increases the cortical excitability2,3. 

· Placebo: For sham-tDCS the same protocol above is used. However, the current will be applied for 30 seconds. This is one of the advantages of tDCS compared to other non-invasive brain stimulation methods. Since the arising sensations resulted from active-tDCS tend to occur only at early stages of application, this sham method makes it difficult for the patient to distinguish the placebo from active tDCS application. This initial and brief stimulation is a reliable method of placebo23.
Rationale for using tDCS in chronic pain:

The fact that multiple therapeutic pharmacological modalities provide only modest relief for chronic pain patients raises the possibility that the cause for the persistence of this debilitating disorder may lie within plastic changes in pain related neural networks. Interestingly, modulation of cortical activity can be achieved non-invasively by tDCS, as described earlier, which has been reported to produce lasting therapeutic effects in chronic pain due to changes in cortical plasticity.

Clinical effect of tDCS in chronic pain:

It has been shown that tDCS applied to the motor cortex changes the local cortical excitability (Figure 6)6. More precisely, anodal stimulation results in an increase of neuronal excitability, whereas cathodal stimulation has opposite results6. Indeed, anodal tDCS application over M1 leads to a greater improvement in visual analogue scale (VAS) pain ratings than sham tDCS. This therapeutic effect on pain following M1 stimulation, although transient, was reproduced in several groups of patients with neuropathic pain syndromes like trigeminal neuralgia, poststroke pain syndrome24, back pain and fibromyalgia25. Interestingly, clinical trials in neuropathic pain, due to spinal cord injury, stimulation of the motor cortex by tDCS showed pain improvement and cumulative analgesic effect that lasted two weeks after the stimulation. There is also evidence of its analgesic effect in fibromyalgia patients26 that is still significant after three weeks of follow-up for anodal tDCS of the M1 compared with sham stimulation, and as well as stimulation of the DLPFC26. Although the effects of anodal tDCS over DLFPC for pain improvement have not been explored extensively, it was shown it can be used to modulate pain thresholds in healthy subjects27. Nevertheless, stimulation of this brain area is a reliable technique for enhancing working memory 10, increasing performance on memory tasks in Alzheimer disease9 and reducing cue-provoked smoking craving significantly28 for instance; therefore it is also conceivable that this might be a useful strategy to modulate affective-emotional cognitive networks associated with pain processing in patients with chronic pain.
Acknowledgments:
DaSilva AF received funding support from CTSA high-tech funding grant, University of Michigan to complete this review. Volz MS is funded by a grant scholarship from Stiftung Charité.
Disclosures: The City University of New York holds patents on brain stimulation, on which Marom Bikson is an inventor. Marom Bikson is a co-founder of Soterix Medical Inc.
References:
1
Fregni F, Pascual-Leone A, Technology insight: noninvasive brain stimulation in neurology-perspectives on the therapeutic potential of rTMS and tDCS. Nat Clin Pract Neurol Jul;3 (7), 383 (2007b)).

2
Wagner T, Valero-Cabre A, Pascual-Leone A, Noninvasive human brain stimulation. Annu Rev Biomed Eng 9, 527 (2007b)).

3
Fregni, F, Freedman, S, Pascual-Leone, A, Recent advances in the treatment of chronic pain with non-invasive brain stimulation techniques. Lancet Neurol 6 (2), 188 (2007a).

4
Lefaucheur JP, Antal A, Ahdab R, Ciampi de Andrade D, Fregni F, Khedr EM, Nitsche M, Paulus W, The use of repetitive transcranial magnetic stimulation (rTMS) and transcranial direct current stimulation (tDCS) to relieve pain. Brain Stimul Oct;1 (4), 337 (2008).

5
Antal A, Paulus W, Transcranial magnetic and direct current stimulation in the therapy of pain. Schmerz Apr;24 (2), 161 (2010a)).

6
Nitsche MA, Paulus W, Excitability changes induced in the human motor cortex by weak transcranial direct current stimulation. J Physiol Sep 15;527 (Pt 3), 633 (2000).

7
Nitsche MA, Cohen LG, Wassermann EM, Priori A, Lang N, Antal A, et al, Transcranial direct current stimulation: state of the art. Brain Stimul 11, 642 (2008).

8
Merrill, D. R., Bikson, M., and Jefferys, J. G., Electrical stimulation of excitable tissue: design of efficacious and safe protocols. J Neurosci Methods 141 (2), 171 (2005).

9
Boggio PS, Khoury LP, Martins DC, Martins OE, de Macedo EC, Fregni F, Temporal cortex direct current stimulation enhances performance on a visual recognition memory task in Alzheimer disease. J Neurol Neurosurg Psychiatry Apr;80 (4), 444 (2009).

10
Fregni, F, Boggio, PS, Nitsche, M, Bermpohl, F, Antal, A, Feredoes, E, Marcolin, MA, Rigonatti, SP, Silva, MT, Paulus, W, Pascual-Leone, A, Anodal transcranial direct current stimulation of prefrontal cortex enhances working memory. Exp Brain Res 166 (1), 23 (2005a).

11
Dundas JE, Thickbroom GW, Mastaglia FL., Perception of comfort during transcranial DC stimulation: effect of NaCl solution concentration applied to sponge electrodes. Clin Neurophysiol May;118 (5), 1166 (2007).

12
Antal A, Terney D, Poreisz C, Paulus W, Towards unravelling task-related modulations of neuroplastic changes induced in the human motor cortex. Eur J Neurosci Nov;26 (9), 2687 (2007).

13
Datta, A. et al., Gyri-precise head model of transcranial direct current stimulation: improved spatial focality using a ring electrode versus conventional rectangular pad. Brain Stimul 2 (4), 201 (2009).

14
Minhas P, Bansal V, Patel J, Ho JS, Diaz J, Datta A, Bikson M., Electrodes for high-definition transcutaneous DC stimulation for applications in drug delivery and electrotherapy, including tDCS. J Neurosci Methods Jul 15 (190(2)) (2010).

15
Nitsche MA, Doemkes S, Karaköse T, Antal A, Liebetanz D, Lang N, Tergau F, Paulus W, Shaping the effects of transcranial direct current stimulation of the human motor cortex. J Neurophysiol. Apr;97 (4), 3109 (2007).

16
Wagner T, Fregni F, Fecteau S, Grodzinsky A, Zahn M, Pascual-Leone A, Transcranial direct current stimulation: a computer-based human model study. Neuroimage Apr 15 (35(3)), 1113 (2007a)).

17
Boggio PS, Rigonatti SP, Ribeiro RB, Myczkowski ML, Nitsche MA, Pascual-Leone A, Fregni F., A randomized, double-blind clinical trial on the efficacy of cortical direct current stimulation for the treatment of major depression. Int J Neuropsychopharmacol Mar;11 (2), 249 (2008a).

18
Williams JA, Pascual-Leone A, Fregni F, Interhemispheric modulation induced by cortical stimulation and motor training. Phys Ther Mar;90 (3), 398 (2010); Lang, N, Nitsche, MA, Paulus, W, Rothwel,l JC, Lemon, RN, Effects of transcranial direct current stimulation over the human motor cortex on corticospinal and transcallosal excitability. Exp Brain Res Jun;156 (4), 439 (2004).

19
Bikson, M., Datta, A., Rahman, A., and Scaturro, J., Electrode montages for tDCS and weak transcranial electrical stimulation: role of "return" electrode's position and size. Clin Neurophysiol 121 (12), 1976.

20
Palm, U, Keeser D, Schiller C, Fintescu Z, Nitsche M, Reisinger E, Padberg Skin lesions after treatment with transcranial direct current stimulation (tDCS). Brain Stimul Oct;1 (4), 386 (2008).

21
Datta A, Elwassif M, Bikson M, Bio-heat transfer model of transcranial DC stimulation: comparison of conventional pad versus ring electrode. Conf Proc IEEE Eng Med Biol Soc 2009, 670 (2009).

22
Miranda, P. C., Faria, P., and Hallett, M., What does the ratio of injected current to electrode area tell us about current density in the brain during tDCS? Clin Neurophysiol 120 (6), 1183 (2009).

23
Gandiga PC, Hummel FC, Cohen LG, Transcranial DC stimulation (tDCS): a tool for double-blind sham-controlled clinical studies in brain stimulation. Clin Neurophysiol Apr;117 (4), 845 (2006).

24
Fregni F, Boggio PS, Mansur CG, Wagner T, Ferreira MJ, Lima MC, Rigonatti SP, Marcolin MA, Freedman SD, Nitsche MA, Pascual-Leone A, Transcranial direct current stimulation of the unaffected hemisphere in stroke patients. Neuroreport 16 (14), 1551 (2005b).

25
Antal A, Terney D, Kühnl S, Paulus W, Anodal transcranial direct current stimulation of the motor cortex ameliorates chronic pain and reduces short intracortical inhibition. J Pain Symptom Manage May;39 (5) (2010 b).

26
Fregni F, Gimenes R, Valle AC, Ferreira MJ, Rocha RR, Natalle L, Bravo R, Rigonatti SP, Freedman SD, Nitsche MA, Pascual-Leone A, Boggio PS, A randomized, sham-controlled, proof of principle study of transcranial direct current stimulation for the treatment of pain in fibromyalgia. Arthritis Rheum Dec;54 (12) (2006a).

27
Boggio PS, Zaghi S, Lopes M, Fregni F, Modulatory effects of anodal transcranial direct current stimulation on perception and pain thresholds in healthy volunteers. Eur J Neurol Oct;15 (10), 1124 (2008b).

28
Fregni F, Liguori P, Fecteau S, Nitsche MA, Pascual-Leone A, Boggio PS, Cortical stimulation of the prefrontal cortex with transcranial direct current stimulation reduces cue-provoked smoking craving: a randomized, sham-controlled study. J Clin Psychiatry Jan;69 (1), 32 (2006c).



