Multifocal Electroretinograms

Author:

Donnell J. Creel

John A. Moran Eye Center

University of Utah

donnell.creel@hsc.utah.edu
Keywords: 

Multifocal electroretinograms, mfERGs

Short Abstract:

An important development in electroretinograms is the multifocal electroretinogram (mfERGs).  Multifocal electroretinograms are a mathematical average of an approximation of a b-wave.  Software programs can derive ERGs from more than a hundred retinal areas in a few minutes per eye.  Scotomas and retinal dysfunction can be mapped and quantified. 

Long Abstract: 

A limitation of the traditional full-field ERG is that unless about 20% or more of the retina is affected the electroretinograms (ERGs) are often normal.  A person that can be legally blind with macular degeneration or other scotomas will often have normal full-field ERGs.  An important development in ERGs is the multifocal ERG (mfERG).  Erich Sutter adapted the mathematical sequences called binary m-sequences enabling the isolation from a single electrical signal an electroretinogram representing less than each square millimeter of retina.

Multifocal electroretinograms look like a flash electroretinogram but are a mathematical average of an approximation of a b-wave.  Multifocal ERG programs can derive ERGs from more than a hundred retinal areas in a few minutes per eye.  Scotomas and retinal dysfunction can be mapped and quantified.

The methodology to record mfERGs using bipolar speculum contact lens is depicted.
Protocol Text: 

1.) Select a room with little electrical interference.  A sink is convenient to wash your hands and to clean electrodes.  An internet connection is also preferable to share data and for manufacturer of your system to update software or help solve problems.
1.1) Organize your dilating drops, balanced salt solution, topical anesthetic and cleaning products so that they are readily accessible.

1.2) Be compulsive about conditions when testing subjects.  Place subjects in the same position.  Use the same room lighting for every subject.

1.3) After introducing yourself and the procedure dilate the subject’s eyes.  Only mydriatic dilation is necessary.  Once eyes dilate to at least 7 mm testing can begin.

1.4) There are several types of recording electrodes.  Popular electrodes include the ERG Jet, DTL silver wire, gold foil and speculum contact lens.  If using reusable electrodes follow the manufacturer’s recommendations for sterilization. 

1.5) The convention is for recording electrode touching cornea or sclera to be positive pole, which produces signal with b-wave up.  If using a monopolar recording electrode choose a location on head for the negative pole such as forehead, mastoid or earlobe.  Ground electrode may be placed any location on body.

1.6) Clean the skin electrode locations well with alcohol and/or a skin preparation product.

1.7) If the recording electrode contacts the eye numb the eye with several drops of a topical anesthetic. If using contact lens or speculum contacts be very careful not to scratch the cornea.

1.8) If your mfERG system displays stimulus on video monitor place all subjects same distance from stimulus monitor.  Make testing situation the same for every subject.
1.9) Pay attention that the subject is comfortable and emphasize importance of maintaining fixation and being relaxed.  Choose fixation target that subject can maintain fixation.  Verbally coach patients during testing to maintain fixation and stay relaxed.
1.10) Multifocal electroretinograms are useful to diagnose and quantify progress of most retinal diseases, ocular trauma, and retinal drug toxicity.
1.11 )  Back up your data at the end of each recording session in addition to saving in system’s computer memory.
Representative Results: An example of patients with early macular degeneration, which show reduced ERG amplitudes in central macular area (Figure 1).

Plaquenil drug toxicity can be detected prior to patient noticing loss.  Plaquenil and chloropromazine toxicity usually begin with small areas of a few square millimeters affected in region 5-15 degrees from fovea (Figure 2).

Stargardt’s disease and cone dystrophies show conspicuous reduction in mfERGs in central macular where cones dominate.

Retinal trauma producing field defects produce abnormal mfERGs consistent with visual field loss. 
Birdshot chorioretinopathy is an uncommon retinal disorder that is peculiar in that the first electrophysiological anomaly is slowing of cone physiology (Figure 3).  The progression of this slowing can be followed as the disease progresses from nasal retina and spreads across the retina (Figure 4).
Discussion: The multifocal electroretinogram is not just one analysis representing the mathematical approximation of b-wave amplitudes.  Look at implicit times.  Retinal diseases vary in whether the implicit times are affected and at what stage in the disease process implicit times are abnormal.  There are specialty protocols provided by most manufacturers of mfERG recording systems.  Be flexible in the recording and analysis protocols you choose.

I encourage those recording multifocal electroretinograms to broaden their horizons and apply new protocols and analyses appropriate for the clinical condition. Multifocal electroretinograms comprise a system of many possible protocols and analyses. 
Figure 1. Multifocal ERGs in early macular degeneration.

(Use Figure labeled AMD.jpg)

Figure 2.  Color scale transformation mfERGs “amplitudes” in patient showing early pockets of retinal toxicity 5-15 degrees from fovea.

(Use photo labeled Plaquenil1.jpg)

Figure 3.  Fundus photo of nasal retina left eye of patient with birdshot chorioretinopathy.

(Use photo labeled birdshotfundus.jpg

Figure 4. Color transformation of “b-wave” implicit times of left eye of patient with birdshot chorioretinopathy showing slow times.

(Use photo labeled birdshot.jpg)
References
Birch, D. G, Williams P. D., Callanan D., Wang R., Locke K. G., & Hood D. C.. Macular atrophy in birdshot retinochoroidopathy: an optical coherence tomography and multifocal electroretinography analysis.  Retina 30, 930-7 (2010).

Creel, D. J. The electroretinogram: Clinical applications. http://webvision.med.utah.edu/clinicalerg.html (2011).

Feigl, B., Brown, B., Lovie-Kitchin, J. & 

HYPERLINK "http://eprints.qut.edu.au/view/person/Swann,_Peter.html"
Swann, P. Adaptation responses in early age-related maculopathy Adaptation responses in early age-related maculopathy. Investigative Ophthalmology and Visual Science, 46, 4722-4727 (2005).
Hansen, R. M., Moskowitz, A. & Fulton, A. B. Multifocal ERG responses in infants. Invest Ophthalmol Vis Sci. 50, 470-475 (2009)
Lai, T. Y., Chan, W. M., Li, H., Lai, R. Y. & Lam, D. S. Multifocal electroretinographic changes in patients receiving hydroxychloroquinetherapy. Am J Ophthalmol. 140, 794-807 (2005).
Marmor, M. F., Hood, D., Keating, D., Kondo, M., Seeliger, M. W. & Miyake, Y. Guidelines for basic multifocal electroretinography (mfERG). Doc Ophthalmol 106,105-115 (2003).

Sutter, E. E. Noninvasive Testing Methods: Multifocal Electrophysiology. In: Darlene A. Dartt, editor. Encyclopedia of the Eye, Vol 3. Oxford: Academic Press pp. 142-160 (2010).

Michaelides, M., Stover, N. B.,  Francis, P. J. & Weleber, R. G. Retinal toxicity associated with hydroxychloroquine and chloroquine: risk factors, screening, and progression despite cessation of therapy. Arch Ophthalmol 129, 30-39 (2011).

