Synthetic, multi-layer, self-oscillating vocal fold model fabrication

Authors:

Preston R. Murray1, Scott L. Thomson2

Department of Mechanical Engineering

Brigham Young University

1prestomurray@gmail.com

2thomson@byu.edu

Corresponding author:

Scott Thomson

Keywords:

Vocal folds, larynx, voice, speech, artificial biomechanical models

Short Abstract:

The methodology for fabricating synthetic vocal fold models is described. The models are life-sized, mimic the multi-layer structure of the human vocal folds, self-oscillate at pressures comparable to lung pressure, and demonstrate flow-induced vibratory responses that are similar to those of human vocal folds.

Long Abstract:

Sound for human speech is produced via flow-induced vocal fold vibration. The vocal folds consist of several layers of tissue, each with differing material properties. Normal voice production relies on healthy tissue and vocal folds, and occurs as a result of complex coupling between aerodynamic, structural dynamic, and acoustic physical phenomena. Voice disorders affect up to 7.5 million annually in the United States alone1 and often result in significant financial, social, and other quality-of-life difficulties for affected individuals. Understanding the physics of voice production has the potential to significantly benefit voice care, including clinical prevention, diagnosis, and treatment of voice disorders.

Existing methods for studying voice production include in vivo experimentation using human and animal subjects, in vitro experimentation using excised larynges and synthetic models, and computational modeling. Owing to hazardous and difficult instrument access, in vivo experiments are severely limited in scope. Excised larynx experiments have the benefit of anatomical and some physiological realism, but parametric studies involving geometric and material property variables are limited. Further, they are typically only able to be vibrated for relatively short periods of time (typically on the order of minutes, not hours).

Overcoming some of the limitations of excised larynx experiments, synthetic models are emerging as a complementary tool for studying voice production. In this paper, fabrication of a model with multiple layers of silicone of differing material properties is discussed. The model layers simulate the multi-layer structure of the human vocal folds, including epithelium, superficial lamina propria, intermediate and deep lamina propria (i.e., ligament; a fiber is included for anterior-posterior stiffness), and muscle (i.e., body) layers2. The model motion features several improvements over that of prior synthetic models, including reduced inferior-superior motion, onset pressure closer to human onset pressure, and evidence of a “mucosal wave,” or a vertically-traveling wave that is characteristic of healthy human vocal fold vibration.

Synthetic models such as that presented here can be fabricated with systematic changes to geometry and material properties, allowing for the study of healthy and unhealthy human phonatory aerodynamics, structural dynamics, and acoustics. For example, they have been used to study left-right vocal fold asymmetry3,4, clinical instrument development5, laryngeal aerodynamics6-9, vocal fold contact pressure10, and subglottal acoustics11 (a more comprehensive list can be found in Kniesburges et al.12).

Protocol Text:

The fabrication sequence (see Fig. 1) consists of making molds for vocal fold model layers, sequentially casting the layers, and mounting the models for testing. The model has four distinct layers: body, ligament, superficial lamina propria, and epithelium, in addition to a single fiber. A backing layer is added for mounting purposes. The cross section is shown in Fig. 2.

- Mold fabrication and preparation

  - Create solid models of three vocal fold layers: superficial lamina propria, ligament, and body layers. This is typically done by creating 3D computer-aided design (CAD) models with the desired geometries, exporting the CAD models as stereolithography (STL) files, and sending the STL files to a custom machine shop for rapid prototyping.

  - Create a box-shaped mold form using thin pieces of acrylic material. Approximate dimensions are 2.54 cm high × 5.72 cm wide × 6.35 cm deep. Make the bottom of the form by adhering it to a flat acrylic plate. Seal all interior edges with vacuum grease.

  - Place small amount of vacuum grease on lateral side of the solid model of desired geometry (i.e., the body, the ligament, or the superficial lamina propria). Press into bottom of mold form cavity, vacuum grease side down, so that the vacuum grease holds the part in place. Liberally coat mold form and solid model with release agent. Using a paint brush, ensure release agent reaches into all corners of the mold form cavity.

  - Mix 10 parts A and one part B of platinum silicone rubber (parts measured by weight) and pour into mold form cavity. Place mold form with uncured silicone rubber in a vacuum chamber and reduce pressure (e.g., to around 23 inHg below atmospheric pressure) for approximately 3 minutes to remove air bubbles. Remove from vacuum chamber and place on level surface. Allow to cure for 24 hours and remove mold from mold form.

  - Repeat steps 1.1 through 1.4 to create molds for each of the superficial lamina propria, ligament, and body layers.

  - Cut ligament layer mold at the center of the medial surface in the anterior-posterior direction with a straight razor to allow for fiber insertion.

- Casting of each layer

    - Body Layer: Apply thin layer of release agent to body mold cavity with paint brush. Mix one part B, one part A of Ecoflex® Supersoft Platinum Silicone (by weight). Add one part silicone thinner (by weight) to reduce the eventual cured stiffness of the material. Mix together for 30 seconds and place in vacuum chamber for one minute to remove entrapped air. Remove mixture from vacuum and pour into the body mold cavity, but do not fill to the top of the entire mold cavity. Place in oven at 250 °F for 30 minutes. Remove from oven and cool.

    - Backing: Mix one part B and one part A of Dragon Skin® and add one part silicone thinner (by weight). Mix vigorously for 30 seconds, place in vacuum for 1 minute, and pour into the body mold cavity until full. Place in oven at 250 °F for 30 minutes. Remove mold from oven and cool. Remove model from mold, allow to cool to room temperature, and remove any release agent on the surface of the body layer with paper towel.

    - Ligament Layer: Apply thin layer of release agent on ligament mold cavity surface with paint brush. Place a 30 cm thread in the mold by pushing it into the cut from the straight razor. Thoroughly mix one part B and one part A of Ecoflex® and four parts of silicone thinner (by weight). Place in vacuum chamber to remove air bubbles and pour mixture into ligament mold cavity.

    - Ligament Layer (continued): Press body-backing model (from Steps 2.1.1 and 2.1.2) into the ligament mold cavity. Begin insertion at one side and gently move to the other so that the model pushes the excess uncured silicone and air bubbles out of the mold cavity. If air bubbles are present, remove the model from the mold cavity, refill with uncured silicone, and repeat pressing model into the mold. Place mold in oven for 30 minutes, remove, and cool to room temperature. Remove model from mold. Remove excess release agent with paper towel.

    - Cover Layer: Apply thin layer of release agent on cover mold cavity surface with paint brush. Mix one part B, one part A of Ecoflex®, and 8 parts silicone thinner by weight. Vacuum as previous and pour into cover mold cavity. Use the same process described in Step 2.1.4 to insert the ligament-body-backing model into the cover mold cavity. Place in an oven at 250 °F and cure for one hour. Remove from oven and allow to cool. Remove model slowly and with extreme care so that the cover remains intact.

    - Epithelium Layer: Place vocal fold model on a flat surface with the backing down. Remove support material with a straight razor. Suspend threads in air by attaching them to an object of greater height than the model. Mix one part B and one part A of Dragon Skin® with one part of silicone thinner, mix, vacuum, then pour over the model and allow to cure for one hour. Repeat the process to have two layers. Remove excess material with a straight razor.

    - Optional: If each layer is desired to be a different color (for visible inspection of different layers), add dye to part B of either the Ecoflex® or Dragon Skin® during the mixing process.

    - Optional: If material property data will be collected, create tensile and rheological specimens simultaneously with fabrication of each model layers. Do this by pouring extra uncured material into release agent-treated molds of desired material property specimen shape and size.

    - Optional: If measurements of layer thickness are desired, cut a cross section of the model with a straight razor and inspect with microscope.

- Final model preparation for testing

  - Mount each completed vocal fold model into an acrylic mounting plate by first applying a thin layer of silicone glue on the back (lateral) and side (anterior-posterior) model surfaces. Insert model into recessed cut of mounting plate. Align the model medial surface with the top of the acrylic plate. Wipe away excess glue. Allow glue to cure for one hour.

  - Apply talc powder to the model surface to reduce surface tackiness.

  - For medial surface tracking use a fine-point Sharpie pen to mark dots on the model. Best results occur if marking is done after application of talc powder.

  - Place long bolts through holes of the mounting plate with the threaded ends pointing toward the model to which the existing model will be paired. Lay threads over the bolts. Put closed cell foam over the bolts to close any air gaps. Ensure that the threads are compressed over the bolts.

  - Pair prepared model with another vocal fold model previously mounted to an acrylic holder. Tighten screws to compress the foam and bring the medial surfaces together until the desire pre-vibratory gap is reached. Ensure both sets of threads are placed over the bolts and extend outward from the acrylic plates in the anterior-posterior direction.

  - Mount vocal fold pair on air supply tube.

  - Tie the anterior threads together to form a loop. Repeat for the posterior threads. Hang desired weight on the loops simultaneously.

  - Models are now ready for testing and data collection.

Representative Results:

Multi-layer synthetic models as described herein can be expected to vibrate with an onset pressure around 300 to 600 Pa, a vibration frequency around 90 to 120 Hz, and a flow rate around 250 to 600 ml/s; these variables will depend on factors such as pre-vibratory gap, individual layer stiffness, fiber tension, and geometry. High-speed imaging can be used to capture the model vibratory response. For example, high-speed videokymography can be used to explore evidence of phase difference between inferior and superior glottal margins (Fig. 3). High-speed medial surface stereoscopic imaging can yield evidence of vibratory features such as convergent-divergent motion, reduced inferior-superior motion, and mucosal wave-like motion (Fig. 4).

Tables and Figures:

Figure 1: Synthetic vocal fold model fabrication process. First, molds for each layer are created. Each layer is then cast, beginning with the body layer and ending with the epithelium layer. Finally, models are mounted in acrylic plates for testing.

Figure 2: Synthetic vocal fold model cross section. Distinct body, cover, ligament, and epithelium layers are shown.

Figure 3: High-speed kymogram of model vibration. Phase differences between the inferior and superior margins are evident.

Figure 4: Medial surface profile of synthetic vocal fold model in a hemilarynx arrangement, captured at two different instances of time while vibrating. Ink markers were placed on the medial surface (as shown in the right image), imaged using two high-speed cameras, and tracked over the vibratory cycle.  The left plot shows a convergent glottis during the opening phase, and the right plot shows a divergent glottis during the closing phase.

Discussion:

Demonstrated is a method for fabricating synthetic vocal fold models that exhibit vibratory behavior similar to that of human vocal folds. The multi-layer concept provides significant advantages over previous one- and two-layer models, primarily in terms of improved model motion. The method presented here is demonstrated on a somewhat idealized model in terms of geometry, but it can be applied to models with vastly different geometries. A model based on human imaging (e.g., MRI13, CT) geometric data can thus be studied.

The materials used here have linear stress-strain response characteristics, and an anticipated future modification to this model includes the incorporation of nonlinear stress-strain materials. Use of biological rather than synthetic materials in this fabrication is also possible. Because of the extreme flexibility of the lamina propria layer, the model is less robust under vibration than previous one- and two-layer models. However, occasionally applying talc powder to minimize surface adhesion and avoiding excessive operating pressures should allow for the model to be used for durations that typically far exceed those possible using excised larynges.

Acknowledgments:

The authors gratefully acknowledge Grants R03DC8200, R01DC9616, and R01DC5788 from the National Institute on Deafness and Other Communication Disorders for support of synthetic model development.

Disclosures:

The authors have nothing to disclose.

Table of specific reagents and equipment:

| Name of the reagent | Company | Catalogue number | Comments (optional) |
| --- | --- | --- | --- |
| High Vacuum Grease | Dow Corning | 01018817 |  |
| Pol-Ease 2300 | Polytek | Pol-Ease2300-1 | Release agent |
| Smooth-Sil 950 | Smooth-On | Smooth-Sil 950 | Mold making material |
| Vacuum Pump | Edwards | E2M2 |  |
| Vacuum Chamber | Kartell | 230 |  |
| Pressure Gage | Marsh Bellofram | 11308252A |  |
| Straight Razor | Husky | 008-045-HKY |  |
| Ecoflex 00-30 | Smooth-On | Ecoflex 00-30 |  |
| Silicone Thinner | Smooth-On | Silicone Thinner |  |
| Dragon Skin | Smooth-On | Dragon Skin 10 FAST |  |
| Thread | Omega | OmegaCrys | Use only clear fibers |
| Silicone Dye | Smooth-On | Silc Pig Black |  |
| Silicone Glue | Smooth-On | Sil-Poxy |  |
| Talc Powder | Western Family |  |  |

References:

1	National Institute on Deafness and Other Communicative Disorders http://www.nidcd.nih.gov/health/statistics/vsl.asp (2010).

2	Hirano, M., & Kakita, Y. Cover-body theory of vocal fold vibration. In Speech Science: Recent Advances, p. 1-46 (1985).

3	Pickup, B.A., & Thomson, S.L. Influence of asymmetric stiffness on the structural and aerodynamic response of synthetic vocal fold models. Journal of Biomechanics, 42 (14), 2219-2225 (2009).

4	Zhang, Z. Vibration in a self-oscillating vocal fold model with left-right asymmetry in body-layer stiffness. Journal of the Acoustical Society of America. 128 (5), EL279-EL285 (2010).

5	Popolo, P. S., & Titze, I. R. Qualification of a Quantitative Laryngeal Imaging System Using Videostroboscopy and Videokymography. Annals of Otology, Rhinology & Laryngology. 117 (6), 4014-412 (2008).

6	Thomson, S.L., Mongeau, L., & Frankel, S.H. Aerodynamic transfer of energy to the vocal folds. Journal of the Acoustical Society of America. 118 (3), 1689-1700 (2005).

7	Neubauer, J., Zhang, Z., Miraghaio, R., & Berry, D.A. Coherent structures of the near field flow in a self-oscillating physical model of the vocal folds. Journal of the Acoustical Society of America. 121 (2) 1102-1118 (2007).

8	Drechsel, J.S., & Thomson, S.L. Influence of supraglottal structures on the glottal jet exiting a two-layer synthetic, self-oscillating vocal fold model. Journal of the Acoustical Society of America. 123 (6), 4434-4445 (2008).

9	Becker, S., et al. Flow-structure-acoustic interaction in a human voice model. Journal of the Acoustical Society of America. 125 (3), 1351-1361 (2009).

10	Spencer, M., Siegmund, T., & Mongeau, L. Experimental study of the self-oscillation of a model larynx by digital image correlation. Journal of the Acoustical Society of America 123 (2) 1089-1103 (2007).

11	Zhang, Z., Neubauer, J., & Berry, D. The influence of subglottal acoustics on laboratory models of phonation.  Journal of the Acoustical Society of America. 120 (3) 1558-1569 (2006).

12	Kniesburges, S., et al. In vitro experimental investigation of voice production. Current Bioinformatics. (In Press).

13	Pickup, B.A, & Thomson, S.L. Flow-induced vibratory response of idealized vs. magnetic resonance imaging-based synthetic vocal fold models. Journal of the Acoustical Society of America 128 (3), EL124-EL129 (2010).