A Targeted Mass Spectrometry Platform for Cancer Cell Metabolomics

Min Yuan1, Susanne Breitkopf1 and John M. Asara1, 2 *

- Div. of Signal Transduction/Mass Spectrometry Core, Beth Israel Deaconess Medical Center, Boston, MA 02115

- Department of Medicine, Harvard Medical School, Boston, MA 02115

* Corresponding author:  John M. Asara, jasara@bidmc.harvard.edu

Keywords: (mass spectrometry, metabolomics, cancer, selected reaction monitoring, LC/MS/MS, label-free, profiling)

Short Abstract: We present a platform for quantitative profiling of endogenous polar metabolites from cells, tissues and bodily fluids using selected reaction monitoring (SRM) tandem mass spectrometry. The platform targets >250 metabolites from a single 18 minute LC/MS/MS acquisition.

Long Abstract:  The revival of cancer cell metabolism in recent years has prompted analytical platforms for studying metabolites from in vivo sources (1). We implemented a quantitative polar metabolomics profiling platform using selected reaction monitoring (SRM) with a 5500 QTRAP hybrid triple quadrupole mass spectrometer. The platform uses normal phase chromatography with positive/negative ion switching in order to quantitatively target more than 250 metabolites from a single 18 minute LC/MS/MS experiment. The platform is compatible with polar metabolites from any source including tissues, cell lines and bodily fluids. No internal standards are used and Q1/Q3 transitions are integrated to produce peak areas based on total ion current (TIC) that can be used for statistical analyses and pathway analyses across biological sample conditions.

Sample Preparation Protocol

Polar metabolite extraction from cell lines:

- Be sure to have HPLC grade or higher methanol stored at -80C to be used for polar metabolite extraction. For cell lines, approximately 2-3 million cells should be prepared (one 10cm2 plate is sufficient). For tissue, at least 50-100 micrograms should be collected and 100-150 microliters of serum or urine should be collected for robust data.

- Grow cells. Wash plate with respective medium 1X 5 mL. Add 10 mL respective medium and incubate for 2 hours prior to metabolite collection

- Aspirate off medium. Tilt plate, aspirate off all media. Keeping plated tilted, wait a few seconds to allow any additional media to collect in corner of plate, and them aspirate to get off as much as possible

- Immediately add 4 mL 80% methanol (-80°C)

- Immediately transfer to -80°C. During transfer travel, place plates with 80% methanol on dry ice

- Incubate plates at -80°C for 15 minutes

- Scrape plates on dry ice with cell scraper

- Transfer cell lysate/methanol mixture to 15 mL conical tubes on dry ice

- Centrifuge at full speed for 5 minutes in cold room to pellet cell debris and proteins

- Transfer supernatant to 50 mL conical tubes on dry ice

DO NOT THROW AWAY 15 mL TUBES

- Add 500 μL 80% methanol (-80°C) to 15 mL tubes and resuspend pellet. Resuspending the pellet is a little difficult and may require a combination of vortexing and
pipetting up and down

- Transfer mixture to 1.5 mL Eppendorf tube on dry ice

- Spin in microcentrifuge at full speed for 5 minutes in cold room

KEEP 1.5 ML EPPENDORF TUBE

- Transfer supernatant to 50 mL conical tubes on dry ice (from step 9)

- Add 500 µL 80% methanol (-80°C) to 1.5 mL tubes and resuspend pellet

- Spin in microcentrifuge at full speed for 5 minutes in cold room

- Transfer supernatant to 50 mL conical tubes on dry ice (from step 9)

- After pooling the three extractions, the samples are completely dried or below.

Modification of metabolite extraction for tissue samples:

- Start with ~50-100 mg of solid tissue or the equivalent of at least 2-3 million cells

- add -80C of 80% HPLC grade methanol 400 µL to frozen tissue pieces

- grind for 30sec

- centrifuge at 14K for 10min

- transfer supernatant to new tube

- add ppt new -80C 80% methanol 400 µL

- grind for 30 sec

- transfer and combine supernatant

- additional centrifuge at 14K for 10min

- transfer supernatant to new 1.5mL microfuge tube

- speed vac/lyophilize to a pellet

Modification of metabolite extraction from bodily fluids (serum/urine/cerebrospinal fluid):

- Start with 100-150uL of fluid

- Add enough pure methanol -80C to make a 80% solution

- Store in the -80C freezer (-20C may also be sued) overnight or at least 6 hours

- Centrifuge at 14K for 10 min in cold room

- transfer supernatant to new 1.5mL microfuge tube

- speed vac/lyophilize to a pellet

- resuspend in 30uL of LC/MS grade water just prior to analysis

Modification of metabolite extraction from formalin fixed paraffin embedded (FFPE) tissue blocks:

- Take  ~30-50 µm FFPE slice for extraction (discard the top slice)

- Place curled slice in a  1.5 mL microfuge tube

- Prepare a water bath or heat block at 70C

- Add ~ 1mL of 80% methanol to the tube and place at 70C for 30-45 min

- Remove from heat and put tube in ice to congeal the wax

- Centrifuge cold (4-8C) at 14K for 10 min

- Transfer supernatant to a new 1.5 mL microfuge tube

- Centrige again cold at 14K for 5 min

- Transfer to a new tube and speed vac/lyophilize to dryness

- Add ~50 µL of LC/MS grade water to resuspend for MS analysis

Instrumental Analysis

Use a triple quadrupole mass spectrometer capable of positive/negative switching at a speed of 50 msec or below and sub 5 msec dwell time such as the 5500 Triple Quad and 5500 QTRAP (AB/Sciex) (1) or other triple quadrupole style mass spectrometer with similar capabilities. The fast polarity switch time and low dwell times allow for fast duty cycles for non-scheduled SRM and MRM runs. An HPLC is needed capable of flow rates in the several hundred microliters per minute range (~300 uL/min). A normal phase (HILIC) chromatography column is needed such as an XBridge Amide (3.5 µm) 4.6 mm i.d. x 100 mm length (Waters Corp.). Create an instrument method in selected reaction monitoring (SRM) mode according to manufacturer’s instructions.  Supplemental Table 1 lists the 298 Q1/Q3 transitions representing 254 unique metabolites using both negative and positive ion modes. Table 2 lists the KEGG identifiers for most of the targeted metabolites. It is not necessary to use scheduled SRM mode for this method since the instruments cited capable of scanning the SRM transition list in less than 1.8 seconds using 3 msec dwell times, producing at least 8-10 scans per metabolite, on average. Chromatographic scheduling is required if duty cycles are higher than ~2 sec because insufficient data points will be acquired per metabolite. The HPLC buffers should be prepared as follows: A buffer (95% LC/MS grade water, 5% LC/MS grade acetonitrile, 20mM ammonium hydroxide, 20 mM ammonium acetate), B buffer (100% LC/MS grade acetonitrile). The gradient should be run as follows:

Flow-rate, 300 µL/min
85% B, 0 min.

85% B to 35% B, 3.5 min.

35% B to 2% B, 11.5 min.

2% B, 16.5 min.

2% B to 85% B, 17.5 min.

85%, 25 min.

Peak Integration and Raw Results

Use a suitable software platform for peak area integration such as MultiQuant 2.0 from AB/SCIEX. Ideally, standard compounds from the metabolites of interest should be tested on the platform to assure proper chromatographic elution times, however, the largest peak in the spectrum via SRM is typically correct for ~90% of the targeted metabolites for carefully selected Q1/Q3 transitions typically between 4 min - 13 min in elution time using Amide Xbridge columns (Waters Corp.) (2). MultiQuant 2.0 software from AB/Sciex is used to integrate the peak areas from the Q3 TIC values across the chromatographic elution (3). If sufficient sample is present, a single dominant peak will typically be present that corresponds to the total ion current (TIC) from the Q1/Q3 transition. If multiple peaks are present per Q1/Q3 transition, it is important to run a standard compound to be sure of the proper chromatographic elution time. This is especially important when incorporating new compounds into your transition list. Figure 2 shows an example of a chromatographic Q3 peak for hexose-phosphate from two different biological samples from H929 multiple myeloma cells as well as the standard parameters used for peak integration for metabolomics analyses. Figure 3 shows a small portion of the text based peak area integration table exported from MultiQuant 2.0 software showing glycolytic intermediates and several different drug treatments used in an experiment with H929 cells. This output is considered the raw results from the metabolomics profiling experiment. Be aware that some metabolites and some biological conditions will show no data (N/A in table). This is due to either concentrations of metabolites below the instrument threshold for detection, degradation, poor fragmentation or poor ionization efficiency for a metabolite. From 254 unique metabolites (298 transitions), we typically acquire robust data for ~200 metabolites. It is important to note that the data do not represent absolute concentrations of metabolites so peak area values can only be used relatively to a control (reference) sample. However, one can perform a standard curve analysis using known concentrations of a standard compound for absolute quantification, if needed. Alternatively, an isotope labeled metabolite can be spiked into a sample for reference.

Data Processing and Representative Results

A variety of data analysis tools can be used including commercial software, freeware and in-house biostatistics tools such as MarkerView for PCA analysis, MatLab for a variety of statistics and graphical tools and a freeware online tool at www.Metaboanalyst.ca can be used to perform statistical analyses, graphical outputs and pathway mapping. Figure 3A shows an unsupervised hierarchical heat map from several conditions of growth factor stimulation in cancer cells including epidermal growth factor (EGF) and insulin at time 0’, 10 minutes and 30 minutes of stimulation. These data represent typical results from pour platform. We can routinely quantify approximately 200-220 metabolites from our targeted list of 254 metabolites using pos/neg switching with a single column setup.  This heat map was created using the Clustergram feature in MatLab software. Similar to microarray data, red represents up-regulated metabolites and green represents down-regulated metabolites across sample conditions. In addition, one can visualize quantitative metabolomics data using principal components analysis (PCA). We use the commercial software MarkerView from AB/Sciex for this purpose. Figure 3b shows a PCA clustergram from both 293T cells and H929 multiple myeloma cells from cells treated with either EGF or Insulin for 0, 15 min and 30 min using MarkerView. The data show that we can define three distinct clusters based on the metabolic profile. H929 cells are less responsive to growth factor stimulation, 293T are very responsive and form two distinct clusters separated by the time of stimulation rather than the specific growth factor used. These results may be explained because 293T have significant levels of both pERK and pAKT driving their proliferation while H929 cells may not be dependent on these signals. If one further looks into H929 cells in terms of the glycolytic intermediates in Figure 4, one will see that during serum stimulation for 20 minutes after starving cells of serum overnight. In 293T cells that expresses high levels of pAKT, the level of glycolytic intermediates are highly elevated compared to H929 cells that show lower levels of pAKT. (5)

Discussion

This protocol is useful for steady-state profiling of polar endogenous metabolites from cells and tissues. We have applied these analyses in several studies including The mode of quantification that is suggested is relative quantification across biological conditions. Alternatively, if one wishes to gain insight into the One should note that while the Waters Xbridge Amide columns are very robust, reproducible and have a long life for metabolomics analyses (~1500 analyses), they suffer from poor peak resolution on most HPLC systems. Alternatively, the Luna Amino columns from Phenomenex Corp. can be used since they produce higher resolution with normal flow HPLCs, however these columns have  a very short life at pH=9.0 with typically less than 500 successful analyses with high reproducibility. High pH based normal phase HILIC columns provide the best resins for capturing the largest numbers of polar metabolites during a single acquisition. Other modes of chromatography such as reversed phase C18 can be used effectively but HILIC captures the highest number of metabolites from a single analysis. Our platform is unique in that we take advantage of positive/negative ion switching during a single run to capture all metabolites in one analysis. The 5500 QTRAP has the ability to switch polarity on the fly in less than 50 msec. Similar platforms have been implemented mainly by the Josh Rabinowitz lab at Princeton Univ. (4, 5), though previous platforms utilize multiple runs and platforms per sample for analysis.

Acknowledgments

The authors would like to thank Matthew Vander Heiden, Edward Driggers and Joshua Rabinowitz for help with method development. We thank Jason Locasale and Leon Peshkin for help with data analysis tools, and Xuemei Yang for help with mass spectrometry experiments. The authors thank BIDMC for support and help with funding the mass spectrometry instrumentation. This research was supported by grants from the National Institutes of Health 5P01CA120964-05 (J.A.) and 5P30CA006516-46 (J.A.) and a generous donation by BIDMC.

Disclosures: I have nothing to disclose.

Table of specific reagents and equipment:

| Name of the reagent | Company | Catalogue number | Comments (optional) |
| --- | --- | --- | --- |
| Amide XBridge column Luna Amino column Ammonium hydroxide (25%) Ammonium acetate LC/MS grade water LC/MS grade acetonitrile 5500 QTRAP Prominence UFLC SpeedVac concentrator 12x32mm capped autosampler vials MultiQuant 2.0 software MarkerView 1.1 software | Waters Phenomenex Fisher Scientific Sigma-Aldrich Fisher Scientific Fisher scientific AB/SCIEX SHIMADZU Thermo electron corp. National Scientific AB/SCIEX AB/SCIEX | 186004868 00F-4378-B0 AC25521-0025 631-61-8 MWX00011 AC61514-0025 C4000-87 | Robust Qq-IT MS HPLC |

References:

- Bayley JP, Devilee P. Warburg tumours and the mechanisms of mitochondrial tumour suppressor genes. Barking up the right tree? Curr Opin Genet Dev. 2010, 20:324-9.

- http://www.absciex.com/Products/Mass-Spectrometers/QTRAP-Systems/AB-SCIEX-QTRAP-5500-LCMSMS-System

- http://www.absciex.com/Products/Software/MultiQuant-Software

- Bajad SU, Lu W, Kimball EH, Yuan J, Peterson C, Rabinowitz JD. Separation and quantitation of water soluble cellular metabolites by hydrophilic interaction chromatography-tandem mass spectrometry. J Chromatogr A. 2006, 1125:76-88.

- Lu W, Bennett BD, Rabinowitz JD. Analytical strategies for LC-MS-based targeted metabolomics. J Chromatogr B Analyt Technol Biomed Life Sci. 2008, 871:236-42.

- http://www.younglin.com/brochure_pdf/waters/XBridge%20Amide.pdf

- Plas DR, Talapatra S, Edinger AL, Rathmell JC, Thompson CB. Akt and Bcl-xL promote growth factor-independent survival through distinct effects on mitochondrial physiology. J Biol Chem. 2001, 276:12041-8.