TITLE  Aligning and calibrating optical tweezers

# Authors:

Steven C. Wasserman

Department of Biological Engineering

Massachusetts Institute of Technology

scwass@mit.edu

Sean P. Robinson

Department of Physics

Massachusetts Institute of Technology

spatrick@mit.edu

Jens Schumacher

Thorlabs Inc.

jschumacher@thorlabs.com

# Corresponding author: Jens Schumacher

Keywords:

optical tweezer, force calibration, optical trapping

# Short Abstract:

Optical tweezers are an invaluable tool for applying and measuring forces on biological systems at the molecular scale. For small displacements, the optical trapping force can be modeled as a linear spring.  Quantitative measurements require an accurate estimate of the spring constant. Three methods for obtaining the constant are demonstrated.

# Long Abstract:

The capability of optical tweezers to exert measurable forces on micron scale dielectric particles offers a unique and valuable tool for studying cell components such as biological polymers and molecular motors. In many investigations, optical tweezers apply force to functionalized microspheres that have been attached to molecules of interest. Optical tweezers apply a force toward the focus of the trapping laser beam with a magnitude proportional to the distance of the particle from the focus, for small displacements from the center of the trap. This allows the optical tweezers to be modeled by Hooke’s law, F = -kd, where d is the displacement from the center of the trap. As the laser beam passes through a trapped particle, it will be deflected by an amount that depends on the position of the particle. The deflection is converted to an electrical signal by a quadrant photodiode, which produces a voltage proportional to particle position v = Rd, for small displacements from the center of the trap. Accurate force measurements depend on precise calibration of the spring constant, k, and the responsivity of the particle position detector, R, which vary with laser power and particle properties.

R is measured by translating a fixed microsphere around the center of the trap with a piezo actuated stage and recording the resultant voltage waveforms. Three methods for ascertaining k are presented: equipartition, PSD rolloff, and Stokes’ drag. Since each method relies on a different physical principle, the combined results provide a convenient way to verify the calibration. The PSD rolloff method offers a particularly effective way to discover an inaccurate position detector calibration, since it does not depend on the detector responsivity like the other two methods.

The equipartation method equates the average potential energy of the particle in the trap to the thermal energy of the solvent molecules. In this method, a sample of particle positions is obtained by recording the position detector output over time and converting the voltages to distances using the detector responsivity, R. For PSD rolloff, the power spectral density of a time series of trapped particle positions is computed and fit to the response of a harmonic oscillator with known damping due to the viscosity of the solvent. In the Stokes method, the sample is translated in a range of velocities. A force balance between viscous drag on the particle and the trap force is computed.

# Protocol Text:

- Set up the piezo stage for computer control.

You will need a two-channel instrumentation amplifier, two piezo controllers, two strain gage readers, a DAQ card with two analog outputs, two BNC cables, and four SMA to BNC cables. You can use a commercial amplifier such as the Tektronix AM502 or build your own. The design files for the circuit board shown here are available on BatchPCB.com. If you are using a commercial amplifier, set the gain to 1 with no filtering and DC coupling on both the inverting and non-inverting inputs.

Connect DAQ output A0 to the inverting input of the instrumentation amplifier. The output monitor of the X-axis strain gage reader goes to the non-inverting input. Connect the output of the amplifier to the external input of the strain gage reader. Similarly, connect DAQ output A1, the Y-axis strain gage reader, and Y-axis controller to the second amplifier.

- Make a sample with a combination of stuck and freely diffusing microspheres.

You will need: microspheres; measurement medium (deionized water in this case); 1M NaCl; an adjustable pipettor; a sample cell; and valap or vacuum grease. Make the calibration sample as optically similar as possible to the sample you plan to measure.

Make about one cc each of two microsphere suspensions: a 1:1000 dilution in 1M NaCl, and a 1:50,000 dilution in the measurement medium. Here we are using 1 micron, silica microspheres at a concentration of 10% by weight. Adjust the dilutions, if you are using a different stock.

Fill the sample cell with about 15 microliters of the salt/microsphere solution and incubate in a humidor with the cover slip down for 10 minutes to 1 hour at room temperature. The high salt concentration will help the microspheres stick to the coverslip.

Remove the saltwater and remaining free beads by flushing the channel with about 300 microliters of measurement medium. Use a pipettor to build up water on one end of the channel and draw it through slowly with a vacuum or a wedge of filter paper. Don’t flow the rinse through too quickly or allow any bubbles go through the channel.

Now flow about 15 microliters of the second microsphere suspension into the channel. The small number of freely diffusing microspheres should stay suspended for some time. Seal the ends of the channel with vacuum grease or valap.

Apply a drop of microscope oil to the objective and mount the slide on the trap.

Assess the quality of the sample. There should be a small number of stuck beads in most fields of view. It should take at least a minute to find a free bead. Too many free microspheres will make it difficult to trap only one sphere for the duration of a measurement. If you find stuck beads only near edges, it is likely that you drew the rinse water through too quickly or allowed a bubble to pass through the channel.

- Align the instrument.

You will need laser safety goggles, two mounted fluorescent infrared alignment disks, and an IR card or viewer. The viewer is helpful, but the procedure can be completed with the card alone. If the trap is already aligned, skip this step.

The beam path will be open during the alignment procedure. Use best laser safety practices, as recommended by your institution. Wear appropriately specified laser safety goggles to protect against accidental exposure to the powerful and invisible trapping laser. Know beam path at all times. Use the minimum necessary power. Remove reflective clothing items and jewelry. Keep reflective tools away from the bench while the laser is on.

Remove the beam expander, objective, and condenser lenses. This is a good opportunity to clean them. Also remove the dichroic mirrors and LED illuminator. Cover the illuminator port for to prevent the beam from accidentally escaping.

Adjust the fiber coupler to produce a collimated beam. Set up two alignment disks and adjust the fiber coupler and steering mirror to center the beam in the horizontal path.

Ensure that the front surface of the horizontal dichroic mirror is centered in its mount and replace it in the cube. Use two mounted pinholes to measure the vertical path of the beam. Adjust the dichroic and right angle mirrors so that the beam passes through the center of both targets.

Replace the beam expander lenses. Small adjustments to the dichroic and steering mirrors may be necessary. Adjust the distance between the beam expander lenses to produce a collimated beam. Secure the lenses. You will refine this adjustment in a later step.

Replace the sample, objective and condenser. Temporarily replace the LED illuminator and bring an empty area of a sample into focus. Remove the LED and adjust the condenser height to collimate the beam that emerges from the LED illuminator port. Adjust the condenser X/Y translator so that beam is vertical. Replace the LED illuminator.

Move the detector lens to the middle of its range and replace the detector branch dichroic. Adjust the dichroic to zero the X/Y position reading.

Replace the IR blocking filter in front of the camera with a filter that has an optical density of about 1 for infrared light, such as the FM01 from ThorLabs. This will allow you to visualize the laser spot directly. Bring the reflection of the laser spot on the cover slip into focus. Adjust the camera and beam path to center the spot in the field of view. Assess the quality of the beam by observing the spot as you move the sample up and down. Replace the tubes enclosing the beam path and the position detector dichroic.

Trap a single, freely diffusing microsphere. Mark the location of the center of the trap. Now fine-tune the distance between the beam expander lenses to bring the microsphere into focus. Save an image of a trapped bead for use during calibration.

Use the detector X/Y translator to center the QPD position reading. If you can’t center the signal, adjust the kinematic platform and then refine the adjustment with the X/Y translator. Center the detector signal each time you change samples.

- Acquire calibration data.

With the single microsphere still in the trap, save several seconds of X/Y position data. This will be used to compute the trap stiffness by the equipartition and PSD roll off methods. Here, the sample rate is 100 KHz and we are saving five seconds of data.

For Stokes calibration, first set the stage oscillation amplitude and frequency. Save several seconds of X/Y position data with the stage in motion. Do this for both X and Y axes.

Now measure the QPD detector responsivity. Discard the trapped bead and bring a stuck microsphere into focus at the center of the trap. Move the stuck bead as close as possible to the trap center in all three dimensions. The image should look as similar as possible to the one you saved earlier. Take an X/Y scan and save the data. Here we are scanning 20 lines with 250 samples per line. Run the scan in both directions.

Repeat these steps for several microspheres at several power levels.

- Analyze the data.

To estimate the QPD position detector responsivity, load the XY scan data from a stuck bead into Matlab or another numerical computing environment. The file has four, tab-delimited columns that correspond to the X-QPD, Y-QPD, X strain gage, and Y strain gage voltages. Convert the strain gage voltages to distance units using the responsivity of the strain gage reader, in this case 2.2 microns per volt.

Plot the X and Y QPD voltages versus position along the corresponding axis. Locate the scan line where the trap passed closest to the center of the stuck microsphere. This line has the greatest X deflection and smallest Y deflection. Estimate the slope of X QPD voltage curve near the zero crossing to find the position detector responsivity in volts per micron. Reduce the uncertainty of your estimate by averaging several datasets for each power level. Repeat the process using an YX scan to estimate the Y responsivity.

Estimate the trap stiffness using the equipartition theorem. Load time data for a trapped microsphere. Convert the QPD and strain gage voltages to distance units. Subtract the average value from the signal. Find the trap stiffness by multiplying Kb by temperature and dividing by the mean squared displacement.

Compute the power spectral density of the X signal. Use pwelch in Matlab or a similar function. Estimate the cutoff frequency. Use the formula for a second order overdamped system in the Appleyard reference to find the trap stiffness.

Estimate the trap stiffness using Stokes’ theorem. Load a dataset taken with the stage oscillating in the X direction. Convert voltages into distance units. Compute the stage velocity by taking the difference of the stage position using, for example, the Matlab diff command. Plot X displacement versus velocity and estimate the slope of the line. The trap stiffness is 6 pi times the radius of the microsphere times the viscosity of the solvent times the slope of the line. Repeat the process for the Y direction.

Repeat the process for a number of beads and power levels. 
Representative Results:

Discussion:

Acknowledgments: List acknowledgements and funding sources.

Disclosures: Author Wasserman has no conflicts to disclose. Author Schumacher is an employee of ThorLabs.

Table of specific reagents and equipment:

| Name of the reagent | Company | Catalogue number | Comments (optional) |
| --- | --- | --- | --- |
| OTKB Optical Trap Kit | ThorLabs | OTKB |  |
| Wide Band Hot Mirror | ThorLabs | FM01 |  |
| 1µ silica microspheres | Bangs Labs | SS04N/9857 |  |
| DAQ Interface | National Instruments | USB-6212 |  |
| Instrumentation amplifier | Tektronix | AM502 | Or build your own from BatchPCB.com catalog number 57271 |
| Numerical analysis software | Mathworks | Matlab |  |

# References

- Ashkin, A. Acceleration and Trapping of Particles by Radiation Pressure. Phys Rev Lett 24,156-159 (1970)

- Ashkin, A. Forces of a single-beam gradiend laser trap on a dielectric sphere in the ray optics regime. Biophys J. 61, 569-582 (1992).

- Svoboda, K. & Block, S. M. Biological applications of optical forces. Annu. Rev. Biophys. Biomol. Struct. 23, 247-285 (1994)

- Appleyard, D. C., Vandermeulen, K. Y., Lee, H., Lang, M. J. Optical trapping for undergraduates. Am. J. Phys. 75, 5-14 (2007)

- Berg-Sorensen, K., Flyvbjerg, H.. Power spectrum analysis for optical tweezers. Rev. Sci. Instrum. 75(3), 594-612 (2004)

- Neuman, K. C., Block, S. M. Optical trapping. Rev. Sci. Instrum. 75(9), 2787-2809 (2004)

- Moffitt, J. R., Chemla, Y. R., Smith, S. B., & Bustamante, C. Recent Advances in Optical Tweezers. Annu. Rev. Biochem. 77, 205-228, (2008).

- ThorLabs Optical Trapping Kit Manual. http://www.thorlabs.com/Thorcat/19500/19590-D02.pdf. (2011)

- Batchelor, George Keith. An Introduction to Fluid Dynamics. Cambridge, United Kingdom: Cambridge University Press. 2000

- National Instruments

- Equipartition