Please suggest names of 5 peer reviewers with their institutional affiliation and email address.

Prof Stefan Debener, University of Oldenburg, Institute of Psychology, stefan.debener@uni-oldenburg.de

Dr Tom Eichele, University of Bergen, Faculty of Psychology, tom.eichele@psypb.uib.no

Dr David Carmichael, UCL Institute of Child Health London, d.carmichael@ucl.ac.uk

Dr Andrew Bagshaw, University of Birmingham, School of Psychology, a.p.bagshaw@bham.ac.uk

Prof Louis Lemieux, UCL Institute of Neurology London, louis.lemieux@ucl.ac.uk

JoVE Article Template

TITLE

Best practice for recording simultaneous EEG-fMRI data.

Authors:

Karen J Mullinger, Robert Störmer , Pierluigi Castellone, Richard W Bowtell

Authors: institution(s)/affiliation(s) for each author:

Karen J Mullinger

Sir Peter Mansfield Magnetic Resonance Centre, School of Physics and Astronomy

University of Nottingham

karen.mullinger@nottingham.ac.uk

Robert Störmer

Sales & Support

Brain Products GmbH

Robert.Stoermer@brainproducts.com

Pierluigi Castellone

Sales & Support

Brain Products GmbH

Pierluigi.Castellone@brainproducts.com

Richard W Bowtell

Sir Peter Mansfield Magnetic Resonance Centre, School of Physics and Astronomy

University of Nottingham

Richard.bowtell@nottingham.ac.uk

Corresponding author:  Richard W Bowtell

Keywords:  neuroscience, EEG, fMRI, simultaneous, recording

Short Abstract:

Simultaneous electroencephalography (EEG) and functional Magnetic Resonance imaging (fMRI) is a powerful neuroimaging tool.  However, the inside of an MRI scanner forms a difficult environment for EEG data recording and safety must be considered whenever operating EEG equipment inside a scanner.  Here, we present an optimised EEG-fMRI data acquisition protocol.

Long Abstract:

Simultaneous EEG-fMRI allows the excellent temporal resolution of EEG to be combined with the high spatial accuracy of fMRI.  There are a number of ways in which the data from these two modalities can be combined (Killner), but all rely on the acquisition of high quality EEG and fMRI data.  EEG data acquired during simultaneous fMRI are affected by several artefacts, including the gradient artefact (due to the changing magnetic fields required for fMRI), the pulse artefact (linked to the cardiac cycle) and movement artefacts (resulting from flux cutting by moving leads, as well as muscle artefacts) (Allen et al. 2000).  Post-processing methods for successfully correcting the gradient and pulse artefacts require a number of criteria to be adhered to during data acquisition.  Minimizing head motion during EEG-fMRI is also imperative for limiting the induction of artefacts (Marjie).

Interactions between the radio frequency pulses required for MRI and the EEG hardware may also occur and can cause heating, which is a potential hazard for the experimental subject (Lemieux).  This is only a significant risk if safety guidelines are not satisfied. Issues which must be considered include hardware design and set-up, as well as careful selection of which MR sequences are run with the EEG hardware present.

The above issues highlight the importance of the choice of the experimental protocol which is used when performing a simultaneous EEG-fMRI experiment.  Based on previous research (Mullinger et al. 2008, 2011; Mandelkow et al. 2006) we demonstrate an optimal experimental set-up which provides high quality EEG data during simultaneous fMRI when using commercial EEG and fMRI systems for data acquisition, with safety risks to the subject minimised.  We demonstrate this set-up in conjunction with a simple visual stimulus. However, it can also be used with much more complex stimuli.   This video shows the EEG-fMRI set-up using a Brain Products GmbH (Gilching, Germany) MR plus 32 channel EEG system with a Philips Achieva (Best, Netherlands) 3T MR scanner although many of the techniques should be transferable to other systems.

Protocol Text:

- Preparing the Experimental Setup

  - Prior to the subject’s arrival the EEG equipment must be set up in the control room where the scanner operator will sit. The laptop should be connected to the EEG hardware as shown in Figure 1. Note: all triggers from presentations or the MR scanner must be more than 200μs to be detected by the EEG software.

  - The stimulus computer should also be set up; here we are using a visual stimulus with markers at the beginning and end of the stimulation period to be read into BrainVision Recorder.

  - Markers for the scanner and presentation should be checked to confirm they are being read in correctly. The synchronistation of the scanner clocks can be turned on through the options in BrainVision Recorder (show the button clicking process). Successful synchronization is checked by the green icon and “Sync On” marker which will appear if set up is correct.

  - Set up the MR scanner in the conventional way; here we’re using a body transmit and 32 channel head receive coil. We’re using this specific coil because the channel in the top of it allows the cables from the EEG cap to run along a straight path out of the scanner.

  - Make sure you have the MR sequences that you wish to run all set up. The EPI sequence must use a slice TR which is a multiple of the EEG clock period (200 μs). If using a Philips MR system the Philips Timing calculator can be used to determine the possible slice and TR combinations.

  - Make one final check that all equipment is recording as expected.

- Subject arrival

  - Ask the subject to arrive with clean hair and in comfortable non-metallic clothing.

  - Explain to the subject the purpose of the experiment and what will happen.

  - Ask the subject to fill out paper work to confirm they are safe to go in an MR scanner and they give consent for the experiment. Check this before proceeding.

  - Measure head and select appropriate size cap.  Place cap on head and position correctly , i.e. FCz electrode between the ears and aligned with the nasion and inion.

  - To connect the electrodes to the head: move hair out of the way, apply alcohol and then Abralyte gel. The ECG electrode, used to measure the heart beat, should be attached to the base of the back using similar methods to that for the cap electrodes.

  - It is important to get the impedances of electrodes below 10kΩ (show working down the scale) apart from the cardiac monitoring electrode.

  - To check the data quality is satisfactory visually inspect the data on the monitor screen.

- Recording outside of the MR scanner (optional: only required if you wish to compare EEG data quality from inside and outside the MR scanner)

  - Set up presentation apparatus and EEG equipment in non-magnetic field.  Ensure setup for the subject is as similar as inside the MR scanner as possible (way stimulus is presented and subject in supine position).

  - Perform experiment and record the data in similar way to inside the scanner (see section 4).

- Setting subject up inside the MR scanner

  - Seat the subject whilst you set up the EEG equipment in the MR room.

  - Take the amplifiers into the room and place them on a table at the back of the scanner and attach them to a long fiber optic which feeds data out of the scanner room through a waveguide.

  - Register the patient in the MR software.

  - Take the subject into the room and ask them to lie on the bed.

  - Give the subject earplugs, head phones, call button and make sure that they are comfortable.

  - Put the head coil over the subject making sure the EEG cables leave the head coil along the shortest path. Then pad the subject’s head to minimize head movement. (Show head immobilization in head coil.)

  - Move the subject into the scanner bore ensuring that electrodes Fp1 and Fp2 will be at isocentre (in the z-direction).

  - Now go around the back of the scanner and attach the EEG cap to the amplifier. Ensure that there are no wire loops (as this can cause heating and larger EEG artefacts to be induced) and that cabling is isolated from MR scanner vibrations as much as possible; here we use a cantilevered beam to achieve this.

- Recording inside the scanner

  - Now the MR scanner operator will talk to the subject from the console room and ask if they can hear them and are OK.

  - A second experimenter will start the EEG monitoring and check for noisy electrodes in the traces as well as the “Sync On” dot at the bottom of the screen.

  - The clear effect of the cryo-pumps on the recording can be seen (show traces with them on and off and the online FFT in RecView). Therefore these pumps should be switched off during data acquisition.

  - The importance of keeping the head still can be seen by asking the subject to move their head by a small amount.

  - The recording of neuronal activity can also be tested by asking the subject to open and close their eyes and look for occipital alpha activity. This should be done to test you are measuring physiological signals rather than noise.

  - The pulse artifact can clearly be seen in the raw data particularly on electrodes over the temples but can be corrected for in RecView in real time using the ECG trace.

  - Once you are happy with the data quality and the subject is happy you can start the MRI preparatory scans and plan the slice positioning for the fMRI.

  - As soon as each MRI scan begins the gradients will cause a large artefact in the EEG data.

  - When you are ready for the fMRI experiment to begin with the presentation of stimuli ready then the EEG data should be saved following the steps shown.

  - Now start the experiment checking that you can see the markers from your presentation and the MR scanner in BrainVision Recorder.

  - The EEG data quality will look poor however it can be cleaned either online in RecView or during post processing. In order for the gradient artefact correction to work the stimulus must not be locked to the TR and the frequency of the stimulus must not be equal to the slice repetition frequency.

  - Gradient artefact correction must occur prior to the pulse artefact correction. Data then may be segmented according to your stimuli and analysed with numerous techniques; the simplest of which is averaging to look at evoked responses. (Show each of these stages in RecView.)

- Debriefing the subject

  - Once the scanning is complete take the subject out of the scanner and help them to take off the EEG cap.

  - Provide them with the equipment required to wash their hair and ask them to wash it.

  - When they return they are free to leave.

- Set down

  - Unset the EEG equipment as required by your laboratory. If the MR manufacture requires it make sure that the synchronization hardware is unplugged at the end of each session and not left attached to the scanner.

  - Finally the EEG cap must be cleaned. To do this soak the cap in a water and disinfectant mixture and then use a tooth brush to clean away residual gel. It is very important to clean the cap fully to ensure the proper performance of the cap in subsequent uses.

- Analysis

  - Here we have shown real time EEG analysis, however it is also possible and sometimes desirable to post-process the EEG data.  This can be done in a number of analysis packages such as Brain Products Analyzer2 or EEGLAB.

  - Gradient and pulse artifact correction should then be performed which can be done using a variety of methods such as: average artifact subtraction (commonly used for gradient correction and often used for pulse artifact correction), independent component analysis  or optimal basis sets (for pulse artifact correction) .

  - Data may then be analysed in the time or frequency domain to look at evoked responses as well as on-going oscillatory activity.

  - Here we have recorded the ECG trace using the Brain Products system to obtain a cardiac trace, necessary for pulse artifact correction.  In the standard setup the ECG trace is recorded by means of a dedicated electrode placed on the back of the subject. At our lab we use also a non-standard solution, which employs a vectrocardiogram to get the cardiac trace (this solution is available only with the Philips physiological monitoring equipment). We have found this can be useful if a clean trace cannot be obtained using the conventional ECG.

Representative Results:

To be written

Tables and Figures (Required):

Figure 1: Experimental set up

Figure 2: Raw data

Figure 3: Gradient artifact correction (time and frequency plots)

Figure 4: Puls artifact correction (time and frequency plots)

Figure 5: Segmented data (showing advantage of single trial)

Figure6: Average evoked response

Discussion:

EEG-fMRI is a powerful tool for studying brain function as the high temporal resolution of the EEG can be combined with the high spatial resolution of fMRI.  To date a number of studies have used this multi-modal approach to gain a better understanding of brain function.  It has been used both on healthy volunteers to investigate the correlation between oscillatory rhythms (measured with EEG) and blood oxygenation responses (using blood oxygenation level dependant (BOLD) fMRI) (Laufs).  It has also been used to study if characteristics of the evoked signal can explain the variance in the BOLD signal on a trial-by- trial basis (Debener, Eichele).  In clinical studies the main use of the technique has been to investigate the foci of epileptic seizers which are inherently difficult to localize non-invasively (Lemieux).  These examples show the power of this multi-modal imaging tool, however, to enable the study of phenomena such as these it is important to get the best possible quality of EEG and MRI data.To achieve this inside the MR scanner it is important to have the best experimental set-up and also choose the most appropriate analysis methods. Whilst optimal analysis methods will to some extent depend upon the research question of interest the experimental set-up will be relatively independent of this and should follow the guidelines outlined in this work.

The most important steps in the process of EEG data acquisition during simultaneous fMRI are to ensure that all external noise sources have been minimized (e.g. cyrocooler pumps and vibration of the EEG equipment).  Then to ensure optimal gradient artifact correction it is important to ensure that your EEG and MR scanner clocks are synchronized and the slice TR is a multiple of the scanner clock period and the subject is optimally positioned.  To ensure optimal pulse artifact correction many techniques require a clean cardiac trace to detect R-peaks from, we suggest that this can be best achieved using a VCG although is also possible with a well positioned ECG lead.

Since the physical layout of all scanner rooms is different we recognize that you may not be able to position your EEG amplifiers out of the bore of the magnet and in this case would suggest that they were put on some significant padding to attempt to decouple them from the scanner vibrations.  If you find that the gradient artifact correction is not working well then check the times between your volume or slice markers as it is likely that the TR you are inputting to the MR console is not precisely the TR that is output.  In this case you will need to contact your MR manufacturer for further assistance.

Even with this optimal set-up combined with the best available post-processing methods for cleaning the EEG data it is still not currently possible to remove all of the gradient or pulse artifacts and therefore these possess limits on the frequency bands which can easily be studied inside the MR environment.  Additionally, motion artefacts in the EEG data are often difficult to correct for and small ones can even be difficult to identify although they still may dominate neuronal signals.  These motion artefacts can cause spurious but plausible correlations with the fMRI data (Jansen 2012) and therefore it is imperative to design tasks to minimize the possibility of subject movement and also restrain the subject as much as possible.

Whilst considerable research is on-going to improve data acquisition and analysis methods for simultaneous EEG-fMRI it is possible once the technique outlined here is mastered to ask some important neuroscience questions combining the power of the high spatial resolution of fMRI with the temporal resolution of EEG.

Acknowledgments:

We would like to thank Brain Products GmbH for providing their equipment, expertise and help.

Disclosures:

The production of this article was sponsored by Brain Products GmbH. Robert Störmer and Pierluigi Castellone are employees of Brain Products GmbH, which manufactures some instruments and software used in this article.

Table of specific equipment/software:

| Name of the equipment | Company | Comments |
| --- | --- | --- |
|  | Brain Products GmbH |  |

References:

1.	Allen, P.J., O. Josephs,R. Turner, A Method for removing Imaging Artifact from Continuous EEG Recorded during Functional MRI. Neuroimage, 2000. 12(2): p. 230-239.

2.	Allen, P.J., G. Poizzi, K. Krakow, D.R. Fish,L. Lemieux, Identification of EEG Events in the MR Scanner: The Problem of Pulse Artifact and a Method for Its Subtraction.  . Neuroimage, 1998. 8(3): p. 229-239.

3.	Debener, S., K.J. Mullinger, R.K. Niazy,R.W. Bowtell, Properties of the ballistocardiogram artefact as revealed by EEG recordings at 1.5, 3 and 7 Tesla static magnetic field strength. Int. J. of Psychophys., 2008. 67(3): p. 189-199.

4.	Debener, S., A. Strobel, B. Sorger, J. Peters, C. Kranczioch, A.K. Engel,R. Goebel, Improved quality of auditory event-related potentials recorded simultaneously with 3-T fMRI: removal of the ballistocardiogram artefact. Neuroimage, 2007. 34: p. 587-597.

5.	Debener, S., M. Ullsperger, M. Siegel,A.K. Engel, Single-trial EEG-fMRI reveals the dynamics of cognitive function. Trends in Cog. Sci., 2006. 10(12): p. 558-563.

6.	Naizy, R.K., C.F. Bechmann, G.D. Iannetti, J.M. Brady,S.M. Smith, Removal of fMRI environment artifacts from EEG data using optimal basis sets. Neuroimage, 2005. 28(3): p. 720-737.

7.	Mullinger, K.J., M.J. Brookes, C.M. Stevenson, P.S. Morgan,R.W. Bowtell, Exploring the feasibility of simultaneous EEG/fMRI at 7 T. MRI, 2008. 26(7): p. 607-616.

8.	Mullinger, K.J., P.S. Morgan,R.W. Bowtell, Improved Artefact Correction for Combined Electroencephalography/Functional MRI by means of Synchronization and use of VCG Recordings. JMRI, 2008. 27(3): p. 607-616.

9.	Mullinger, K.J., W.X. Yan,R.W. Bowtell, Reducing the Gradient Artefact in Simultaneous EEG-fMRI by Adjusting the Subject's Axial Position. NeuroImage, 2011. 54(3): p. 1942-1950.

10.	Mandelkow, H., P. Halder, P. Boesiger,D. Brandeis, Synchronization facilitates removal of MRI artefacts from concurrent EEG recordings and increases usable bandwidth. Neuroimage, 2006. 32(3): p. 1120-6.

11.	Jansen, M., T.P. White, K.J. Mullinger, L. E.B., P.A. Gowland, S.T. Francis, R.W. Bowtell,P.F. Liddle, Motion-related artefacts in EEG predict neuronally plausible patterns of activation in fMRI data. Neuroimage, 2012. 59: p. 261-270.