Title:  Hollow microneedle-based sensor for multiplexed transdermal electrochemical sensing
Authors: Philip R. Miller1,2, Shelby A. Skoog1, Thayne L. Edwards2, David R. Wheeler2, Xiaoyin Xiao2 ,Susan M. Brozik2, Ronen Polsky*2, Roger J. Narayan*1 

1Joint Department of Biomedical Engineering, University of North Carolina and North Carolina State University, Raleigh, North Carolina 27695-7115, USA
2Department of Biosensors and Nanomaterials, Sandia National Laboratories, Albuquerque, New Mexico 87185, USA

Corresponding author: Ronen Polsky, Roger Narayan
Keywords: Microneedle, Microneedle sensors, multiplexed detection, electrochemistry, stereolithography 

Short Abstract: 

This article details the creation of a multiplexed microneedle-based sensor. The device is being developed for in situ sampling and electrochemical analysis of multiple analytes in a rapid and selective manner. We envision clinical medicine and biomedical research uses for microneedle-based sensors.
Long Abstract: 

The development of a minimally invasive multiplexed monitoring system for rapid analysis of biologically-relevant molecules could offer individuals suffering from chronic medical conditions facile assessment of their immediate physiological state. Furthermore, it could serve as a research tool for analysis of complex, multifactorial medical conditions. In order for such a multianalyte sensor to be realized, it must be minimally invasive, sampling of interstitial fluid must occur without pain or harm to the user, and analysis must be rapid as well as selective. 

Initially developed for pain-free drug delivery, microneedles have been used to deliver vaccines and pharmacologic agents (e.g., insulin) through the skin.1-2 Since these devices access the interstitial space, microneedles that are integrated with microelectrodes can be used as transdermal electrochemical sensors. Selective detection of glucose, glutamate, lactate, hydrogen peroxide, and ascorbic acid has been demonstrated using integrated microneedle-electrode devices with carbon fibers, modified carbon pastes, and platinum-coated polymer microneedles as transducing elements.3-7 

This microneedle sensor technology has enabled a novel and sophisticated analytical approach for in situ and simultaneous detection of multiple analytes. Multiplexing offers the possibility of monitoring complex microenvironments, which are otherwise difficult to characterize in a rapid and minimally invasive manner.  For example, this technology could be utilized for simultaneous monitoring of extracellular levels of, glucose, lactate and pH, which are important metabolic indicators of disease states7-12 (e.g., cancer proliferation) and exercise-induced acidosis.13
1.) Microneedle Fabrication
1.1) Using three-dimensional modeling software Solidworks® (Dassault Systemes SA, Velizy, France), design a pyramidal-shaped hollow microneedle array (Figure 1).3-5 
1.2) Design a support structure for the microneedle array using Magics® RP 13 software (Materialise NV, Leuven, Belgium). The support structure allows the resin to drain from the device during fabrication and provides a base on which the microneedles are built. An example support structure is shown in Figure 1.

1.3) The linked support and microneedle array files are uploaded into the Perfactory RP software (EnvisionTEC GmbH, Gladbeck, Germany), which controls the fabrication process. Within this software package, choose the number of microneedle arrays to be fabricated and determine the placement of devices on the fabrication plate.

1.4) Run calibration in ultraviolet mode at 180 mW for the Perfactory rapid prototyping manufacturing system and verify the deviation in energy is within
[image: image1.wmf]  

±

2

mW

.

1.5) Once fabrication is complete, remove the microneedle arrays from the baseplate and develop in isopropanol for 15 minutes. Dry the arrays with compressed air and cure the needles at room temperature for 50 seconds in the Otoflash® Postcuring System (EnvisionTEC GmbH, Gladbeck, Germany) to ensure complete polymerization.

1.6) Validate microneedle fabrication via microscopy and verify that each needle bore is hollow and unobstructed. Fully-fabricated microneedles are shown in Figure 2.

2.) Fabrication of Carbon Paste Electrode Arrays

2.1) Use a 60W Model 6.75 CO2 raster/vector laser system (Universal Laser Systems, Inc., Scottsdale, AZ) to cut holes and expose underlying individually addressable connecting copper wires in a flat flexible cable (21039-0249) from a commercial source (Molex Connector Corp., Lisle, IL) (Figure 3 (A and B)).  Place the flat flexible cables in a jig to properly align them on the laser ablation plate. Use a rastering approach to create 500 m diameter cavities in the insulating portion of the flexible cable. Patterns for ablation are created in CorelDraw (Corel, Ottawa, Ontario) and sent to the laser system.  

2.2) Clean the modified flat flexible cables with an airbrush that sprays acetone at 40 psi. Finish cleaning them by rinsing with isopropanol and deionized water. Verify under a microscope that no insulating film remains over the exposed copper strips. 

2.3) The next step is to create a holding cavity for packing of carbon pastes. Melinex® tape (0.002 inch thickness coated single-side with pressure sensitive acrylic adhesive) is ablated with the same pattern as the electrode strips, oriented over the ablated electrode strips, and compressed at 3000 psi for 2 minutes to ensure a proper connection. In this case, the cavity diameter is 750 m.
2.4) An additional layer of double-sided adhesive (Melinex®) tape (thickness= 0.004”) is subsequently ablated in the same pattern as the single-sided adhesive tape and is used after alignment to bond the microneedle arrays to the carbon paste electrode arrays. 

3) Synthesis of Functional Carbon Pastes and Packing of Electrode Cavities

3.1) The glucose sensitive carbon paste: Mix 10 mg of glucose oxidase and 2.2 mg of poly(ethylenimine) until a homogeneous mixture is obtained. To this mixture, 60 mg of rhodium on carbon powder (5% loading) is added. 40 mg mineral oil is added and subsequently mixed. The pastes are stored at 4o Celsius until use; the pastes are used up to one week after preparation. 
3.2) The pH sensitive carbon paste: Mix 30% (w/w) mineral oil and 70% (w/w) graphite powder. Pack paste into electrode cavity as described in section 3.4. Make solution of 10 mM Fast Blue RR diazonium salt (4-benzoylamino-2,5-dimethoxybenzenediazonium chloride hemi (zinc chloride) salt) in 0.5 M phosphoric acid. Place 20 (l drop of this solution over the packed paste electrode for 30 minutes to spontaneously chemisorb the Fast Blue PR. Rinse with deionized water and store in buffer or deionized water when not in use.

3.3) The lactate sensitive carbon paste: Mix 2.5 mg of rhodium on carbon powder and 2.5 mg lactate oxidase, alternating between 5 minutes of sonication and 5 minutes of vortexing for five rotations. 

3.4) Packing of the modified pastes into the prepared flat flexible cable is accomplished by applying the respective pastes over electrode cavities.  Using a thin piece of plastic (e.g., an edge of a plastic weigh boat) as a trowel and pack the paste until a smooth surface is achieved. Repeat with a second clean weighing boat until excess paste is removed. Wash with deionized water. A schematic showing laser ablation to create cavities, packing of carbon pastes, and microneedle integration (described in section 2 and 3) is presented in Figure 3.

4)  Detection and sensor calibration
4.1) Lactate detection is accomplished by measuring the chronoamperometric response of the sensor at -0.15 V and recording the current after 15 seconds in 0.1 M phosphate buffer.
4.2) Glucose detection is performed in a similar manner by measuring the chronoamperometric response of the sensor at -0.05 V and recording the current after 15 seconds in 0.1 M phosphate buffer.

4.3)   pH is monitored  by running cyclic voltammetric scans from -0.7 V to 0.8 V at 100 mV/s and recording the position of the oxidative peak potential. 

4.4) Calibration curves for glucose and lactate sensors can be created by successive additions of the respective analyte; chronoamperometric measurements are performing after each analyte addition as described in sections 5.1 and 5.2. Alternatively, fixed potential chronoamperometric measurements can be made under stirring while allowing sufficient time (~10-100 seconds) in between each analyte addition for current stabilization.

4.5) pH calibration curves can be created by measuring the position of the  oxidative peak potential over a series of known pH values from 5 to 8 in 1.0 pH unit increments and recording cyclic voltammograms as described in section 5.3.

5. Representative Results: 

When obtaining chronoamperometric curves (e.g., for glucose detection or lactate detection) in quiescent solutions with modified carbon paste-filled needles, the current will immediately decrease upon application of the respective detection potential. It will eventually decay to a steady state value. This result is shown in Figure 4; this result was obtained from 2 mM additions of lactate and recording at the lactate needle. The solution must be briefly stirred after each lactate addition. The current after 15 seconds rises upon increasing the concentration of lactate; the current response can then be used to determine the concentration of lactate in an unknown solution. Alternatively, continuous monitoring can be used in a stirred solution (or in a flowing solution) as demonstrated for a solution with an increasing glucose concentration (Figure 5).  Again, the increase in current upon increasing the concentration of glucose can be used to standardize the glucose response of an unknown solution. Sufficent time must be allowed after each spike to allow the solution to stabilize. Cyclic voltammograms at the pH sensitive needle in 0.1 M phosphate buffer are shown in Figure 6 over four different pH solutions from 5 to 8 in 1 pH unit increments. The oxidative peak potential will shift with increasing pH; this phenomenon is used as an indicator of the pH value.

6. Discussion

Multiple aspects of the design of this microneedle-based sensor were considered prior to device fabrication. In order for this sensor to be used for real-time detection, the  response time of the sensor must be rapid; in this protocol, each tested sensor exhibited a response time below fifteen seconds. Pastes used in this protocol were also chosen for their selectivity within in vivo environments, which contain electroactive biomolecules that can interfere with electrode response. In addition to paste composition, the operating potentials were chosen to minimize the influence of  interfering electroactive species. Successful fabrication of the microneedle array involves the selection of a suitable microneedle design and microneedle material. These two aspects will determine if the needle can puncture the skin, protect the electrodes from physical damage, and preclude electrode-tissue contact. It should be noted that an external Ag/AgCl and Pt reference and counter electrodes were used during measurements; in vivo use of this device with human or animal subjects would require that these electrodes be incorporated within the device. 

Each component of the microneedle-based sensor has features that must be validated to ensure proper functionality. Quality control during modification to the flat flexible cable (Figure 3 B) involves ensuring that the insulating layer is completely removed from the surface of the tin-plated copper conductors after laser ablation (Figure 3). Failure to remove the insulating layer from the surface of the copper conductor after laser ablation can cause irregular responses due to incomplete electrical contact. The laser-ablated Melinex® tape needs to be examined with a microscope to ensure that the diameter of each opening is consistent as it defines the working area of the electrode. When applying carbon paste to the laser-ablated Melinex® tape cavities, the paste must conform to the exact hole diameter with no excess in order to avoid signal variations due to differences in surface area. During chronoamperometric measurements with modified carbon pastes, the signal must stabilize to a limiting value before current is recorded. These results can slightly vary due to mixing effects. Mechanical testing of the microneedle arrays was done prior to sensor incorporation; in a previous study, our group showed these arrays were capable of puncturing porcine skin, which is often used as an analog for human skin.3 Microneedle arrays should not undergo deformation or fracture during skin penetration since these processes could lead to electrode damage.
This protocol has detailed the creation of a novel transdermal device for electrochemical monitoring. We envision future efforts involving microneedle sensors with an even greater number of individually addressable microneedles and a greater variety of transducers. This device was designed for analysis of interstitial fluid in humans; use with animals is also possible with appropriate species-specific modifications to microneedle design. Future directions with this technology include but are not limited to remote patient monitoring/diagnosis, monitoring of dynamic movements, or coupling with a drug delivery device for combined sensing-drug delivery.
Acknowledgments: Sandia is multiprogram laboratory operated by Sandia Corporation, a Lockheed Martin Company, for the United Stated Department of Energy’s National Nuclear Security Administration under Contract DE-AC04-94AL85000. The authors acknowledge funding from Sandia National Laboratories' Laboratory Directed Research & Development (LDRD) program. 

Disclosures: We have nothing to disclose.  

Table of specific reagents and equipment:

	Name of the reagent
	Company
	Catalogue number
	Comments (optional)

	DMEM
	Invitrogen
	ABCD1234
	


References:  

1. Henry, S., McAllister, D. V., Allen, M. G. & Prausnitz, M. R. Microfabricated microneedles:

a novel approach to transdermal drug delivery. J. Pharm. Sci. 87, 922-925 (1998).
2. Prausnitz, M. R. Microneedles for transdermal drug delivery. Adv. Drug Deliv. Rev. 56, 581-587 (2004).
3. Miller, P. R., Gittard, S. D., Edwards, T. L., Lopez, D. M., Xiao, X., Wheeler, D. R., Monteiro-Riviere, N. A., Brozik, S. M., Polsky, R. & Narayan, R. J. Integrated carbon fiber electrodes within hollow polymer microneedles for transdermal electrochemical sensing. Biomicrofluidics 5, 013415 (2011).

4. Windmiller, J. R., Zhou, N., Chuang, M. C., Valdés-Ramírez, G., Santhosh, P., Miller, P. R., Narayan, R. & Wang, J. Microneedle array-based carbon paste amperometric sensors and biosensors. Analyst 136, 1846-1851 (2011). 

5. Windmiller, J. R., Valdés-Ramírez, G., Zhou, N., Zhou, M., Miller, P. R., Jin, C., Brozik, S. M., Polsky, R., Katz, E., Narayan, R. & Wang, J. Bicomponent microneedle array biosensor for minimally-invasive glutamate monitoring. Electroanal. 23, 2302-2309 (2011).

6. Ricci, F., Moscone, D. & Palleschi, G. Ex vivo continuous glucose monitoringwith microdalysis technique: The example of GlucoDay. IEEE Sensors J. 8, 63-70 (2008).   

7. Zimmermann, S., Fienbork, D., Flounders, A. W. & Liepmann, D. In-device enzyme immobilization: Wafer-level fabrication of an integrated glucose sensor. Sens. Actuat. B 99, 163-173 (2004).
8. Rofstad, E. K. Microenvironment-induced cancer metastasis. Int. J. Radiat. Biol. 76, 589-605 (2000).

9. Vander Heiden M. G., Cantley L. C. & Thompson, C. B.  Understanding the Warburg effect: the metabolic requirements of cell proliferation. Science 324, 1029-1033 (2009). 

10. Warburg, O., Wind, F., Negelein, E. The metabolism of tumors in the body. J. Gen. Physiol. 8, 519-530 (1927).

11. Goode, J. A. & Chadwick, D. J., Eds. The Tumour Microenvironment: Causes and Consequences of Hypoxia and Acidity (Novartis Foundation Symposium 240, John Wiley & Sons, Ltd., 2008). 

12. Cardone, R. A., Casavola, V. & Reshkin, S. J. The role of disturbed pH dynamics and the Na+/H+ exchanger in metastasis. Nature Rev. Cancer 5, 786-795 (2005). 
13. Robergs, R. A., Ghiasvand, F. & Parker, D. Biochemistry of exercise-induced metabolic acidosis. Am. J. Phys. 287, R502-R516 (2004).

[image: image2.png]



Figure 1: Images of the STL file of the microneedle array created in Solidworks (A) and of the print screen, which shows the linked STL file and the support structure (B). 

[image: image3.png]i

Analytical Instrumentation Facility NCSU

5.0 KV EM Mag 70X
1mm 400um

Analytical Instrumentation Facility NCSU
5.0 KV EM Mag 25X





Figure 2: Scanning electron micrographs of the microneedle array (A) and a single microneedle within this array (B). 

[image: image4.png]Flexible Flat Cable

carbon paste

/ tape





Figure 3: Schematic of flat flexible cable assembly. The steps involved include modifying the flat flexible cable (A), ablating the patterned circles (B), adding the initially ablated Melinex® layer, which the carbon paste is squeezed into (C), as well as adding the second ablated Melinex® layer and mating the microneedle array (D).

[image: image5.png]Current (A)

000e00

soue0s

“Loce07

“s0e07

20007

2s0c07

o007

asee0r

asee0r

Time (s)

1

%

—omm
— 2w
—amm
— e
— e
—1omm
12w




Figure 4: Calibration of lactate sensitive paste with 15 second chronoamperometric scans at -

0.15V in 0.1 M phosphate buffer. Each increase in current corresponds to a 2 mM addition of lactate.

[image: image6.png]Current (A)

Time (s)

-4E-08 1

-5E-08 1

-6E-08 1

-7E-08 4

350

450 550 650 750

2mM Glucose

<

2mM Glucose

2mM Glucose

2mM Glucose

2mM Glucose

-8E-08




Figure 5: Calibration of glucose sensitive paste running chronoamperometric scans at -0.05 V. Each decrease in the response corresponds to a 2 mM addition of glucose. The 0.1 M phosphate buffer solution was stirred during calibration. External Ag/AgCl and Pt reference and counter electrodes were used in this study. 
[image: image7.png]



Figure 6: Cyclic voltommogram (CV) of pH sensitive carbon paste in 0.1 M phosphate buffer over pH 5-8 in 1 pH unit increments. A fifth CV was used for analysis versus an Ag/AgCl reference and Pt wire counter electrodes. 

_1255629656.unknown

