OFDI Imaging of Ex Vivo Pulmonary Resection Specimens: Obtaining One to One Image to Histopathology Correlation 

Authors: 
Lida P. Hariri, Matthew B. Applegate, Mari Mino-Kenudson, Eugene J. Mark, Brett E. Bouma, Guillermo J. Tearney, Melissa J. Suter

Authors: institution(s)/affiliation(s) for each author:

Lida P. Hariri

Department of Pathology

Wellman Center for Photomedicine
Harvard Medical School
Massachusetts General Hospital, Boston, MA

lhariri@partners.org

Mari Mino-Kenudson

Department of Pathology
Harvard Medical School
Massachusetts General Hospital, Boston, MA

Eugene J. Mark

Department of Pathology
Harvard Medical School
Massachusetts General Hospital, Boston, MA

lhariri@partners.org

Matthew B. Applegate

Pulmonary and Critical Care Unit

Massachusetts General Hospital, Boston, MA

Brett E. Bouma

Wellman Center for Photomedicine
Harvard Medical School
Massachusetts General Hospital, Boston, MA

Guillermo J. Tearney

Department of Pathology

Wellman Center for Photomedicine
Harvard Medical School
Massachusetts General Hospital, Boston, MA

Melissa J. Suter

Pulmonary and Critical Care Unit
Wellman Center for Photomedicine

Harvard Medical School
Massachusetts General Hospital, Boston, MA

Corresponding author: Melissa J. Suter
Keywords: Bronchoscopic imaging, In vivo optical microscopy, Optical imaging, Optical coherence tomography, Optical frequency domain imaging, Histopathology correlation
Short Abstract: 

A method to image ex vivo pulmonary resection specimens with optical frequency domain imaging (OFDI) and obtain precise correlation to histopathology is described, which is essential to developing specific OFDI interpretation criteria for pulmonary pathology. This method is applicable to other tissue types to obtain precise imaging to histopathology correlation for accurate image interpretation and assessment.  

Long Abstract: 

Lung cancer is the leading cause of cancer-related deaths7





2-7

. OFDI has been utilized in a variety of applications, including evaluation of coronary artery atherosclerosis


 ADDIN EN.CITE 


1

. Squamous cell and small cell cancers typically arise in association with the conducting airways, whereas adenocarcinomas are typically more peripheral in location. Lung malignancy detection early in the disease process may be difficult due to several limitations: radiological resolution, bronchoscopic evaluation of tissue underlying the airway mucosa and identification of early pathologic changes, and small sample size and/or incomplete sampling in histopathological biopsies. High resolution imaging modalities, such as optical frequency domain imaging (OFDI), provide non-destructive, large area 3-dimensional views of tissue microstructure to depths approaching 2 mm in real time (Fig. 1) HYPERLINK  \l "_ENREF_1" \o "Jemal, 2007 #70" 


 ADDIN EN.CITE ,8
 and esophageal intestinal metaplasia and dysplasia7


 ADDIN EN.CITE ,9-11
. Bronchoscopic OFDI has been demonstrated in vivo as a safe technique 
Here, we present a method for precise, one to one correlation between high resolution, volumetric optical imaging and histopathology in ex vivo lung resection specimens. We performed airway centered OFDI in ex vivo surgical resection specimens with a specialized custom built bronchoscopic 2.4 French (0.8 mm diameter) catheter. Tissue samples were marked with tissue dye, visible in both OFDI and histopathology, and carefully oriented to precisely correlate imaging and histological sampling locations. This study is the first demonstration of volumetric OFDI with precise correlation to tissue-based diagnosis for evaluating pulmonary pathology. This straightforward, effective technique may be extended to other tissue types to provide precise imaging to histopathological correlation needed to determine fine imaging features of both normal and diseased tissues.

Protocol Text:  (Video demonstrating methods will be performed on swine tissue due to difficulty in scheduling video with availability of surgical tissue)
1) Imaging System 
The technical details of OFDI have been described previously
2) Imaging System Set-up

2.1) Turn on imaging system
2.2) Set and record imaging parameters (rotational speed, pullback speed, image acquisition rate, etc.).
2.3) Attach catheter to rotary junction and pullback device.

2.4) Spin catheter and check for image quality.  Adjust system alignment, and offset as needed.
3) Tissue Preparation

3.1) Place a table top disposable absorbent pad on your benchtop and set lung specimen on pad.

3.2) If imaging a surgical ex vivo specimen from a patient, be sure to consult the pathology department to ensure that all resection margins (bronchial, vascular, and parenchymal margins) have been assessed, documented, and/or removed by a pathologist.

3.3) Identify the bronchial airway entering the resection specimen. Remove any visible mucus within the airway using a syringe bore. If necessary, attach a longer segment of plastic tubing to the syringe bore to suction deeper within the airway.

3.4) Palpate the exterior surface of the specimen to identify the lesion of interest.

3.5) Using a fine metal probe, gingerly navigate through the bronchial tree until near the lesion of interest.

3.6) Open the airway along the probe until the lesion of interest is visible or palpable under the airway mucosa.

3.7) Carefully remove any blood or mucus from the airway mucosa with a cotton-tipped applicator.

3.8) Place imaging catheter parallel to the length of the airway mucosa and obtain an image to confirm lesion is underlying the airway mucosa and to identify a high-quality imaging region of interest for histopathological correlation.

4) Tissue Marking
4.1)  Select region of interest on airway based on previous imaging findings in Step 3.8. 
4.2) Choose two points on the tissue along the desired line of imaging. Points may be parallel to either the longitudinal (Fig. 2) or circumferential (Fig. 3a) aspect of the airway, depending on desired results). Space dots no more than 1.5 cm apart so that the portion of tissue may fit in one histology block for processing.
4.3) Dip a fine tipped open bore needle (i.e. 25 gauge 7/8" long) into the tissue marking dye. 
4.4) Carefully wipe excess ink off the outside of the needle with gauze, leaving tissue marking ink only within the needle bore.

4.5) Puncture the tissue perpendicular to the airway mucosa at the chosen point along the line of imaging. 

4.6) Repeat steps 3.3 to 3.5 for second point on airway mucosa.  

4.7) If the ink runs over the mucosal surface away from the puncture site, use a cotton tipped applicator to carefully remove the excess ink. 
4.8) Remove mucus or blood on the surface of the airway mucosa with a cotton tipped applicator. 
4.9) If the ink dots are placed circumferentially within an airway, it is useful to pin open the two sides of the airway to flatten the tissue in the imaging field (Fig. 3a).

5) Imaging Tissue
5.1) Place the catheter optics over each ink mark and image to ensure the marks are visible on OFDI. Marks should appear as focal disruptions within the tissue structure with overlying highly scattering particles and underlying rapid signal attenuation, which corresponds to the ink within the puncture site (Fig. 3b).
5.2) If ink mark(s) are not visible on OFDI, repeat steps 4.3 to 4.7 for those marks. If ink marks are visible with OFDI, proceed to step 5.3.

5.3) Place the catheter parallel to the two ink marks on the airway mucosal surface such that the catheter optics overlie the tissue beyond the first ink mark (Fig. 2b). Anchoring the proximal end of the catheter with a light weight object and securing the distal end can help reduce motion artifacts.
5.4) Proceed with collecting OFDI pullback.
5.5) View OFDI pullback images to ensure both marks are visible in imaging and to check for motion artifacts (Fig. 3b). If they are not visible, repeat steps 5.3 to 5.4.
6) Collecting and Processing Tissue

6.1) Place green ink dot on airway mucosal tissue to mark the beginning of the imaging scan, 3 mm away from the black ink mark that appeared first in the imaging pullback (Fig. 2c).

6.2) Remove portion of airway containing the two dots by cutting parallel to the line between dots on both sides. Tissue should include green ink mark as well as both black ink dots. Tissue thickness perpendicular to the dots should be no more than 0.3 cm and tissue depth should be no more than 1 cm.
6.3) Place tissue in a histology processing cassette and fix in 10% formalin for at least 48 hours. 

6.4) Process tissue in a tissue processor, available through any histology department.
6.5) Embed tissue in paraffin such that the cut sections will be parallel to the two black ink marks on the airway surface. 

6.6) Use a tissue microtome to face paraffin block until either an ink mark is visible or the entire tissue section is visible, whichever comes first.

6.7) Once both black ink marks are visible, cut one 5 micron section and mount onto a glass slide.

6.8) Continue to cut and mount 5 micron sections every 50 microns until the black ink marks are no longer visible or the tissue ends, whichever comes first.

6.9) Follow standard hematoxylin and eosin staining protocol to stain and coverslip slides.
7) Image Processing

If images were acquired with a benchtop scanner, or other scanning technique where both ink marks were visible in a single cross-sectional image, then the image can be directly correlated with corresponding histology. If volumetric datasets were acquired with a helical scanning catheter, the images will need to be re-interpolated so that a single 2D image bisects both ink marks for correlation with histology. This can be accomplished using ImageJ or other image processing software.
Representative Results: The black ink marks should be between 1 – 1.5 cm apart to indicate the imaging region of interest. The green ink mark should be placed at the beginning of the imaging scan, before the first black ink mark to orient the specimen (Fig. 2 and Fig. 3a). Tissue ink marks should be visible on both OFDI imaging and histopathology (Fig. 3).  

Discussion:  Assessment of early lung malignancies can be extremely challenging due to lack of symptoms and the inability to visualize early neoplastic changes radiologically or bronchoscopically. OFDI provides near histologic resolution, large area 3-dimensional views of tissue microstructure in real time
The tissue dye used to mark the imaging region of interest is clearly visible in both OFDI and histopathology. By using straightforward techniques to orient the tissue in both imaging and histopathology, ink marks in both OFDI and histology may be correlated to allow one to one comparisons of OFDI features and histopathological findings to determine the identifiable imaging characteristics of the tissue pathology. The technique is inexpensive and practical, thus making it useful in many OFDI applications. 

This study is the first demonstration of volumetric OFDI with precise correlation to tissue-based diagnosis for evaluating pulmonary pathology. This technique is presented as an application to pulmonary resection specimens, but can be applied to almost any tissue type to provide the precise imaging to histopathological correlation needed to determine fine imaging features of both normal and pathological tissues. 

Acknowledgments: The authors would like to thank Mr. Sven Holder and Mr Stephen Conley for their invaluable assistance in this study. This work was funded in part by the National Institute of Heath [Grant number R00CA134920] and CIMIT.

Disclosures: We have nothing to disclose.  

Table of specific reagents and equipment:

	Name of the reagent
	Company
	Catalogue number
	Comments (optional)

	Tissue marking dye
	Triangle Biomedical
	TMD-BK, TMD-G
	


References:  

1.
Jemal, A. et al. Cancer statistics, 2007. CA Cancer J Clin 57, 43-66 (2007).

2.
Fujimoto, J. G. et al. Optical biopsy and imaging using optical coherence tomography. Nat Med 1, 970-2 (1995).

3.
Tearney, G. J. et al. In vivo endoscopic optical biopsy with optical coherence tomography. Science 276, 2037-9 (1997).

4.
Yun, S., Tearney, G., Bouma, B., Park, B. & de Boer, J. High-speed spectral-domain optical coherence tomography at 1.3 mum wavelength. Opt Express 11, 3598-604 (2003).

5.
Yun, S., Tearney, G., de Boer, J. & Bouma, B. Removing the depth-degeneracy in optical frequency domain imaging with frequency shifting. Opt Express 12, 4822-8 (2004).

6.
Yun, S., Tearney, G., de Boer, J., Iftimia, N. & Bouma, B. High-speed optical frequency-domain imaging. Opt Express 11, 2953-63 (2003).

7.
Yun, S. H. et al. Comprehensive volumetric optical microscopy in vivo. Nat Med 12, 1429-33 (2006).

8.
Tearney, G. J. et al. Three-dimensional coronary artery microscopy by intracoronary optical frequency domain imaging. JACC Cardiovasc Imaging 1, 752-61 (2008).

9.
Suter, M. J. et al. Image-guided biopsy in the esophagus through comprehensive optical frequency domain imaging and laser marking: a study in living swine. Gastrointest Endosc 71, 346-53 (2010).

10.
Suter, M. J. et al. Comprehensive microscopy of the esophagus in human patients with optical frequency domain imaging. Gastrointest Endosc 68, 745-53 (2008).

11.
Vakoc, B. J. et al. Comprehensive esophageal microscopy by using optical frequency-domain imaging (with video). Gastrointest Endosc 65, 898-905 (2007).

12.
Lam, S. et al. In vivo optical coherence tomography imaging of preinvasive bronchial lesions. Clin Cancer Res 14, 2006-11 (2008).

13.
Michel, R. G., Kinasewitz, G. T., Fung, K. M. & Keddissi, J. I. Optical coherence tomography as an adjunct to flexible bronchoscopy in the diagnosis of lung cancer: a pilot study. Chest 138, 984-8 (2010).

14.
Williamson, J. P. et al. Using optical coherence tomography to improve diagnostic and therapeutic bronchoscopy. Chest 136, 272-6 (2009).



Figure Legend

Figure 1. OFDI of swine airway. In vivo images obtained from a swine airway under mechanical ventilation. Catheter diameter is 0.8 mm and tick marks represent 500 (m. Although different layers of the airway wall are discernible in the OFDI images, it is difficult to precisely interpret the anatomic correlate of the OFDI signals without directly registered histology. e: epithelium, lp: lamina propria, sm: submucosa, c: cartilage, a: alveolar attachments.

Figure 2. Tissue marking of swine airway. (a) Opened airway with two black ink marks on the luminal surface placed parallel to the longitudinal aspect of the airway, 1.5 cm apart. (b) OFDI catheter placed over two black ink marks to include both marks within the OFDI pullback. (c) Airway with additional green ink mark to orient the specimen as to the beginning of the imaging scan. 

Figure 3. OFDI and histopathology of swine airway demonstrating precise correlation using tissue marking. (a) Opened airway with two black ink marks on the luminal surface placed parallel to the circumferential aspect of the airway. Pins are used to further open the airway.  (b) OFDI of swine airway with both ink marks visible (asterisks). (c) Precisely correlated histology with black ink marks visible on respiratory epithelium (asterisks). Scale bar: 2 mm.  (d) Higher magnification view of OFDI image and (e) corresponding histopathology. e: respiratory epithelium, cl: connective tissue in the lamina propria, sm: smooth muscle in the lamina propria, c: cartilage rings (histological artifact has resulted in artificial separation of the cartilage rings), s: salivary gland tissue, d: salivary duct entering epithelium.  Scale bar: 250 µm  
