JoVE manuscript 4383: Response to reviewers

Reviewer 1:

Major Concerns

The main concern here is that the ‘Discussion’ section could be expanded: the result figures show not the same concentrations of glucose and glutamine as mentioned in the discussion; the changes in concentrations can be predicted, but the deviations from the prediction could be discussed;

The concentration range in the discussion is from the calibration experiment, the results of which are presented in Table 2. Figure 4 corresponds to the fed-batch experiment, which was used to validate the sensor’s ability to quantitatively monitor intracellular glucose and glutamine concentrations. Deviations are discussed in more detail in the second paragraph of the ‘Discussion’ section.

what are potential further opportunities for other metabolites?

The Discussion section has been expanded to include information on the potential uses of FIBS for monitoring other small molecules and metabolites.

Minor Concerns

- can the authors give an indication of how fast (order of magnitudes) the ‘fast response’ of FIBS can be?

We have included a reference to Hu et al (Nat Protoc. 2011 Oct 27;6(11):1818-33. doi: 10.1038/nprot.2011.392), which demonstrate that the response is immediate (within minutes).

- the reference to figure 1 in the long abstract does not seem to match with the context of the figure

The figures have been swapped around

- for miniaturization improving the fluorescence read-out S/N ratio Dittrich and Manz Anal Bioanal Chem (2005) 382: 1771–1782 is a good citation to consider

This citation has been included.

- long abstract, second paragraph, space between ‘interesting option’

Corrected

- hemocytometer/haemocytometer misspelled;

Corrected

- one reference to Figure 4 in the ‘Representative Results’ section is wrong; in the penultimate line of the 2nd paragraph, it should likely be ‘4c’.

Corrected

Reviewer 2:

Major Concerns

1) It is not clear what kind of FRET biosensor they used. It is well known that it is difficult to obtain the cells both expressing CFP and YFP, probably due to their similar DNA sequences. Therefore, transpose systems such as PiggyBac or Tol2 are now utilized to establish mammalian cells expressing CFP-YFP pair biosensors.

It is true another blue fluorescent protein TFP with low homology to GFP (as they cited as reference 8) is now available. In that case, retroviral system is applied to establish TFP-YFP paired FRET biosensors efficiently.

Detailed information on the sensors’ plasmids has been included in the first and second paragraphs of the ‘Representative results’ section on page 7. The plasmids have been modified as now described in the text to optimise FRET efficiency.

2) It is also important to estimate the concentration of biosensor for the analysis, since it affects FRET efficiency in the cell. For this, the vector information, such as enhancer, promoter, and antibiotics for selection, is required to obtain reproducible results.

Vector information and antibiotic concentration for selection are now described in the first and second paragraphs of the ‘Representative results’ section. We have conducted a qRT-PCR analysis on cell extracts to quantify the mRNA levels of the biosensors, as now described in paragraph 3 of the same section.

3) Since cells have auto-fluorescence, the proper control for background subtraction is “cells without biosensor expression”, not “no cells”.

This has now been corrected. No background subtraction is necessary.

Minor Concerns

Name and company of the plate reader and filters should be addressed.

The plate reader is in Table 1 (InfinitePRO 200). The filters have also been included. Information on excitation and emission wavelengths has also been included in the second paragraph of the ‘Representative results’ section on page 8.

Qiagen MAXI prep column has no fluorescence (Fig1 left top).

This was a photo of our glucose FIBS in the presence of glucose that has now been removed so as not to confused readers.