Protocol for Relative Hydrodynamic Assessment of Tri-leaflet Polymer Valves
Sharan Ramaswamy1*, Manuel Salinas1, Rob Carrol2, Karla Landaburo1, Xavier Ryans1, Cynthia

Crespo1, Ailyn Rivero1, Faris Al-Mousily3,4, Curt DeGroff3, Mark Bleiweis3,Hitomi Yamaguchi2
1Tissue Engineered Mechanics, Imaging and Materials Laboratory (TEMIM Lab),

Department of Biomedical Engineering, Florida International University, Miami, FL.

2Department of Mechanical and Aerospace Engineering, University of Florida, Gainesville, FL. 

3College of Medicine, University of Florida, Gainesville, FL. 

4King Faisal Specialty Hospital and Research Center (KFSH&RC), Jeddah, Saudi Arabia.

*Corresponding author: 
Sharan Ramaswamy, Ph.D., F.A.H.A

Florida International University

Department of Biomedical Engineering

College of Engineering and Computing

10555 W. Flagler Street, EC 2612

Miami, FL 33174

Tel: (305)-348-2532

FAX: (305)-348-6954

Email: sramaswa@fiu.edu
Keywords: valve disease; valve replacement; polymer valves; pulse duplicator; modification; tri-leaflet geometries; hydrodynamic studies; relative assessment.
Short Abstract

There has been renewed interest in developing polymer valves.  Here, the objectives are to demonstrate the feasibility of modifying a commercial pulse duplicator to accommodate tri-leaflet geometries and to define a protocol to present polymer valve hydrodynamic data in comparison to native and prosthetic valve data collected under near-identical conditions.
Abstract
Limitations of currently available prosthetic valves, xenografts and homografts has prompted a recent resurgence of developments in the area of tri-lealfet polymer valve prostheses. However, identification of a protocol for initial assement of polymer valve hydrodynamic functionality is paramount during the early stages of the design process.  Traditional in vitro pulse duplicator systems are not configured to accomodate flexible tri-leaflet mateirals; in addition, assessment of polymer valve functionality needs to be made in a relative context to native and prosthetic heart valves under identical test conditions so that variability in measurements from different instruments can be avoided.  Accordingly, we conducted hydrodynamic assessment of i) native (n = 4, mean diameter, D = 20 mm ), ii) bi-leaflet mechanical (n= 2, D = 23 mm) and iii) polymer valves (n = 5, D = 22mm) via the use of a commercially available pulse duplicator system (Vivitro Labs Inc, Victoria, BC) that was modifed to accomodate tri-leaflet valve geometries.  Tri-leaflet silicone valves developed at the University of Florida comprised the polymer valve group.  A mixture in the ratio of 35:65 glycerin to water was used to mimic blood physical properites.  Instaneous flow rate was measured at the interface of the left ventricle and aortic units while pressure was recorded at the ventricular and aortic positions.  Bi-leaflet and native valve data from the literature was used to validate flow and pressure readings.  The following hydrodynamic metrics were reported: forward flow pressure drop, aortic root mean square forward flow rate, aortic closing, leakage and regurgitant volume, transaortic closing, leakage and total energy losses.  Representative results indicated that hydrodynamic metrics from the three valve groups could be successfuly obtained by incorporating a custom-built assembly into a commercially available pulse duplicator system and subsequently, objectively compared to provide insights on functional aspects of polymer valve design.
Introduction

Heart valve disease often results from degenerative valve calcification 
 ADDIN EN.CITE 
1
, rheumatic fever 2, endocarditis 
 ADDIN EN.CITE 
3,4
 or congenital birth defects.  When valve damage occurs, causing stenosis and/or regurgitation valve prolapse and cannot be surgically repaired, the native valve is usually replaced by a prosthetic valve.  Currently available options include mechanical valves (cage-ball valves, tilting disk valves, etc.), homograft, and bioprosthetic valves (porcine and bovine valves).  Mechanical valves are often recommended for younger patients based on their durability; however the patient is required to remain on anticoagulant therapy to prevent thrombotic complications 5.  Homograft and Biological prosthetic valves have been effective choices to avoid blood thinner therapy; however, these valves have elevated risk for fibrosis, calcification, degeneration, and immunogenic complications leading to valve failure 6.  Tissue-engineered valves are being investigated as an emerging technology 
 ADDIN EN.CITE 
7-9
, but much still remains to be uncovered. Alternative durable, biocompatible, prosthetic valves are needed to improve the quality of life of the heart valve disease patients.  Again, this valve design could replace the bioprothesis used in transcatheter valve technology, with transcatheter approaches showing the potential for transforming the treatment of selected patients with heart valve disease 
 ADDIN EN.CITE 
10
.
As stated by current standards, a successful heart valve substitute should have the following performance characteristics: "1) allows forward flow with acceptably small mean pressure difference drop; 2) prevents retrograde flow with acceptably small regurgitation; 3) resists embolization; 4) resists hemolysis; 5) resists thrombus formation; 6) is biocompatible; 7) is compatible with in vivo diagnostic techniques; 8) is deliverable and implantable in the target population; 9) remains fixed once placed; 10) has an acceptable noise level; 11) has reproducible function; 12) maintains its functionality for a reasonable lifetime, consistent with its generic class; 13) maintains it functionality and sterility for a reasonable shelf life prior to implantation." 11.  Some of the shortcomings of existing valve prostheses may potentially be overcome by a polymer valve.  Biocompatible polymers have been considered top candidates based on biostability, anti-hydrolysis, anti-oxidation, and advantageous mechanical properties such as high strength and viscoelasticity. In particular, elastomeric polymers may provide material deformation resembling native valve dynamics.  Elastomers can be tailored to mimic soft tissue properties, they may be the only artificial materials available that are bio-tolerant and that can withstand the coupled, in vivo, fluid-induced, flexural and tensile stresses, yet, move in a manner resembling healthy, native valve motion.  Moreover, elastomers can be mass produced in a variety of sizes, stored with ease, are expected to be cost-effective devices and can be structurally augmented with fibrous reinforcement.  
The concept of the use of polymer materials to assemble a tri-leaflet valve is not new and has been the subject of several research investigations over the last 50 years 12, which were abandoned largely due to limited valve durability.  However, with the advent of novel manufacturing methodologies 13,14, the reinforcement of polymer materials 
 ADDIN EN.CITE 
15,16
 and potentially seamless integration of polymer valve substitutes with transcatheter valve technology, there has recently been a renewed interest and activity in developing polymer valves as a potentially viable alternative to currently available commercial valves.  In this light, a protocol for enabling testing of these valves to assess hydrodynamic functionality is the first step in the evaluation process; yet commercially available pulse simulator systems generally do not come equipped to accommodate tri-leaflet valve designs and contain an annular spacing to insert commercially available heart valves (e.g. tilting disc, bi-leaflet mechanical heart valves).  Secondly, polymer valves are an emerging technology whose hydrodynamics can only be assessed in a relative context.  Even though native heart valve pressure and flow data is available, it is important to conduct testing of native aortic porcine valves, which are biologically similar to human valves, using the same pulsatile simulator that is used to evaluate the polymer valves so as to account for measurement differences that may be system dependent.  Thus, the goal of this study was to demonstrate how a commercially available pulse simulator can be fitted with an assembly to accommodate tri-leaflet valve constructs and to systematically evaluate polymer valve hydrodynamic metrics in a relative context in comparison to mechanical and native porcine heart valve counterparts.  In our case, novel tri-leaflet silicone polymer valves previously developed at the University of Florida 13 comprised the polymer valve group.
Protocol
1. Preparation
1.1 Design and fabricate an assembly to accommodate a tri-leaflet valve geometry. This will 
       at minimum include a valve holder to suture-in the valve leaflets, a tube to house the 
       valve holder and surrounding accessories to secure the assembly onto the pulse 
       duplicator system.  In our case, we utilized a commercially available pulse duplicator 
       system available from ViVitro Labs Inc. (Victoria, BC). Valve holder design as well as pre 
       and post assembly configurations are depicted in Fig. 1.
1.2 The entire loop will need to be primed prior to usage.  This involves two steps: i) 
       cleansing of the entire loop system using soap solution and water, including 
       replacement of any degraded tubes prior to use and ii) calibration of instruments 
       connected to the loop, namely the pump being used, the flow probe and the pressure 
       transducers (generally measured at atrial, aortic and ventricular locations). Calibration 
       can initially be performed using 1% saline solution and should be repeated prior to using 
       blood-analog glycerin solution.
2. Native aortic valve dissection

2.1 Obtain 4 fresh pig hearts with the aorta intact from a USDA approved slaughter 
       house (Institutional Animal Care and Use Committee (IACUC) approval may be 
       required).  In our case, our dissection protocol was approved by the IACUC at Florida 
       International University (Protocol Approval Number: 11-020).  Rinse the heart with 
      deionized water and place it in a receptacle filled with the 1% antimycotic/antibiotic and 
      sterile phosphate buffered saline (PBS) solution and transport on ice to the 
      hydrodynamic testing laboratory.

2.2 Place hearts in a dissecting pan and carefully remove the pericardium.  Position the   

       heart such that ventral side is facing you. Visually inspect and identify the four  

       chambers of the heart and locate the aortic arch on the intact aorta.
2.3 Separate the heart into two halves by cutting across horizontally at approximately ¾” below the annulus, i.e., the junction between the aorta and the left ventricle.  Carefully isolate the intact aorta still attached to the left ventricular tissue segment.
2.4 Examine the aortic valve located in the aortic root, the region between the ascending aorta and the lower annulus, ensuring that there is no damage or any signs of calcification.
2.5 Split the aorta at ~ 1” above the annulus and separate the left ventricular tissue segment below the annulus to isolate the aortic valve (Fig.2a).

3. Polymer and Native Valve Suturing process

       3.1 Place the heart valve inside the valve holder such that the base of each valve aligns    

with the base of the post holder. Secure the valve in place at each post temporarily with 

a paper clip, but be careful not to damage the commissures or the cusps.

       3.2. Insert the suture in the needle.  Begin suturing at the bottom of the valve holder by 
passing the needle through the first hole, from the outside to the inside such that the    

needle may be easily pulled from the bottom.  In a looping fashion, start suturing the valve vertically up the posts of the valve holder.

3.3 Progress with suturing (Fig. 2b) along the circumference of the holder and secure with 

additional suture around the tips of the holder posts.  Paper clips (Fig. 2c) can be removed when the valve is completely secured using sutures to the 3 posts and at the circumference of the valve holder (Figs. 2d and 2e). 

4. Hydrodynamic evaluation (note actual protocol will vary depending on specific pulse duplicator system being used; all information contained herein used the Vivitro Pulse duplicator system, (Vivitro Labs Inc., Vancouver, BC).
4.1 Bi-leaflet valve

4.1.1 Set heart rate of pulse duplicator system to 70 beats/min.
4.1.2 Select a flow waveform to drive the pump (in the case of the Vivitro system the S35 waveform was chosen for all hydrodynamic tests).  The specific waveform utilized in our experiments is illustrated by Lim et al (2001) 17.
4.1.3 Turn on amplifier and piston pump. Warm up for 15 min.
4.1.4 Place bi-leaflet valve (Fig. 2f) in the aortic position.
4.1.5 Smear vacuum grease on all junctions of the device where leaks could occur.
4.1.6 Pour glycerin/saline liquid in the atrial compartment.  Note that the pulsatile duplicator system runs on 2 L of liquid with: 35%/0.7 L Glycerin and 65%/1.3 L of Saline solution.  The saline solution is prepared using common salt well-dissolved in deionized water at a concentration of 9 mg/ml (weight/volume).
4.1.7 Turn on the flow transducer that has been placed in the aortic position.
4.1.8 Calibrate the pump.

4.1.9 Proceed with the flow transducer calibration followed by the pressure transducers. Similarly to the pump, simply follow the instructions given by the Vivitest software (Vivitro Labs Inc.) for each flow and pressure under the calibrate tab.

4.1.10 Once calibration is complete, start the pump at a low rpm until the fluid fills the aortic compartment.   Check for leaks. Use additional vacuum grease if necessary.
4.1.11 Turn the two stop-cocks (aortic and ventricular transducers) to open position.
4.1.12 Increase the rpm of the pump until the stroke volume reaches 80 ml/beat.
4.1.13 Permit the system to run for 10 minutes until flow has stabilized. Flow stabilization can be verified by observing the flow and pressure waveforms displayed in the screen. Low to none variation between cycles is a good indicator of system stabilization. 

4.1.14 In the Vivitest software select acquire mode.
4.1.15 Click on collect 10 cycles.
4.1.16 From the analyze mode, click on table and save the file.  Also save an image of the waveforms using the photo-snap option in Vivitest.
      4.2 Native and Polymer valves

4.2.1 For polymer and animal valve, follow the same steps 3.1.1 to 3.1.3 from the 
          bi-leaflet valve instructions.

4.2.2 Place the valve holder with the sutured valve inside the glass tube from the  
          custom made assembly. Sandwich the tube with the top and bottom pieces   

                       and secure in-place with lateral screws and nuts.

4.2.3 Place assembly between the aortic chamber and the original aortic valve 
 

          holder.

4.2.4 Continue with steps 3.1.5 to 3.1.16 from the bi-leaflet valve instructions.
5. Post Processing
      5.1 Flow and Pressure Waveforms
5.1.1 Average the data collected for each of the waveforms collected, i.e., aortic 

                       pressure (AP), ventricular pressure (VP) and flow rate (Q).


5.1.2 For each group of valve (polymer, porcine native aortic valve and bi-
                                     leaflet), plot the corresponding AP, VP and Q versus time relationships on  

                                     the same plot.

5.1.3 For the AP, superimpose normal, native aortic valve 18 and bi-leaflet  

          prosthetic valve 19 plots from the literature for validation purposes.

      5.2 Hydrodynamic metrics
5.2.1 For each valve tested, the following hydrodynamic metrics should be 
computed:  a) Forward flow pressure drop and maximum transvalvular pressure (TVP), b) the aortic root mean square   (RMS) forward flow rate, c) aortic forward flow, closing, leakage and total regurgitant volume, d) valve end orifice area (EOA), e) transaortic forward flow, closing, leakage and total energy losses.
5.2.1.1 Forward flow pressure drop is computed from TVP readings and can be 
             categorized into 3 time intervals, P: interval that starts and ends with 0 
             TVP, F: interval with forward flow and H: interval starting with 0 TVP and 
             ending with 0 flow.  Maximum TVP is the maximum pressure gradient 

             recorded across the valve from the aortic and ventricular pressure 

             readings.
5.2.1.2 The RMS forward flow rate (Qrms) provides a useful metric for quantifying 
              the magnitude of forward flow rate as follows: 
[image: image1.png]1
Q7+ QT 07 0





(1)
Where ‘n’ is the total number of time points collected, ‘Qi’ is the     instantaneous flow rate measurement collected in order ‘i’.
5.2.1.3
The aortic forward, closing and leakage volumes are computed based on    the following time intervals, Forward:  beginning of forward flow through the valve (to), to the end of forward flow (t1); Closing: from t1 till the instance of valve closure (t2); Leakage: from t2 till the end of the cardiac cycle (t3).  Total regurgitant volume is simply the sum of closing and leakage volumes.
5.2.1.4 The EOA based on blood properties can be computed for the 3 intervals, P, F and H from the mean TVP during each of these periods as20: 
[image: image3.png]= rms
EOA = 51.6Vmean TVP









    (2)


5.2.1.5 Energy losses are defined as follows21:
[image: image5.png]Promvarariow = Iy, Quarve AP dt





(3)
[image: image7.png]Perosing riow = Jo Quatoe - AP dt






(4)
[image: image9.png]Breakage riow = Ju: Quatve - AP dt






(5)
Representative Results
Representative flow and pressure waveforms are shown in Figs 3, 4 and 5.  The plots were averaged over the sample size of valves tested for each group, which was, n = 5, 4 and 2 valves for polymer, native porcine and bi-leaflet groups respectively.  The mean hydrodynamic metrics and the standard error of the mean for these sample sizes are presented in Table 1. 
Tables and Figures
Figure 1: (a) Schematic of the Vivitro pulse duplicator system showing the primary components that implement a Windkessel model for physiologically relevant flows (figure presented here with permission from Vivitro Systems, Inc, BC, Canada). (b) Rapid prototyped valve holder configuration to suture and secure silicone or native porcine valves in-place.  (c) (b) Modification of the Vivitro pulsatile loop to accommodate tri-leaflet valve constructs;

Figure 2: (a) Native porcine valve. (b) Top view of polymer valve leaflets.  (c) Side view of polymer valve after suturing and securing in-place within valve-holder. (d) Saint Jude Bi-leaflet mechanical valve.

Figure 3: Mean instantaneous flow rates of the 3 valves tested (n = 5, 4 and 2 valves for polymer, native porcine and bi-leaflet respectively).  Flow rate was measured using an electromagnetic flow meter connected to a noninvasive flow probe placed at the interface location of the ventricle and aortic chambers (see Fig. 1a).

Figure 4:  Mean instantaneous ventricular pressure of the 3 valves tested (n = 5, 4 and 2 valves for polymer, native porcine and bi-leaflet respectively).  Ventricular pressure was measured in the ventricle chamber using a micro-tip pressure transducer.  Superimposed literature ventricular pressure values for native and bi-leaflet valves (Diameter: 29 mm) were obtained from 18 and 19 respectively.

Figure 5: Mean instantaneous aortic pressure of the 3 valves tested (n = 5, 4 and 2 valves for polymer, native porcine and bi-leaflet respectively).  Aortic pressure was measured just downstream from the aortic valve position using a micro-tip pressure transducer.  Superimposed literature aortic pressure literature values for native and bi-leaflet (Diameter: 29 mm) valves were obtained from 18 and 19 respectively.

Table 1:  Mean and Standard Error of the Mean (SEM) Hydrodynamic metrics computed for the heart valves tested (n = 5, 4 and 2 valves for polymer, native porcine and bi-leaflet respectively).  The following intervals should be noted: P: interval that starts and ends with 0 TVP, F: interval with forward flow and H: interval starting with 0 TVP and ending with 0 flow.  Mean diameters of the valves were as follows: Polymer valve (n=5): 22 mm; Native porcine valve (n=4): 20 mm; Bi-leaflet (n=2): 23mm.  Small sample size for bi-leaflet valve was due to limited samples available for research use; the two bi-leaflet valves tested were previously donated to the Biomedical Engineering Department at Florida International University by Saint Jude Medical (Saint Paul, MN).
Table of Materials: Table of specific reagents and equipment used in this protocol
Discussion

In this study, we have demonstrated the utility of modifying a commercially available pulsatile duplicator unit to accommodate tri-leaflet valve geometries so that hydrodynamic testing of polymer and native porcine valves can be performed.  Specifically in our case, the system modified was a Vivitro left heart and systemic simulator system (Fig. 1a) controlled via the ViViTest data acquisition system (Vivitro Systems, Inc, Victoria, BC, Canada).  However, the system is not unlike several in vitro, pulsatile flow loops which all utilize a 2-component Windkessel model to mimic flow and pressure waveforms of relevance to the human circulation 
 ADDIN EN.CITE 
22-25
.   These 2-component Windkessel systems typically consist of a pulsatile pump, a compliance chamber that mimics the distensibility of the arteries, and a peripheral resistance controller that can be used to regulate the vascular resistance.  The equation that describes the 2-component Windkessel model is:






[image: image11.png]w=em -

at









(6)

where [image: image13.png]


 is the compliance, [image: image15.png]


 the resistance, [image: image17.png]Q(®)



 is the volumetric flow rate as a function of time and P is the arterial pressure (i.e., either in the pulmonary artery or aorta).  In this context, we believe that a similar modification can be made to accommodate tri-leaflet valves in other pulsatile simulators as well.  Specifically in our case, to house a tri-leaflet valve structure in the aortic valve location, an assembly primarily of acrylic plastic (plexiglass) casing that contained a rapid prototype valve holder and sutured tri-leaflet valve (Figs. 1b and 1c)  could be easily integrated and removed from the primary Vivitro system. Hydrodynamic testing was subsequently conducted similar to other studies performed by Baldwin et al. 26 and Wang et al. 27.  Instantaneous flow rate was measured using an electromagnetic flowmeter system (Fig. 3).  Real-time measurement of pressure was recorded at the ventricular and conduit location using microtip transducers at a set heart rate of 70 beats/min (Figs. 4 and 5).    The testing fluid was a blood-analog liquid, comprising, deionized water to glycerin in a 65% to 35% ratio and 9 g/l of NaCl, mimicking blood viscosity (~ 3.3 cP).

We initially tested a mechanical bi-leaflet valve and the obtained mean pressure wave forms were compared to literature values 19.  Some ventricular pressure variability was observed possibly owing to different pump mechanisms in place to drive fluid flow as well as geometry and specific settings of different pulse duplicator systems such as size of the ventricular chamber, specific valve mimicking the mitral valve location, heart rate chosen, physiological flow waveform selected, etc.  On the other hand, the aortic waveforms were found to be very similar and system-independent.  This exercise was repeated for native porcine valves and again, larger variability in ventricular pressure was observed when comparing our results to the literature 18.  However, it is important to note that within our system, instantaneous flow rates as well as both ventricular and aortic pressures were similar regardless of the valve that was tested, i.e., polymer and native with assembly or bi-leaflet without assembly.  This exercise is important to perform because one needs to ensure that modifications to the duplicator system with an assembly do not considerably alter local flow and/or pressure conditions. Secondly, these results indicate that as a means of system validation, at minimum, comparable aortic pressures need to be derived across pulse duplicator platforms or the valve being tested.  The interpretation of the hydrodynamic variables themselves is a matter of individual polymer valve design specifics. Standards such as ISO (International organization for standardization) 5840 used in the evaluation of cardiac valve prostheses can serve as a guide to assess various parameters associated with the polymer valve geometry, manufacturing and material properties.  These parameters can be further optimized and hydrodynamic testing subsequently re-visited to ensure that the standards needed for FDA submission are met.
For example, in our polymer valves, comparable energy losses and lower regurgitant volumes versus native and bi-leaflet valves suggested acceptable workloads on the left ventricle 21 and efficient valve closure (Table 1).  However, the closing dynamics resulted in a relatively higher polymer valve maximum TVP gradient (versus bi-leaflet valves) which in our case, warrants further mechanical evaluation of silicone material being used to fabricate the valves to ensure that the higher stress does not cause leaflet rupture, and that a sufficient factor of safety can be put in place.  In conclusion, we have demonstrated that an assembly consisting of a housing unit, glass tube and a valve holder can be fabricated to accommodate tri-leaflet structures such as polymer valves which can be sutured in-position.  Comparative flow and pressure waveforms across native, prosthetic and polymer valve that is being developed needs to be obtained.  Secondly the pressure waveforms need to be validated with literature values.  A limitation of our approach is that ventricular waveforms are pulse duplicator system specific and are likely to show differences; however aortic pressure waveforms should be comparable across platforms or valve being tested if sufficient valve functionality exists.  A future direction of this work is to further optimize the polymer valve material, manufacturing process and geometry.  Hydrodynamics tests will subsequently be repeated under identical conditions so as to determine if functional improvements are quantitatively observed by comparing the current and previous hydrodynamic metrics computed.
Disclosures
The authors have nothing to disclose.

Acknowledgments

A seed grant from the University of Florida - College of Medicine is gratefully acknowledged.  Graduate studies (Manuel Salinas) were supported through a minority opportunities in biomedical research programs - research initiative for scientific enhancement (MBRS-RISE) fellowship: NIH/NIGMS R25 GM061347.  Financial support from the Wallace H. Coulter Foundation thru Florida International University’s, Biomedical Engineering Department is also gratefully acknowledged.  Finally, the authors thank the following students for their assistance during various stages of the experimental process:  Kamau Pier, Malachi Suttle, Kendall Armstrong and Abraham Alfonso.
References
1
Rajamannan, N. M. et al. Calcific aortic valve disease: not simply a degenerative process: A review and agenda for research from the National Heart and Lung and Blood Institute Aortic Stenosis Working Group. Executive summary: Calcific aortic valve disease-2011 update. Circulation 124, 1783-1791, doi:124/16/1783 [pii] 10.1161/CIRCULATIONAHA.110.006767 (2011).

2
Marijon, E., Mirabel, M., Celermajer, D. S. & Jouven, X. Rheumatic heart disease. Lancet 379, 953-964, doi:S0140-6736(11)61171-9 [pii] 10.1016/S0140-6736(11)61171-9.

3
Karaci, A. R. et al. Surgical treatment of infective valve endocarditis in children with congenital heart disease. J Card Surg 27, 93-98, doi:10.1111/j.1540-8191.2011.01339.x (2012).

4
Knirsch, W. & Nadal, D. Infective endocarditis in congenital heart disease. Eur J Pediatr 170, 1111-1127, doi:10.1007/s00431-011-1520-8 (2011).

5
Korossis, S. A., Fisher, J. & Ingham, E. Cardiac valve replacement: a bioengineering approach. Biomed Mater Eng 10, 83-124 (2000).

6
Ghanbari, H. et al. Polymeric heart valves: new materials, emerging hopes. Trends Biotechnol 27, 359-367, doi:S0167-7799(09)00076-6 [pii] 10.1016/j.tibtech.2009.03.002 (2009).

7
Mol, A., Smits, A. I., Bouten, C. V. & Baaijens, F. P. Tissue engineering of heart valves: advances and current challenges. Expert Rev Med Devices 6, 259-275, doi:10.1586/erd.09.12 (2009).

8
Ramaswamy, S. et al. The role of organ level conditioning on the promotion of engineered heart valve tissue development in-vitro using mesenchymal stem cells. Biomaterials 31, 1114-1125, doi:S0142-9612(09)01106-5 [pii] 10.1016/j.biomaterials.2009.10.019 (2010).

9
Sacks, M. S., Schoen, F. J. & Mayer, J. E. Bioengineering challenges for heart valve tissue engineering. Annu Rev Biomed Eng 11, 289-313, doi:10.1146/annurev-bioeng-061008-124903 (2009).

10
Zamorano, J. L. et al. EAE/ASE recommendations for the use of echocardiography in new transcatheter interventions for valvular heart disease. J Am Soc Echocardiogr 24, 937-965, doi:S0894-7317(11)00507-4 [pii] 10.1016/j.echo.2011.07.003.

11
ANSI/AAMI/ISO. Cardiovascular Implants - Cardiac Valve Prostheses. Association for the Advancement of Medical Instrumentation 71 (2005).

12
Gallocher, S. L. Durability Assessment of Polymer Trileaflet Heart Valves PhD thesis, Florida International University, FIU Electronic Theses and Dissertations, (2007).
13
Carroll, R. et al. in UF Pediatrics Science Days conference.

14
Pierre, K. K. et al. in Biomedical Engineering Society, Annual Fall Meeting.

15
Cacciola, G., Peters, G. W. & Schreurs, P. J. A three-dimensional mechanical analysis of a stentless fibre-reinforced aortic valve prosthesis. J Biomech 33, 521-530, doi:S0021-9290(99)00222-5 [pii] (2000).

16
De Hart, J., Cacciola, G., Schreurs, P. J. & Peters, G. W. A three-dimensional analysis of a fibre-reinforced aortic valve prosthesis. J Biomech 31, 629-638, doi:S0021929098000633 [pii] (1998).

17
Lim, W. L., Chew, Y. T., Chew, T. C. & Low, H. T. Pulsatile flow studies of a porcine bioprosthetic aortic valve in vitro: PIV measurements and shear-induced blood damage. J Biomech 34, 1417-1427, doi:S0021929001001324 [pii] (2001).

18
Gutierrez, C. & Blanchard, D. G. Diastolic heart failure: challenges of diagnosis and treatment. Am Fam Physician 69, 2609-2616 (2004).

19
Shi, Y., Yeo, T. J., Zhao, Y. & Hwang, N. H. Particle image velocimetry study of pulsatile flow in bi-leaflet mechanical heart valves with image compensation method. J Biol Phys 32, 531-551, doi:10.1007/s10867-007-9035-2 (2006).

20
Chandran, K. B., Yoganathan, A. P. & Rittgers, S. E. Biofluid Mechanics: The Human Circulation. 1st edn,  277-314 (CRC Press, Taylor & Francis Group, 2007).

21
Akins, C. W., Travis, B. & Yoganathan, A. P. Energy loss for evaluating heart valve performance. J Thorac Cardiovasc Surg 136, 820-833, doi:S0022-5223(08)00465-0 [pii]

10.1016/j.jtcvs.2007.12.059 (2008).

22
Fung, Y. C. Biomechanics: Circulation. 2nd ed.,  (Springer, 1997).

23
Keener, J. & Sneyd, J. Mathematical Physiology, II: Systems Physiology. 2nd ed.,  (Springer, 1998).

24
Quick, C. M., Berger, D. S. & Noordergraaf, A. Apparent arterial compliance. Am J Physiol 274, H1393-1403 (1998).

25
Wang, Q., Jaramillo, F., Kato, Y., Pinchuk, L. & Schoephoerster, R. T. Hydrodynamic Evaluation of a Minimally Invasive Heart Valve in an Isolated Aortic Root Using a Modified In Vitro Model. Journal of Medical Devices 3, 011002-011001 - 011002-011006 (2009).

26
Baldwin, J. T., Campbell, A., Luck, C., Ogilvie, W. & Sauter, J. Fluid dynamics of the CarboMedics kinetic bileaflet prosthetic heart valve. Eur J Cardiothorac Surg 11, 287-292, doi:S1010-7940(96)01034-2 [pii] (1997).

27
Wang, Q., Jaramillo, F., Kato, Y., Pinchuk, L. & Schoephoerster, R. T. Hydrodynamic Evaluation of a Minimally Invasive Heart Valve in an Isolated Aortic Root Using a Modified In Vitro Model. Journal of Medical Devices 3, 011002-011006 (2009).



