Submission ID #: 50340
Editor Name: Michael Linnes
Videographer name: Craig Scorgie
Film Date: Jan. 8, 2013
Authors and Affiliations: Mansik Jeon1, Jeehyun Kim2, Chulhong Kim1
1 Department of Biomedical Engineering, The State University of New York at Buffalo. 2School of Electrical Engineering and Computer Science, Kyungpook National University
Title: Photoacoustic Cystography
Corresponding Author: 

Chulhong Kim, Ph.D.

State University of New York at Buffalo

Buffalo, NY UNITED STATES
chulhong@buffalo.edu

Additional Contacts:

mansikje@buffalo.edu
jeehk@knu.ac.kr
A.  Will you require JoVE to record video microscopy, such as filming a complex dissection or microinjection technique? (Y/N) _____N____ If yes, please list make and model of your microscope: ______________________________
B.   Does your protocol include detailed, step-by-step, descriptions of software usage? (Y/N)____N____ 
C.  Which steps of your protocol will viewers benefit most from having filmed? Please list 4-6 steps_______________4___________
D.  What is the single most difficult aspect of this procedure and what do you do to ensure success?  ____Proper injection of contrast agents__________
1. Introduction (Schematic Overview and Interview)
A. Schematic Overview (read by voice talent at JoVE): 
Procedural Narrative:
The overall goal of this procedure is to demonstrate in vivo nonionizing photoacoustic cystography using near-infrared optical absorbents as nontoxic optical-turbid tracers. (Intro)
This is accomplished by first anesthetizing the experimental animal. (P1: Show slide 2 of Presentation-Script-11232012.pptx)
The second step is to prepare the animal for imaging by removing the hair in the abdominal region, positioning the rat atop of a custom-made animal holder, inserting a catheter, and apply ultrasound gel to the imaging region. (P2: Show slide 3 of Presentation-Script-11232012.pptx (*it is animated to point out the different parts of the statement.))
Next, images of the abdominal area are obtained using a photoacoustic cystography system before and after injection of optical absorbents into the bladder. (P3: Beginning with the rat from slide 4, show the animation within the slide of the photoacoustic system scanning the shaved region of the mouse.  With the word “before” go to a clear screen and show the pre-injection image on the left and the post-injection image on the right part of the screen both clearly labeled.)
Ultimately, this process allows for the imaging of a rat bladder using a nonionizing and noninvasive photoacoustic cystography system providing with no radiation exposure and no long-term heavy agent accumulation. (P4: Show slide 5 of Presentation-Script-11232012.pptx beginning with the image alone, then adding the text as it is mentioned. (*it is animated to point out the different parts of the statement.))
B.  Interview: (Said by you on camera. Don’t forget to smile!)  
1.1. Chulhong Kim: Photoacoustic tomography has become a premier biomedical imaging modality because it provides strong optical absorption contrast and high ultrasonic spatial resolution in biological tissues. 
1.2. Mansik Jeon: The key advantages of photoacoustic tomography are that it is completely free from ionizing radiation, the injected agent is not accumulated in the bladder, and it is fast and cost effective. 
Protocol (read by voice talent at JoVE):

2. Configuration of the Deep Reflection Mode Photoacoustic Cystography System
2.1. For this procedure, use a Q-switched neodymium-doped yttrium aluminum garnet laser that has been frequency doubled to generate light at 532 nm.  The laser output is passed through an optical parametric oscillator, or OPO, with a wavelength-tunable range of 680 to 2500 nm. 
2.1.1. MED: Talent points to laser within the setup holds still for a few seconds then moves along the light path to point out the optical parametric oscillator.

2.2. Start by setting the OPO pulse duration to 5 ns for each laser shot and set the repetition rate to 10 Hz. Then, set the optical wavelength to 667 nm for use with the methylene blue contrast agent.  This wavelength may need to be modified if another contrast agent is used.

2.2.1. SCREEN: Talent sets the pulse duration, repetition rate, and finally the optical wavelength as described in the text.
2.3. Deliver the light coming out of the OPO through a right angle prism to a home-made spherical conical lens.   The conical lens is made from a 2.5 cm diameter BK7 lens with a cone angle of 152°. 
2.3.1. MED: Talent points out the light path beginning at the OPO, into the right angle prism, and into the conical lens. Pause at the conical lens for a few seconds

2.3.2. CU: Close up of the conical lens (Lens can be in the setup or shown with the talent holding it.)
2.4. Next, redirect the diverging donut-shaped light beam through a 2.5 cm thick optical condenser made of transparent acrylic… The diameters of the top and bottom surfaces are 6.1 and 4.8 cm, respectively. 
2.4.1. MED: Talent points out the light path beginning at the conical lens and ending at the condenser.
2.4.2. CU: Close up of the condenser lens (Condenser can be in the setup or shown with the talent holding it.)
2.5. Make sure the beam pattern is a perfect ring shape.  If the donut-shaped beam pattern is not properly generated, the photoacoustic signals originating from the skin surface will be dominant. Thus, it will be difficult to achieve deep tissue imaging. 
2.5.1. MED: Talent holds a sheet of paper in front of the condenser and then behind the condenser to show the change in light pattern. 
2.6. Then, prepare a small water container to boost acoustic coupling between the animal and the laser. The bottom opening is wrapped with a clear thin polyethylene film which is optically and acoustically transparent. 
2.6.1. MED: Talent displays the water container for the camera and then wraps it with a clear thin polyethylene film.

2.7. Using a piece of black electrical tape, coaxially align the line-shaped light with the ultrasound focal zone. If these are not well aligned, the system will suffer from low signal-to-noise ratio. 
2.7.1. MED: Talent places the black electrical tape in the focal zone and adjusts the focus of both the light and ultrasound.
2.8. Next, mount a spherically focused ultrasound transducer to detect the photoacoustic waves. For this, a V308 Olympus NDT immersion transducer is used.  It collects data at 5-MHz and has transverse and axial resolutions of 590 and 144 µm, respectively. 
2.8.1. MED: Talent mounts the ultrasound transducer. (TEXT: 25 mm diameter, 19 mm focal length)
2.9. Then, connect the transducer to a broadband ultrasonic pulser/receiver, which will amplify the photoacoustic waves, and connect this to an oscilloscope. 
2.9.1. MED: Talent connects the transducer to the receiver. (TEXT: OlympusNDT, 5072PR, 35MHz bandwidth, 59dB rf gain)
2.9.2. MED: Talent connects the receiver to the oscilloscope. (TEXT: TDS5054, Tektronix)
2.10. Once the system has been set up, obtain one-dimensional time-resolved images, called as A-line images, by measuring the arrival times of the photoacoustic waves through the oscilloscope.
2.10.1. MED Over the Shoulder: Talent works at computer to obtain A-line images.

2.10.2. MED Over the Shoulder: Talent adjusts the oscilloscope as data is incoming.

2.11. Then, mechanically control a linear raster stage using LabVIEW to acquire 2-dimensional, B-Scans, and 3-dimensional photoacoustic images.  The 3-dimensional imaging takes approximately 25 minutes per sample for a field 2.5 cm × 2.4 cm × 1.5 cm in size.

2.11.1. MED Over the Shoulder: Talent uses lab view to begin image acquisition using the mechanical stage.

2.11.2. CU: Mechanical stage moves while acquiring the images.

2.12. To acquire 3-D images, set up the LabVIEW program to control the stage.  Acquire 125 samples in the x-direction with a step size of 0.2 mm, 60 samples along the y-direction with a step size of 0.4 mm, and 500 data points with a 50-MHz sampling rate along the z direction.
2.12.1. MED Over the Shoulder: Talent sets up LabView program and starts the program.
2.12.2. CU: Stage moves based on LabView input.
2.13. Finally, set up a MathWorks MATLAB software system to acquire, covert, and process the raw data into a sequence of cross-sectional B-scan images that represent the volumetric data. The recorded images are then displayed as 2D maximum amplitude projection images.
2.13.1. SCREEN: Talent processes the raw data into 2D max amplitude projection images.
3. Animal Preparation for Imaging Procedures

3.1. To prepare rats for imaging, anesthetize a female Spraque-Dawley rat with a weight between 200 – 250 grams by intraperitoneal injection of a mixture of ketamine and xylazine. 
3.1.1. MED: Talent removes anesthetized rat from cage and places in prepared work area. (TEXT: Ketamine 85 mg/kg, Xylazine 15 mg/kg)

3.2. Then, remove the hairs in a 5 cm square around the abdominal area using hair removal lotion. 
3.2.1. MED: Talent adds hair removal lotion to the rat’s abdomen.

3.2.2. MED: Talent wipes off hair removal lotion and hair.
3.3. Next, position the rat atop a custom-made animal holder. 
3.3.1. CU: Talent places rat into the custom holder.
3.4. Then, prepare a 22-gauge catheter for insertion by coating it with a lubricant and insert the distal end of the catheter horizontally into the urethra until the hub of the catheter reaches the opening and urine begins to void via the catheter.
3.4.1. MED: Talent coats catheter with lubricant

3.4.2. CU: Talent inserts catheter until urine begins to void.
3.5. Finally, position the rat, so that it is on top of the animal holder and below the water container in the photoacoustic cystography system. 
3.5.1. CU: Talent places the rat into the photoacoustic cystography system.
4. In vivo and Ex vivo Photoacoustic Cystography Imaging Procedures
4.1. To begin imaging, fully anesthetize the rat using vaporized isoflurane during the in vivo photoacoustic imaging experiments. 
4.1.1. CU: Talent places vapor mask on the rat.

4.1.2. MED: Talent turns on the gas flow. (TEXT: 1 L/min Oxygen, 0.75% isoflurane) 

4.2. Then, apply ultrasound gel between the animal skin surface and plastic membrane to improve acoustic coupling.
4.2.1. MED: Talent applies ultrasound gel on the animals skin and places the animal under the plastic membrane ensuring good coupling.
4.3. Next, run the prepared MathWorks MATLAB program to obtain a control three-dimensional photoacoustic image prior to injection of contrast agents. 
4.3.1. MED: Talent runs the program on the computer in the background with the experimental setup in the foreground.

4.4. Once the control image is obtained, use a 1-ml-syringe to introduce an aqueous solution of methylene blue to the bladder via the catheter. 
4.4.1. MED: Talent draws up methylene blue solution in syringe. (TEXT: Methylene Blue: 30 mM, 0.8 µl/g body weight)
4.4.2. CU: Talent injects methylene blue via the catheter.
4.5. Following injection of methylene blue, acquire a series of in-vivo photoacustic images every 12 hours until 48 hours post-injection.
4.5.1. MED Over the shoulder: Talent collects data from one of the imaging trials.
4.6. Next, sacrifice the rat by injecting an overdose using 150 mg/kg of pentobarbital. Then, remove the bladder and kidney and place them on a glass plate.  The bladder and kidney from an uninjected animal should be used as a control. 
4.6.1. CU: Talent removes the bladder and kidney from the rat and places them into a glass plate.
4.7. For ex-vivo imaging, position the glass plate below the water container in the photoacoustic cystography system and apply ultrasound gel between the excised organs and plastic membrane to improve acoustic coupling during imaging.
4.7.1. MED: Talent places glass plate below the water container.

4.7.2. CU: Talent adds ultrasound gel between the organs and the membrane.

5. Results: Photoacoustic Cystography Imaging of Rat Bladders
5.1. The image shown here is a 2-dimensional x-y abdominal scan of a rat imaged with 667 nm laser light pre-methylene blue injection. Here the blood vessels can be easily seen as blood absorbs light to generate a photoacustic signal, but there is no visible bladder due to the lack of contrast agent. 
5.1.1. LABMEDIA: Figure 1a (Video Editor: Point out the blood vessels (BV) and the area where the bladder should be located (BD) when they are mentioned in the text. Do not show the dotted line at this point.)
5.2. When imaged 12 minutes after injection of the methylene blue contrast agent, the bladder becomes easily visible with very strong amplitude throughout the bladder. When measured again at 24 and 48 hours, the contrast agent has left the body and is no longer visible. 
5.2.1. LABMEDIA: Figure 1b (Video Editor: Point out the area where the bladder is located (BD) when it is mentioned in the text. Do not show the dotted line at this point.)

5.3. By changing the optical wavelength to 850 nm with the optical parametric oscillator, the methylene blue contrast agent no longer absorbs the energy and becomes invisible. This makes the bladder effectively disappear and acts as another negative control for the system. 
5.3.1. LABMEDIA: Figure 1c (Video Editor: Point out the area where the bladder is located (BD) when it is mentioned in the text.)
5.4. Depth resolved images obtained from the z-plane along the dotted lines shown here provide additional depth data along areas of interest as shown here, or can be performed for the entire image to describe the area in 3-dimensions.
5.4.1. LABMEDIA: Figures 1a, 1b, 1d, and 1e (Video Editor: Begin showing Figure 1a and 1b on the left hand side of the screen without the text labels, but with the dotted lines.  Expand 1d from the dotted line in 1a and 1e from the dotted line in 1b over to the right hand side of the screen.  Show the PA amplitude scale on the bottom center of the screen.

5.4.2. LABMEDIA: Figure 2 (Video Editor: With the word “or”, enlarge figure 1b to take up the left half of the screen and remove all other figures.  Then start with the dotted line on the top of figure 1b and slowly lower it down the image.  At the same time begin to show figure 2 on the right half of the screen unveiling it from top to bottom at the same rate as the dotted line is dropped down figure 1b.  Show the color scale bars below each image from the beginning of the animation and add the text labels to figure 2 once the full image is shown.)

6. Conclusion (said by authors on camera) 
6.1. Mansik Jeon: In this video, we demonstrate the feasibility of a novel and nonionizing process for bladder imaging in-vivo, called photoacoustic cystography. Using a photoacoustic imaging system, we have successfully imaged a rat bladder filled with an injection of optical-opaque tracers in rats. 
6.2. Chulhong Kim: In the future, this technique may also be applied to monitoring urological problems such as vesicoureteral reflux in pediatric patients as the low radiation dose, portability, and inexpensive nature of this procedure are optimal for this type of application. 
( 2012, Journal of Visualized Experiments


