| Time Code (shot number) | Description (plus additional comments) | New voice-over |
| --- | --- | --- |
| 00:00 (4.1) Include as VO | Text overlay not included in this cut because it was read out loud. It can be included in the VO, or it can return to a text overlay – whichever you prefer. | # The Syncardia total artificial heart is the only FDA-approved total artificial heart in clinical use. Total volume displacement is around 400 ml. |
| 00:10 (4.2) | Cut 00:30-00:36 – we don’t use the deairing needles that the camera focuses on | # The device consists of 2 polyurethane ventricle chambers, each with a maximum stroke volume of ~70 ml and 2 tilting disc valves (TEXT: Medtronic Hall, 27 mm inflow, 25 mm outflow) to direct blood inflow and outflow. Each chamber is also pneumatically driven by a separate driveline connected to a driver externally. |
| 00:37 (4.3) | We are missing a shot of the grafts and devices being soaked in rifampin (00:45), so I continued to show footage of the device being sealed. | # To prepare the device for surgery, first seal the Dacron aortic and pulmonary grafts with CoSeal Surgical Sealent and then soak the grafts and devices in rifampin. |
| 00:51 (6.1) | [Table skips Section 5 because it is all lab media and go straight to section 6. There will be a title card at the beginning of this section with the new section title (Implantation)] | # After performing a routine chest and abdomen cardiac surgery prep and a median sternotomy, medially divide the left diaphragm and then dissect underneath the posterior rectus fascia, creating a left upper abdominal peritoneal pocket. |
| 01:11 (6.2) |  | # Next make two 2 cm incisions 5 cm apart below the left costal margin through the rectus fascia and then create intramuscular tunnels for the pneumatic drivelines by incising the skin and fascia and passing a clamp through the rectus muscle into the mediastinum. Maintain the tunnels with two 1 inch Penrose drains. |
| 01:39 (7.1) | (Ann):There is a lot of video to show cannulation of the aorta and superior and inferior vena cava, and I felt it was important to see these steps. We may want to include additional VO detailing the specifics of the steps, and coinciding with the cut video, to avoid this long stretch of video with no audio underneath (although there will be music). Or, we can cut the video down, as you see fit. SURGEON: I think the video could be cut to just showing each of the cannulas being inserted. This part is fairly standard. | Heparin is administered for cardiopulmonary bypass. Minimize unnecessary dissection and mobilization within the pericardium to preserve tissue planes for subsequent heart transplantation Cannulate the aorta. Then, cannulate the superior and inferior vena cava through the right atrium. Flood the surgical field with carbon dioxide to minimize air embolization. |
| 03:38 (7.2) | At 03:56, we are missing the VO that says, “Start the excision of the right ventricle…etc.”, but I included the audio of the surgeon talking. This is where the VO would go. | Cross clamp the aorta. Divide the aortic root and the main pulmonary artery just above the valvar commissures. Start the excision of the right ventricle along the acute margin, 1-2 cm distal and parallel to the atrioventricular groove. Continue the incision anteriorly into the RV outflow tract Divide the intracardiac portion of any defibrillator and/or pacing wires and allow them to retract into the SVC. |
| 04:16 (7.3) |  | Incise the interventricular septum to open the left ventricle. Similarly, extend the left ventricular incision laterally, parallel to the atrioventricular groove and into the LV outflow tract. |
| 04:35 (7.4) | At 04:51, there is a pretty good shot of the aortic outflow tract that the surgeon points out. I included it because I felt it might be good to include VO pointing this out, as the surgeon does. | Then divide and remove the remaining interventricular septum |
| 05:15 (7.5) | There is a good shot of the coronary sinus, and the surgeon is explaining that the next step is to look for a PFO, after which, he points out the coronary sinus. Again, I included this because I thought the shot was good, and we might want to add VO stating essentially what the surgeon says. | Inspect the right atrium through the tricuspid valve, specifically looking for a patent foramen ovale. |
| 05:40 (7.5b) | The VO needs to change, because there was a PFO. I cut it in the order it occurred: Namely, the sewing of the PFO, then the sewing of the coronary sinus. Also, the usage of the stapler to ligate the left atrial appendage came later in the procedure. After oversewing the coronary sinus, the surgical team proceeded to the excision of the mitral and tricuspid valve leaflets. I cut it in the order it occurred, so this VO does not appear here. | If necessary, close any patent foramen ovale with 3-0 Prolene to prevent intra-atrial shunting. Similarly, oversew the orifice of the coronary sinus with 3-0 Prolene. |
| 06:12 (7.6) | The highlighted part of the VO occurs later in the video. | Excise the tricuspid valve leaflets, leaving a several millimeter cuff to the annulus. Similarly, excise the mitral valve leaflets, again leaving several millimeters to the annulus. |
| 06:41 (7.7) | This is the second half of the VO originally meant for 7.5. | This patient was known to have a large atrial thrombus and the left atrial appendage was amputated with a SeamGuard reinforced, endoGIA stapler. |
| 06:54 (7.8) | This is the second half of the VO originally meant for 7.6 | Trim the left and right ventricular cuffs, leaving a 1 cm rim of tissue. Oversew the cuffs with 2-0 Prolene sutures for hemostasis and to reduce the orifice size to match the TAH atrial quick connects. |
| 07:12 (7.9) Agree | The surgical assistant reading aloud from the script says it differently than is listed in the next column. She states, “The total artificial heart atrial quick connects have been trimmed to…” I recommend keeping this change, or something similar to it, since the trimming did not actually happen at that point in the surgical procedure, but rather, at an earlier point. | The left and right TAH atrial quick connects are trimmed to 0.5-1 cm. Invert the atrial quick connects and sew them to their respective left and right ventricular cuffs. |
| 07:45 (7.10) | The surgeon makes a comment at this point that the one common area for leakage is in the corners, and he is uncertain if it was in the script. I included it as it was clear that this warning is important and should be included, with appropriate accompanying VO. | Reinforce the anastomoses with additional 2-0 Prolene. Note that a common area for leak is at the corners of the interventricular septum, which is also reinforced. |
| 08:15 (7.11) |  | Trim the aortic graft quick connect to ~ 2 cml centimeters. Trim the pulmonary graft quick connect to leave it several centimeters longer than the aortic graft, to allow room for the aortic graft to pass below. Sew the pulmonic graft to the pulmonary artery. Then sew the aortic graft to the aorta |
| 09:10 (7.12) |  | Now, line the pericardium to facilitate re-entry and reduce blood loss during subsequent explant for transplantation. Tack a Goretex sheet to the left lateral pericardium prior to connecting the pump, as access to this space afterwards will be limited. |
| 09:35 (7.13) Can probably trim video showing tunneling of drivelines to just one of the drivelines. | The surgical assistant didn’t read the last part of the Vo out loud, but there is plenty of footage to account for it should it wind up in the final script. I kept this cut quite long so you could decide which parts are necessary to show and which aren’t. | Place a strip of Goretex in the transverse sinus, to help protect the right pulmonary artery during re-entry for transplant. Pass a 13 mm Hegar (pronounced “hay-gar”) dilator through the Penrose drains, to dilate the tract and help guide the pneumatic drivelines through the abdominal wall. |
| 10:25 (7.14) Can also cut video showing ligation to just one. | 7.12 actually came before 7.11 in the actual procedure. I cut it in the order it was shot. Also, the last part of this VO was not read aloud, but I put in a text box to coincide with the video that will accompany it in the final script. | Ligate the de-airing nipples on the TAH left and right ventricles. Suction and irrigate the left and right atrium to ensure no residual embolic debris. |
| 10:46 (7.15) | Highlighted VO was not said or done at this time. It occurred later in the procedure. | Use 2 heavy needle drives to grab the pliable quick connects and stretch them over the hard plastic connectors of the artificial ventricles to tightly connect the ventricles to their respective orifices Then, place an aortic root vent. |
| 11:32 (7.16) | This is step 7.11 in the original script, but it occurred here in the actual procedure. | # Replace the original metal connectors of the external console driveline with plastic connectors that will also connect with the portable discharge driver. Connect the large male connector to the left or red BigBlue driveline tubing and the large female connector to the right or blue BigBlue driveline tubing. Reinforce the connections with plastic zip ties. |
| 12:04 (7.17) | Second part of step 7.13, which occurred here. | Initiate low rate, low pressure pumping for de-airing. (TEXT: Left drive pressure 40, right drive pressure 0, % systole 40, rate 40, vacuum 0) Flow through the right side is initially passive. |
| 12:10 (7.18) | Lab media: OR deairing – TEE movie attached | Resume lung ventilation, and then remove the cross clamp Despite the limited visualization provided by TEE, it remains a useful modality to assess de-airing and unobstructed caval and pulmonary venous return to the atria. |
| 12:34 (7.19) | Lab media goes here: Table with early parameters (manuscript 4.13) | # Similar to normal physiology, TAH output is sensitive to changes in preload, afterload, and contractility, or pneumatic ejection. The pneumatic airflow and pressure within the left and right artificial ventricles are displayed graphically on the console. Here, typical early initial TAH parameters are shown: |
| Not in footage (7.20) | # Table from 7.16.1. with left drive pressure highlighted | # The left drive pressure is typically between 180-200 mmHg. Note, that the pneumatic drive pressure is measured at the console and does not reflect the actual blood pressure generated. Once the artificial ventricle has fully ejected, further pressurization of the artificial ventricle is not transmitted to the patient |
| Not in footage (7.21) | Table from 7.16.1. with right drive pressure highlighted | # The right drive pressure is typically between 30-60 mmHg. Patients with biventricular failure often have an elevated pulmonary vascular resistance. While increasing the right drive pressure can overcome the resistance, higher drive pressures increases the risk for significant pulmonary edema. |
| Not in footage (7.22) | Table from 7.16.1. with heart rate, % systole, and vacuum each highlighted as mentioned | # The heart rate should be between 100-120 beats per minute, the percent systole should be 50, and the vacuum should be left at 0 until the chest is sealed to reduce entraining air. |
| 14:05 (7.23) | There was no good shot of the bypass cannulas being removed. However, at 14:21, there is a good shot of the valve being removed. Perhaps Vo could be inserted to accompany that. | Once sealed, increase the vacuum to about 10 to 15 mmHg to improve the ventricular filling and thus the cardiac output. Remove the cannulas and administer protamine to reverse the heparin anticoagulation. |
| 14:38 (7.24) | The second part of this VO happens later. | Once hemostasis has been obtained and the chest is ready to be closed, line the remainder of the pericardium with Goretex Preclude membrane. |
| 14:52 (7.25) | The second part of this VO happens later. Attached is OR TEE videos: OR SVC, OR IVC, OR LUPV, and OR RUPV and RLPV | The TAH can compress the vena cava or pulmonary veins and obstruct inflow. This is particularly true of the IVC and left pulmonary veins. If necessary, place a heavy suture around the artificial right ventricle and secure it to the left costal margin. |
| 15:05 (7.26) | There is no VO for this, but I included it because it’s a good shot of hemostatic sealant being applied. We might want to consider adding VO. | Reinforce anastomotic connections with CoSeal. |
| 15:16 (7.27) | The second part of 7.21 happens here. | The centrally placed TAH does not maintain the normal pericardial space. Place a smooth saline breast implant filled with 200-250 ml of saline at the apex to maintain the pericardial space for a subsequent transplant |
| 15:32 (7.28) | The second part of 7.23 happens here. The surgical assistant did not speak along with this, but this is where the VO would go. | Place a strip of silastic beneath the sternum to protect the pump during re-entry for transplantation |
| 15:45 (7.29) We no longer use a wound vac routinely. | There is no VO for this, but it’s the last footage we have of the procedure, so perhaps accompanying VO? We do not have footage of the closed skin incision being reinforced with a wound vac (7.24) | Close the chest with interrupted steel wire, followed by routine multilayer closure of the fascia, subcutaneous tissue, and skin. |