# Submission ID #: 50575

# Editor Name: Melissa Ceo

# Videographer name:

# Film Date:

# Authors and Affiliations:

Rachel V. Bennett,†1 Chaminda M. Gamage,†1 Facundo M. Fernández1

School of Chemistry and Biochemistry

Georgia Institute of Technology

901 Atlantic Dr. NW

Atlanta, GA 30332

†These authors made equal contributions to the body of work

Title: Imaging of Biological Tissues by Desorption Electrospray Ionization Mass Spectrometry

# Corresponding Author:

Facundo M. Fernández: facundo.fernandez@chemistry.gatech.edu

Authors, please fill out the brief questionnaire below.

A.  Will you require JoVE to record video microscopy, such as filming a complex dissection or microinjection technique? (Y/N) No If yes, please list make and model of your microscope: ______________________________

B.   Does your protocol include detailed, step-by-step, descriptions of software usage? (Y/N)Yes, but can be included in written protocol and not required in video format

C.  Which steps of your protocol will viewers benefit most from having filmed? Please list 4-6 steps 2.2-2.7, 3.5-3.10

D.  What is the single most difficult aspect of this procedure and what do you do to ensure success?  Steps 2.2-2.7; the optimization of the of the source geometry in preparation for imaging

1. Introduction (Schematic Overview and Interview)

A. Schematic Overview (read by voice talent at JoVE):

Authors, please select from “Procedural Narrative” or “Conceptual Narrative” and complete the statements below. Please do not add additional steps.  Then, attach your finished graphic overview.  See accompanying instructions for details and examples.

# Procedural Narrative:

The overall goal of this procedure is to _______(insert goal here)___________________ . (Intro)

This is accomplished by first ________(insert 1st step)___________________________. (P1)

The second step is to _____(insert 2nd step)_____________________. (P2)

Next, the  _______(insert 3rd step)_______________________ . (P3)

The final step is _____(insert 4th step)_________________________. (P4)

Ultimately, _(insert method used to assay - e.g. immunofluorescence microscopy)_is used to show _( insert type of results e.g. changes in protein localization)________ _____. (P5)

# Conceptual Narrative:

The overall goal of the following experiment is to map the spatial distribution of chemicals in a biological tissue using an ambient mass spectrometry technique. (Intro)

This is achieved by using the pneumatically assisted spray of desorption electrospray ionization (DESI) to desorb and ionize the molecules present in the tissue. (P1)

As a second step, the DESI ion source is scanned across the sample area in a controlled fashion, which____________. (P2)

Next, the mass spectral information is correlated to the spatial motion of the sample stage in order to plot the ion’s intensity as a function of x and y position ultimately creating a chemical image. (P3)

Results are obtained that show that there are distinct distributions of lipids in rat brain tissues between the grey and white matter based on the mass spectrometry images of selected lipid ions. (P4)

Paste a copy of your graphic overview here.  The original file should be Adobe Illustrator (preferred) or Powerpoint (see instructions) and should be uploaded through your online submission on the JoVE website. Please keep all layers in the file (i.e., do not flatten the file).

B.  Interview: (Said by you on camera. Don’t forget to smile!)

Authors: Below are statements we would like you to complete that are complementary to the information contained within the schematic overview.

- Only one statement should be chosen and completed per author who will be on camera demonstrating the protocol.

- Enter the name of the individual who will say each line.

- Please choose and fill out the statement(s) that convey the most important fact(s) about your protocol. You may revise the given prompts as necessary to improve the sentence flow.

- If any individuals will be doing demonstrations on camera but are not assigned a speaking part in this interview section, please use statement 1.8 to introduce these demonstrators (for example, the PI introduces students)

  - Author name _________: The main advantage of this technique over existing methods, like vacuum-based ionization techniques, is that DESI is conducted under the open atmosphere, simplifying the tissue preparations for analysis.

  - Author name ________: This method can help answer key questions in the _________ field, such as _________________.

  - Author name _________: The implications of this technique extend toward therapy (or diagnosis) of_______, because ________.

  - Author name Rachel Bennett: Though this method can provide insight into chemical distributions in biological tissues, it can also be applied to other systems, such as thin layer chromatography plates, mineral surface reactions, and marine algae.

  - Author name _______: Generally, individuals new to this method will struggle because ______________.

  - Author name ________: I/We first had the idea for this method, when I/we ___________.

  - Author name _________: Visual demonstration of this method is critical as the ______________ steps are difficult to learn, because _______________.

  - **Author name ________: Demonstrating the procedure will be Rachel Bennett a graduate student from my laboratory. (Add additional mention of demonstrators as necessary).

    - Interview style: Author saying the above

    - The named technician, post doc, student looks up from workbench or desk or microscope and acknowledges the camera.

# Protocol (read by voice talent at JoVE):

Authors:  In order to ensure that your protocol can be filmed in a single day, the protocol text must be limited to 30 steps – each step being defined as 3 lines of 12 pt text in our formatting style below.  This amounts to 3 pages of protocol text.  The scope of the scripted protocol text should include only those aspects of the procedure that require visualization in order to be well understood.

- DESI Optimization

  - Prior to starting this procedure, mount the tissue sample on the sample holder.  Position the tip 3 mm from the sample surface and 5 mm from the MS capillary inlet, with the probe at an angle of 55° with respect to the sample surface. (TEXT overlay: Inner capillary of the DESI probe should extend ~1mm beyond the outer capillary).

  - For optimal transfer of ions, adjust the stage height so that the MS interface capillary hovers over the sample surface; the capillary should be <1 mm from the surface and have a collection angle of ~15°.

  - Next, align the DESI probe in the x dimension with respect to the MS capillary inlet so that they are directly in line.

  - Using the extra tissue section, adjust the y and z positions of the DESI source for optimal sensitivity using the extra tissue section.  Set the y separation between the source tip and the capillary inlet to ~4 mm, and the z separation between the tip and the sample surface to ~1.5 mm. (TEXT overlay: Note: these parameters will be slightly different for each experimental set-up).

  - Following this, adjust the distance that the inner capillary extrudes from the outer capillary of the source for maximum sensitivity and minimal impact spot size. (TEXT overlay: Caution: Ensure that the high voltage is OFF while making this adjustment).

  - To improve the sensitivity during the ionization process, wrap a rope heater around the transfer capillary and heat it to 100 °C.

- Tissue Imaging

  - Then, select the motion parameters of the sample stage based on DESI impact plume size and optimal sensitivity.  Authors, can you specify what motion parameters you used in this experiment (and what you will be demonstrating in the video)?

  - Program the LabView VI control software for desired imaging conditions, using a stage scan speed of 80 μm/s, and a line spacing between rows of 200 μm. When using the Omni Spray source, set the motion parameters for a scan speed of 100-200 μm/s and a line spacing of 200 μm.  Authors, how do you pronounce “LabView VI”?

  - Give the directory path and the file name for the position and the time files to be recorded during images within the LabView VI.

  - In preparation for the mass spectral data acquisition, calculate the total time required for imaging.  Authors, can you describe how you access the total time value by the LabView program since it is automatically calculated/provided? The Labview program automatically calculates this value, as do commercial DESI imaging software packages such as OmniSpray.

  - When using an Omni Spray automated source and stage, each line of the image is acquired as individual runs. The time-per-line and number of lines necessary for given imaging dimensions are calculated using the Omni Spray control software to be input into the mass spectrometer software.  Authors, will you demonstrate how the time-per-line and number of lines are calculated using the software?   	How are these values input into the MS software and will you be demonstrating this in the video?

# Known:Stage Speed μms

# Known:Image size, x dimension, (mm)

# Time-per-line= Stage Speed(Image size/1000)

  - Next, position the spray impact spot in the top-left of the area to be imaged.  Begin the acquisition of the MS data and the stage motion simultaneously.

  - Upon conclusion of the data acquisition, return the mass spectrometer to STANDBY mode.  Turn off the high voltage of the DESI source.

  - Turn off the nitrogen gas followed and the syringe pump.

- Image Processing

  - Using the DataManager within MassCenter, convert the acquired data to centroided data and then export it in .cdf format.

  - Finally, upload the raw .cdf mass spectral data and the two text files, position and time, to the OmniSpect website.  Use BioMAP to visualize the Firefly-processed imaging data.

  - From either software platform, images for individual ions of interest can then be visualized. Additionally, the OmniSpect website non-negative matrix factorization can be used to visualize a user-defined number of components.

- Results: Imaging of Rat Brain Tissue by DESI-MSI

Authors - please feel free to change the results title if you had something more fitting in mind.

# Authors, how should the following be pronounced?

# - [PC 34:1 + K]+ and [PC 36:1 + K]+

  - Shown here is a representative spectrum obtained from an untreated rat brain section. In the positive mode, the mass spectrum is dominated by phosphatidylcholines due to their high ionization efficiencies, which are attributed to the positively charged quaternary ammonium group. In addition, the total ion image of the tissue section shows an abundant signal across the entire brain section.

    - LAB MEDIA: Figure 3 (Video Editor: Please only show image A for the second sentence and image B for the last sentence).

  - The spatial distribution of example lipids shows how the relative abundance of different phosphatidylcholine species varies between the grey and white matter of the brain.  For example, the species [PC 34:1 + K]+ with a m/z (pronounced mass to charge ratio) of 798.5364 shows increased intensity in the cerebellar cortex or gray matter.  In contrast, the species [PC 36:1 + K]+ with a m/z of 826.5558 shows increased intensity in the cerebellar peduncle or white matter. The composite image obtained for the two ions highlights the contrast in lipid distribution across the tissue section.

    - LAB MEDIA: Figure 4 (Video Editor: Please only show image A for the second sentence, image B for the third sentence, and image C for the last sentence).

# INSTRUCTIONS FOR AUTHORS:

Please ensure that the representative results narration is appropriate and correctly describes your images, movies, or figures.  Our editors have ensured that the results are written in our format.

We consider this section a critical aspect of the video, because here is where you provide validation for your experiments.  For example, if this is a cell culture preparation, this section is where the video will show your cells at various time points following culturing.  If this is an imaging prep, then this part is where you will show examples of your imaging experiments.

Please limit the extent of narration to no more than 2-3 lines of text per image or movie file being described.   Figures with multiple panels submitted with the original protocol should be broken up so that each panel is a separate image.   Like the schematic, each image or movie file supplied in the results should be referenced by annotation in parenthesis, however for the results, the specific filename should be given in parenthesis.

Below is an example of results text:

# EXAMPLE REPRESENTATIVE RESULTS

5.  Evaluation of Morpholino Injection and Knockdown

5.1   Representative results of both morpholino injection and mRNA injection are shown here. The

uninjected control at 48 hours post fertilization looks normal, as expected

-LAB MEDIA: 0123_PIname_Figure1.tif  (Replace 0123 with your jove video #)

# 5.2   However, embryos injected with the morpholino heg_e3i3_egfr1, which knocks down Heg isoforms

containing the first of two EGF-like repeats, exhibit brain edema.

-LAB MEDIA: 0123_PIname_Figure2.tif

# 5.3   Injection of heart of glass mRNA also produced an obvious phenotype. At 24 hours post fertilization,

the heads of the uninjected controls look normal

-LAB MEDIA: 0123_PIname_Figure3.tif

5.4   Conversely, some of the embryos injected with the mRNA exhibit cyclopia

-LAB MEDIA: 0123_PIname_Figure4.jpg

# Please visit the following URL to see an example of how the results will look when complete:

http://www.jove.com/video/1597/results-example-mably?access=ksw0bprj

- Conclusion (said by authors on camera)

Authors: Below are statements we would like you to complete that summarize and conclude the video. Only one statement should be chosen and completed per author who will be on camera demonstrating the protocol. In addition to choosing and filling out the appropriate statement, please enter the name of the individual who will say each line. You may revise the given prompts if necessary to better fit your protocol.

  - Author name ________: Once mastered, this technique can be done in ____________ (hours/min) if it is performed properly.

  - Author name ________: While attempting this procedure, it’s important to remember to carefully optimize the DESI-stage-inlet geometry. These conditions are most critical for successful DESI imaging.

  - Author name ________: Following this procedure, other methods like _____________ can be performed in order to answer additional questions like _____________.

  - Author name ________: After its development, this technique paved the way for researchers in the field of __________ to explore _____________ (subdivision of field, disease, natural phenomenon) in __________( model organism, patient demographic, organ system).

  - Author name _________: After watching this video, you should have a good understanding of how to _____________ (restate overall goal of the procedure mention specific steps).

  - Author name _________: Don't forget that working with _____________(reagent, pathogen, instrumentation) can be extremely hazardous and precautions such as ____________ should always be taken while performing this procedure.

# Provided Media

Authors, Please list all images, movie files, or 3-D rendered animations that can be included in the video per editor’s request.  The step in the script/video where these images will be inserted should be specified.   For example:

# 6.2 –  0123_PIname_Figure1.tif -  dual color imaging of tumor angiogenesis at 40X

6.2 –  0123_PIname_Figure2.tif -  dual color imaging of tumor angiogenesis at 100X

Formats:  For static images we prefer .tiff, Illustrator, Powerpoint or Photoshop files at dimensions of at least 720X480 pixels and 300 dpi.  The higher resolution, the better.  Likewise any exported movie files should have at minimum these dimensions and be rendered to .mov, .mp4, or .avi files.

50575_Layered-Figures.pptx

# General Preparation

# It’s critical for a smooth and organized shoot that all reagents are accounted for, in advance.

Any overnight or long incubation steps should be recognized and specimens/samples be prepared in advance so that prior steps can be recorded and shooting can continue with pre-prepared specimens/samples.

# All tubes/flasks should be pre-labeled neatly before we arrive.

Ex. Luciferase assay done in 96 well plates should be labeled with negative/positive control wells and experimental samples are labeled accordingly.

You will receive more detailed preparation instructions, as well as an introduction to your videographer, closer to your filming date.