# Submission ID #: 50604

# Editor Name: Roslyn A. Bauer

# Videographer name: Micah Wolf

# Film Date: 6/3/2013

# AUTHORS:

# Grady E. Carlson

Department of Chemical and Biomolecular Engineering

Russ College of Engineering and Technology

Ohio University

171 Stocker Center

Athens, OH 45701

gc733706@ohio.edu

# Eric W. Martin

Biomedical Engineering Program

Department of Chemical and Biomolecular Engineering

Russ College of Engineering and Technology

171 Stocker Center

Ohio University

Athens, OH 45701

em849611@ohio.edu

Monica M. Burdick, Ph.D.

Department of Chemical and Biomolecular Engineering

Russ College of Engineering and Technology

171 Stocker Center

Ohio University

Athens, OH 45701

burdick@ohio.edu

Tel 740-593-0033

Fax 740-593-0873

CORRESPONDING AUTHOR: Monica M. Burdick, Ph.D.

Title: Simultaneously capturing real-time images in two emission channels using a dual camera emission splitting system: Applications to cell adhesion

Authors, please fill out the brief questionnaire below.

A. Will you require assistance with video microscopy, such as filming a complex dissection or microinjection technique (Y/N, please specify steps by number. Also, please list make and model of your microscope)?  No, Leica DMI 6000B

B.   Does your protocol include detailed, step-by-step, descriptions of software usage (Y/N, please specify steps by number)? Y, (2.3, 2.4, 2.5, 2.6)

C.  Which steps of your protocol will viewers benefit most from having filmed? Please list 4-6 steps

2.3, 2.8, 2.11, 4.3, 4.12, 4.13

D.  What is the single most difficult aspect of this procedure and what do you do to ensure success?

The most difficult aspect of this procedure is camera alignment in the dual camera emission splitting system. To ensure that cameras are aligned correctly we use the manufacturer’s alignment slide to calibrate the cameras into alignment while meticulously following the manufacturer’s instructions.

1. Introduction (Schematic Overview and Interview)

- Schematic Overview (read by voice talent at JoVE):

  - The overall goal of this procedure is to perform a parallel plate flow chamber adhesion assay using a dual camera emission splitting system to simultaneously record real-time image sequences in two colors (Intro).

  - To do this, the two cameras on the imaging setup are first aligned. (P1).

  - Next, the cells and reactive substrate are prepared and (P2) and the parallel plate flow chamber is set up. (P3).

  - The camera settings are then calibrated for two-color image acquisition, and any needed digital corrections are made to camera alignment. (P4).

  - Image sequences are then captured, which demonstrate acquisition of dual color images with a high temporal resolution (P5)

- Interview: (Said by you on camera. Don’t forget to smile!)

  - Grady:  The advantage of the dual camera emission splitting system over single camera systems is that different exposure times can be used for each camera, and the full field of view is retained.

  - Grady: This technique can help answer key questions in the cell adhesion field, such as how might localization of adhesion molecules on the cell surface affect dynamic biological processes like cell migration, inflammation, and cancer metastasis.

# Protocol (read by voice talent at JoVE):

- Installation of the dual camera system

  - The following procedure requires an inverted epifluorescence microscope.  This should be  equipped with a high intensity light source coupled to a light guide, a combination filter cube, and a dual camera emission splitting system camera system.

    - MED:  Talent approaches microscope setup and turns it on

  - The microscope setup should be connected to a computer running imaging software capable of collecting and combining images captured by monochrome CCD cameras.   Here StreamPix 5 Multi-Camera software is used with the SimulPix module.

    - MED:  Talent points to the two cameras

  - Begin by opening StreamPix 5.   In the task ribbon, select “workspace”, then, at far right of the screen open “workspace manager” and select “new workspace”.

    - MED:  Talent opens streamPix

    - SCREEN:  Talent selects workspace, opens workpace manager and selects new workspace

  - After naming the camera, follow software prompts to select, from a pre-populated list, a grabber-matching camera model.   A small camera icon will appear in the workspace to indicate that the camera feed is received. Objects imaged with the microscope may now be viewed on screen.

    - SCREEN:  Talent finishes naming the camera and follows software prompts to select a grabber from a pre-populated list.   Talent selects the grabber matching camera model.   And camera icon appears in workspace

  - Repeat this process for the second camera workspace.

    - SCREEN:  Talent repeats selection of grabber for second camera.

  - Then, for the merged workspace, repeat this once more. A “grabber failed to load” error message will appear. This message is normal.

    - SCREEN:  Talent repeats again for merged workspace, and a grabber failed to load error message appears.

  - Assign cameras from workspace 1 and 2 to the merged workspace in the docking panel located on the right-hand side of the viewing screen.  The images in workspace 1 and 2 will be superimposed in the merged workspace as two pseudocolored images.

    - SCREEN:  Talent assigns cameras from ws1 and 2 to the merged work space and the images are superimposed in the merged workspace.

- Camera Alignment

  - Grady: “The most difficult aspect of this procedure is camera alignment. To ensure success, we use the manufacturer’s alignment slide while meticulously following the manufacturer’s instructions.”

    - MED:  Talent looks up from setting up microscope and says the above line

  - To align the cameras in the dual camera emission splitting system, send bright field or phase contrast light to the microscope eyepiece. Place the metallic-coated calibration grid that was provided by the manufacturer on the microscope stage, and square the grid on the microscope stage.

    - MED:  Talent sends light to the eyepieces and places the calibration on the stage

    - CU:  Talent squares the grid on the stage

  - Display the grid without binning and without auto-scaling and bring it into focus.  Then, displace, but do not remove, the dichroic (pronunciation: dye-croh-ic) housing to align camera 1. Then, using the focus adjustment dial on C-mount port 1, focus the image again.

    - SCREEN:  Talent brings grid into focus.

    - CU:  Talent displaces the dichroic housing to align the camera.

    - CU:  Talent uses focus adjustment on c-mount port 1

    - SCREEN:    Talent focuses the image

  - Next, push the dichroic housing into the dual channel acquisition device fully, so that the image is split by the dichroic to both cameras.

    - MED:  Talent pushes the dichroic housing into the dual channel acquisition device

  - Loosen the three 5/64” set screws and rotate the second camera on female C-mount  port 2 until the orientation of the grid is identical in both images. Using the focus adjustment dial on C-mount port 2, bring the image from camera 2 into focus.

    - CU:  Talent loosens the set screws and rotates the second camera

    - SCREEN:  Show the adjustment of the grid as it is adjusted to match.

    - CU:  Talent uses focus on C-mount port 2

    - SCREEN:  image from camera 2 comes into focus.

  - To finalize the alignment, use the R/L knob on the right side of DC2 to adjust the combined image positions until right/left vertical boundary of the images is precisely aligned in the merged workspace.

    - CU:  Talent rotates R/L knob

    - SCREEN:  Combined image becomes vertically aligned

  - Then, use the U/D knob to adjust the up/down alignment of the second image until the horizontal grid bars are properly aligned.

    - CU:  Talent uses U/D know to adjust the up/down alignment of second image

    - SCREEN:  The horizontal grid bars come into alignment

  - If during the up/down alignment, a slight camera rotation occurs, correct this issue by loosening the three 5/64” inch screws for camera 2 and rotating until the displayed image is properly aligned.

    - SCREEN:  Show what happens if rotation occurs

    - MED:  Talent loosening screws and rotating camera

    - SCREEN:  Image comes into alignment

- Preparation of cells, reactive substrate and parallel plate flow chamber

  - Prepare BT-20 breast cancer cells for parallel plate flow chamber adhesion assay by staining them with anti- CD24 and HECA-452, using AlexaFluor 568 and 488 secondaries as described in the accompanying document. Be sure to prepare the appropriate single-color controls.

    - (This was slated 4.4.1) MED:  Talent working at bench to stain cells (30 s of video)

  - Following staining, dispense cells into the reservoir of the parallel plate flow chamber.

    - CU:  Talent pipettes cells into parallel plate flow chamber

  - Connect two separate lengths of tubing to the inlet and outlet ports of flow chamber. One tube will connect to the reservoir containing the labeled cells suspended in DPBS+ containing 0.1% BSA.

    - CU:  Talent connects tubing to inlet and outlet ports of flow chamber

    - CU:  Talent places flow chamber in 35 mm tissue culture dish

    - CU:  Talent connects to reservoir

  - Connect the other tube to the syringe pump, which will withdraw fluid at a specified flow rate.

    - (This was slated 4.4.1.1) CU:  Talent connects to syringe pump

  - Connect a vacuum line to the flow chamber and position the flow chamber over the 35mm tissue culture dish containing a monolayer of Chinese hamster ovary cells expressing E-selectin, or “CHO-E” cells. (pronounce with hard O sound (Oh) and hard E sound (we) chO-E)  (TEXT: See accompanying document for culture details)

    - MED: Talent connects a vaccum line to flow chamber

    - CU:  Talent positions flow chamber over dish with cells

  - Adjust the flow chamber and flow chamber gasket to ensure an adequate vacuum seal.

4.6b. Withdraw fluid from the reservoir, monitoring the flow chamber assembly for proper seal and no visual evidence of air bubble formation.

    - CU:  Talent adjusts gasket

4.6b.1  Use 4.7.1. CU:  Talent withdraws fluid from the reservoir while monitoring for proper seal (Video Editor: The post shoot notes were confusing here, but if the chamber is already under the microscope for this shot, then please move this -- 4.6b (VO) and 4.7.1 (shot) -- after the next step).

  - Finally, place the flow chamber assembly on the microscope.

    - MED:  Talent places assembly on microscope

4.6.2. MED: Talent places assembly on microscope

4.6.3. CU:  Talent places assembly on microscope

- Two-color image acquisition

  - Power on the light source and microscope setup.  By manual inspection, ensure that the dichroic housing is in the correct, depressed, operating position and is not in bypass mode.

    - MED:  Talent powers on microscope and light source

    - CU:  Talent inspects dichroic housing

  - Manually change to fluorescence mode and select the green/red fluorescence filter cube.

    - MED:  Talent changes to fluorescence mode and selects the proper fluorescence filter cube

  - To determine the exposure time and gain, separately image the cells using red fluorescence in camera 1 and green fluorescence in camera 2. Manually depress the dichroic as needed to obtain the images.

    - MED:  Talent images first cell sample

    - SCREEN:  Talent images second cell sample and both workspaces show samples with detectable emmisions.

  - Next, deposit a suspension of BT-20 cells stained with only red fluorophore-conjugated antibody in the reservoir connected to the inlet port of the parallel plate flow chamber. Use the syringe pump to perfuse BT-20 cells over the CHO-E monolayer in the parallel plate flow chamber.

    - Use 4.2.1. CU:  Talent deposits a suspension of BT-20 cells in reservoir of flow chamber,

    - MED:  Talent starts syringe pump

    - SCREEN:  show perfusion of BT-20 cells onto monolayer (if possible)

  - In the  “live settings” of the imaging software, adjust the exposure time of camera 1 to obtain an image in the red channel.

    - SCREEN:  in “live settings”, talent adjusts exposure time of camera 1 to obtain image in red channel

  - Match the exposure time of camera 2 to the exposure time of camera 1. If an image is observed in the green channel, bleed-through artifacts are present.

    - SCREEN:  Talent matches exposure time of camera 2 to camera 1.      Image is observed in green channel.

  - If this happens, keep the exposure time equivalent in camera 1 and camera 2 and reduce the exposure time until the bleed-through artifact is no longer visible in the green channel. Adjust the gain of camera 1 to improve image visibility in the red channel if necessary.

    - SCREEN:   Talent reduces exposure time until bleed through is no longer visible.  Talent then adjusts gain of camera 1.

  - Next, remove any remaining BT-20 cell suspension from the reservoir and replace the suspension with a DPBS. Use the syringe pump to perfuse the solution over the CHO-E monolayer until the BT-20 cells are no longer visible on the monolayer.

    - CU:  Talent removes BT-20 suspension from reservoir and replaces it with buffered solution.

    - MED:  Talent uses syringe pump to perfuse solution over monolayer

    - SCREEN:  Perfusion clears BT-20 cells from monolayer

  - Refill the reservoir with a suspension of BT-20 cells stained with only green fluorophore-conjugated antibody.

    - Use 4.2.1. CU:  Talent refills reservoir with suspension of BT-20

  - Repeat the camera calibration process using the green single-color controls.   This time, define exposure time with camera 2 to image cells in the green channel without bleed-through artifacts in the red channel collected by camera 1.

    - MED:  Talent, seated at computer repeats camera calibration process

  - Use the imaging software to simultaneously view images captured from the two monochrome CCD cameras in the flow chamber experiment.

    - SCREEN:  images are viewed simultaneously

  - Merge the images collected from both cameras using the SimulPix module of StreamPix.

    - SCREEN:  Talent merges images using simupix module

  - Use digital correction adjustments in SimulPix or alternative software to spatially align the merged images, if necessary. Images can be corrected with the SimulPix module for horizontal, vertical, and rotational adjustments.

    - MED:  Talent seated at computer uses digital correction to spatially align images

  - The system is now ready to be used to capture 2 color images simultaneously.

    - MED:  Talent begins experiment.

- Representative Imaging Results:

  - A parallel plate flow chamber adhesion assay was used to demonstrate the dual camera emission splitting system described in this video.

    - 50604_Burdick_Figure 1

  - As shown here, the system detected BT-20 cells that were fluorescently labeled with anti-human CD24, shown in red, and anti-HECA-452 (pronounced: heck-a four-five-two), which binds to sialofucosylated molecules, shown in green.

    - 50604_Burdick_Figure 2

  - Some rolling cells displayed red, but not green signals.  Others displayed green signals but not red signals, and yet other cells displayed both red and green signals.

    - 50604_Burdick_Figure 2 movie

  - Here, the cameras maintained the optical and temporal resolution to track cell features on a cell tumbling “end over end” as it rolled on the reactive substrate in the direction of flow.

    - 50604_Burdick_ Figures 3movie

  - Merged emission channels revealed the distribution and colocalization of CD24 and sialofucosylated molecules, on the surface of BT-20 cells rolling and adhering to CHO-E monolayers.

    - 50604_Burdick_ Figures 4movie

  - These images show a zoom of a single BT-20 cell in which CD24 and sialofucosylated molecules colocalize and appear as yellow to orange spots when the red and green pseudocolored emission channels are merged.

    - 60604_Burdick_Figure4

  - Taken together, this information demonstrates that the dual camera emission splitting system has the spatial, temporal, and optical resolution needed to reveal cellular and molecular features in applications such as flow chamber adhesion assays, in which cells move rapidly through the field of view.

    - 60604_Burdick_ Figure 5

- Conclusion (said by authors on camera)

  - Grady: After watching this video, you should have a good understanding of how to set up and calibrate a dual camera emission splitting system for real-time two color imaging of a parallel plate flow chamber adhesion assay.

    - Interview style

  - This method can be applied to other in vitro assays that use fluorescence to study cell behavior.

    - Interview style

# Provided Media

# 50604_Burdick_SchematicOverview

# 50604_Burdick_Figure1

# 50604_Burdick_Figure2

50604_Burdick_Figure2movie

50604_Burdick_Figure3

50604_Burdick_Figure3movie

50604_Burdick_Figure4

50604_Burdick_Figure 4movie

50604_Burdick_Figure5

# General Preparation

# It’s critical for a smooth and organized shoot that all reagents are accounted for, in advance.

Any overnight or long incubation steps should be recognized and specimens/samples be prepared in advance so that prior steps can be recorded and shooting can continue with pre-prepared specimens/samples.

# All tubes/flasks should be pre-labeled neatly before we arrive.

Ex. Luciferase assay done in 96 well plates should be labeled with negative/positive control wells and experimental samples are labeled accordingly.