## Slide 1
Open slidebook – new “slide” opens- save.
camera icon= capture window ( to take final images)
F= focus window (controls pre-imaging setup)
Slidebook image capture:

## Slide 2
Set objective- [scope menu] one time application.
Filter selections in color.

## Slide 3
Set Camera intensity- [camera menu] one time application

## Slide 4
Set Z stack size [Z menu]-set top- rotate- set bottom.
Z-step size setting- sometimes we use non-optimal (1.5 um)

## Slide 5
Open capture window (2D image example)
- choose colors to be captured-
click on each- set intensity for each

## Slide 6
Pre-capture- 2D image example- setting channel intensity
( vary exp. time- intensify settings -use test button)

## Slide 7
All three channels set- with intensify setting for each
Ready for 2D image capture

## Slide 8
Capture settings for 3D (Z-stack)- check 3D box-
Select “use top and bottom positions”- settings made above using focus window
will transfer over- note number of planes and step size

## Slide 9
Post-capture- new image in slide- X, Y, Z

## Slide 11
4D imaging setup- select buttons for 3D AND timelapse.
Set time intervals. Z-stack setting applies to each timepoint.
Everything else is the same.

## Slide 12
Post –image capture. 4D image saved.
Note dimensions- (X, Y, Z, T)
Z
t