Porte 52688 redos (13)

2.2 (2:10) - Switch on the venous and arterial pumps.  Follow the manufacturer's instructions on the screen to increase the pump output on both sides.  Remove all the air bubbles from the closed system in a systematic manner. (rewrite)

2.3a (2:22) - Null the pressure meter against atmospheric pressure by following the instructions on the screen.  This ensures that the pressure measured during the perfusion is the real pressure at the level of the portal vein. (rewrite)

2.3d (2:45) - Repeat the pressure nulling procedure on the arterial side of the device. Set the temperature to 37 degrees and the mean arterial pressure at 70 milli-meters of mercury. (rewrite)

2.5b (3:55) - Remove all air bubbles from the blood gas sample. (rewrite)

3.9b (6:04) - Do not insert the catheter too deeply into the bile duct as this may cause injury to the biliary epithelium.

3.10b (6:20) - If the graft has been preserved in the University of Wisconsin solution, flush it out with 2 liters of cold saline solution followed by half a liter of warm saline solution. (rewrite)

3.11 (6:38) - If the graft has been preserved in H-T-K solution, flush out the liver with 1 liter of cold saline solution followed by half a liter of warm saline solution. (rewrite)

4.3b (8:22) - Immediately remove any air bubbles from the syringe.  Store perfusate fluid samples for later analysis in blood collection tubes. (rewrite)

4.4 (8:41) - Insert the 1 milli-liter syringe in the blood gas analyzer and follow the manufacturer's instructions for analysis of blood gas and biochemical parameters. (rewrite)

4.5a (8:51) - Take perfusion fluid samples for later analysis, transfer them to blood collection tubes, and store on melting ice.  Centrifuge the samples at 1,400 times g and 4 degrees C for 5 minutes. (rewrite)

4.6 (9:13) - To evaluate the quality of the liver during perfusion, monitor the macroscopic homogeneity of the liver. (rewrite)

5.3a (9:37) - The hemodynamics of the livers were assessed by monitoring the changes in the arterial and portal flows during normothermic machine perfusion. (rewrite)

5.6 (10:36) - Stable concentrations of hepatic injury markers such as ALT, alkaline phosphatase, gamma-GT, and potassium in the perfusion fluid reflected minimal injury of the grafts. (rewrite)