Submission ID #: 53339
Editor Name: Michael Linnes
Videographer name: Josh Mandell 
Film Date: October 24th, 2015
Authors and Affiliations: Jiang Yang1, Tae-Joon Kwak2, Xiaodong Zhang4, Woo-Jin Chang2,3*, Sundaram Gunasekaran1*

1Department of Biological Systems Engineering, University of Wisconsin-Madison, Madison, WI, USA 2Department of Mechanical Engineering 3School of Freshwater Sciences, University of Wisconsin-Milwaukee, Milwaukee, WI, USA

Title: Iridium Oxide-reduced Graphene Oxide Hybrid Thin Film Modified Screen-printed Electrodes as Disposable Electrochemical Paper Microfluidic pH Sensors

Corresponding Authors: 

Sundaram Gunasekaran
guna@wisc.edu
Woo-Jin Chang
wjchang@uwm.edu
Questionnaire:
A.  Will you require JoVE to record video microscopy, such as filming a complex dissection or microinjection technique? (Y/N)_____N____   
B.   Does your protocol include detailed, step-by-step, descriptions of software usage? (Y/N)___Y_____ If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

C.  Which steps of your protocol will viewers benefit most from having filmed? Please list 4-6 individual steps using the step numbers listed in this document. (Please do not list entire sections.) _______2.4, 2.6, 3.7, 4.5_____

D.  What is the single most difficult aspect of this procedure and what do you do to ensure success?  Please list 1-2 individual steps using the step numbers listed in this document. (Please do not list entire sections.) _2.6: If the pad is handled careless, the hydrophobic PDMS could contaminate the hydrophilic wicking region so that the sensing and fluid delivery will be limited. The vacuum cover and the stamp should be examined to be clean before further stamping.__3.7: because of the relative hydrophobicity of reduced grapheme oxide, the solution-based electrodeposition is not quite efficient. The electrode is recommended to be immersed in the deposition solution for 5 min to reach quasi-stable conditions.________________________

E.  Will the filming need to take place in multiple locations? (Y/N) ___N 
1. Introduction (Experimental Goal and Author Interviews) – As the beginning of your video, the introduction should clearly present the goal of your method to the viewer and its significance.  Other information can be provided according to the various statements below, but the total introduction should not exceed 150 words. 
A. Experimental Goal (read by voice talent at JoVE): 
The overall goal of this method is to use the electrochemical synthesis of graphene-iridium oxide hybrid thin films to fabricate a portable electrochemical paper-fluidic sensor for point-of-care pH measurements. (Intro)
B.  Required Interview Statements: (Said by you on camera. Don’t forget to smile!)   
1.1. Prof. Gunasekaran: This method provides a new way to perform point-of-care pH measurements by combining the synthesis of anodically electrodeposited iridium oxide thin films, assisted by reduced graphene oxide scaffolds. Combined with paper microfluidic pads and digital voltammetric display, it truly provides a portable and disposable pH sensor.
1.2. Prof. Chang: This technique combines the features of paper strip and solid-state pH meters using paper microfluidics and electrochemical detection to offer quick, accurate, and point-of-care measurements of the pH outside labs using low sample volumes.   

C.  Optional Interview Statements: (Said by you on camera. Don’t forget to smile  
1.3. Dr. Jiang Yang: We first had the idea for fabrication of such devices, because we wanted to measure the relatively small pH variations of lake water in different months in Madison, WI and found the pH paper strip cannot tell such a small pH difference and it can hardly be read with the chart. It is also not practical to use a laboratory-based pH meter.  
Protocol (read by voice talent at JoVE): 
2. μPAD and Apparatus Preparation

2.1. To begin, secure a 7 cm x 4 cm x 0.6 cm sheet of ABS or other compatible plastic into the holder of a three-dimensional milling machine [2.1.1 - MED - TXT] that is equipped with a 1.6 mm milling bit. [2.1.2 - CU]
2.1.1. *Film as written (TEXT: ABS: 7 cm x 4 cm x 0.6 cm)
2.1.2. Close-up of milling bit.
2.2. Use the milling machine to engrave a dogbone shaped groove, 500 μm deep, into the surface. [2.2.1 - MED] Create the shape by milling two 1 cm radius circles connected by a 1 cm by 5 mm wide channel.  This pattern will be used as the stamp.[2.2.2 - CU]
2.2.1. Talent starts milling machine.

2.2.2. Close-up of milling machine cutting the dogbone shape.

2.3. Next, place a separate sheet of plastic into the mill to be used as the vacuum cover [2.3.1 - MED] and mill out a groove that is large enough to contain the entire stamp pattern. [2.3.2 - CU]
2.3.1. Talent places plastic sheet into the mill and starts the mill.

2.3.2. Close-up as mill finishes the cover.    
2.4. Next, prepare a mixture of PDMS pre-polymer and cross linker at a ratio of 10 to 1 and mix the two parts with a spatula. [2.4.1 - CU] Then, place the mixture into a vacuum chamber for 5 minutes to degas the mixture.[2.4.2 - MED]
2.4.1. *Film as written

2.4.2. *Film as written

2.5. Apply an appropriate amount of degassed PDMS to the channel groove of the stamp so that it fills the entire groove. [2.5.1 - CU] Scrape away the excess material with a flat piece of plastic. [2.5.2 - CU]
2.5.1. *Film as written

2.5.2. *Film as written

2.6. Then, place a piece of filter paper pad, pre-cut to the desired size, on top of the PDMS filled stamp. [2.6.1 - CU] Set the vacuum cover on the opposite side of the stamp across the paper and apply a light finger pressure on the stamp. [2.6.2 - CU]
2.6.1. *Film as written

2.6.2. *Film as written

2.7. Apply vacuum for up to 30 seconds using a hand-operated vacuum pump. [2.7.1 - MED] When finished, remove the paper pad [2.7.2 - CU] and place it in a convection oven for 10 minutes at 80 °C to harden the patterned PDMS.[2.7.3 - MED - TXT]
2.7.1. *Film as written

2.7.2. Talent opens the chamber and removes the paper pad.

2.7.3. *Film as written (TEXT: 10 min, 80 °C)
2.8. Prof. Chang: “Take special care during the application of PDMS to avoid any possible PDMS contamination to the inner hydrophilic region of the filter paper which is where the sample liquid will be dropped.” [2.8.1 - INT]
2.8.1. Author saying the above interview style.

3. Modification of Screen Printed Electrodes with IrO2-RGO Hybrid Thin Films

3.1. First, obtain a screen printed 3-electrode array [3.1.1 - MED] consisting of a graphitic carbon working electrode, a counter electrode, and a silver-silver chloride pellet reference electrode. [3.1.2 - CU]
3.1.1. Talent places electrode into a petri dish or other container.

3.1.2. Talent uses forceps to point to the electrodes in the order listed above.
3.2. Then, use a micropipette to drop cast 3 microliters of a 1 mg/ml graphene oxide solution onto the graphitic carbon working electrode. Let the solution dry at room temperature.   [3.2.1 - CU]
3.2.1. *Film as written 

3.3. In the meantime, nitrogen purge 10 mL of a pH 5.0 PBS buffer for 20 minutes. [3.3.1 - MED] When the screen printed electrode is dry, dip it into the buffer while the nitrogen continues to flow.[3.3.2 - CU]
3.3.1. *Film as written

3.3.2. *Film as written

3.4. Next, connect the electrode to an electrochemical work station [3.4.1 - CU] and conduct 100 cycles of repetitive cathodic potential cycling from 0.0 to -1.5 V to electrochemically reduce the graphene oxide. [3.4.2 - MED - TXT]
3.4.1. *Film as written

3.4.2. Talent starts the potential cycling from the work station. (TEXT: 100 cycles, 0 to -1.5 V)

3.5. Then, rinse the screen printed electrode with deionized water and dry it at room temperature. [3.5.1 - MED]
3.5.1. Talent rinses the electrode and sets it aside to dry.
3.6. Dr. Jiang Yang: “The homogeneity of the reduced graphene oxide film is important, because it serves as the carbon support for further growth of iridium oxide thin films.”[3.6.1 - INT]
3.6.1. Author saying the above statement interview style

3.7. Next, make 100 ml of iridium oxide deposition solution in deionized water. [3.7.1 - MED - TXT] Gradually add a small amount of anhydrous potassium carbonate while stirring until the pH of the solution reaches 10.5 and the solution turns a yellowish color. [3.7.2 - CU]
3.7.1. (TEXT: IrO2 Deposition Solution: 0.15 g IrCl4, 0.6 mL 50% (w/w) H2O2, 0.5 g C2H6O6)  
3.7.2. *Film as written 
3.8. Age the solution for 48 hours at room temperature. [3.8.1 - CU] During this time, observe as its color turns from yellow to a pale blue. [3.8.2 - CU - TXT]
3.8.1. Close up of solution (yellow) 

3.8.2. Close up of pale blue solution (TEXT: 48 Hours later) (Authors: Please prepare both solutions and replace the yellow solution with the blue one so that it appears to be in the same container.)
3.9. Place the reduced graphene oxide screen printed electrode into the iridium oxide deposition solution [3.9.1 - CU] and use the electrochemical work station to apply a constant potential of positive 0.6 volts for 5 minutes.  [3.9.2 - MED - TXT]
3.9.1. *Film as written

3.9.2. *Film as written (TEXT: Apply +0.6 V, 5 min)

3.10. Confirm the structure of the sensing area by scanning electron microscopy using standard SEM sample preparation and imaging techniques.[3.10.1 - CU]
3.10.1. Talent prepares sample for SEM by mounting specimen on an appropriate stub. 

4. Construction of Inexpensive and Portable Digital pH Meters

4.1. To begin construction of the portable Digital pH meter, collect the required components as shown here, which are listed in the accompanying text protocol under the list of materials.[4.1.1 - CU - TXT]
4.1.1. Show all of the materials that are needed, laid out on the bench. (TEXT: *See accompanying list of materials)
4.2. Then, plug in the INA111 (pronounced: I-N-A-one-one-one) high speed field-effect transistor input instrumentation amplifier onto the breadboard to achieve sufficiently high internal impedance for stable measurements. [4.2.1 - MED/CU - TXT]
4.2.1. *Film as written (TEXT: high input impedance >1012 Ω)

4.3. Next, connect two grounded 9-V alkaline consumer batteries in series to power the pH meter…  Connect the cathode to pin 7 and the anode to pin 4… [4.3.1 - CU]  Then, connect a 10 mega ohm resister between pins 1 and 8.   [4.3.2 - CU]
4.3.1. *Film as written

4.3.2. *Film as written

4.4. Also, connect the positive probe of a digital multimeter to pin 6 and the negative probe to pin 5. [4.4.1 - CU] Set up the multimeter to measure and display the output voltage. [4.4.2 - CU]
4.4.1. *Film as written

4.4.2. *Film as written

4.5. Finally, connect the reference end of the modified screen-printed electrode to pin 2 and the working electrode to pin 3.   [4.5.1 - CU]
4.5.1. *Film as written

5. pH Measurements 

5.1. Prepare 100 ml mixtures of Britton–Robinson buffers by combining varying amounts of 0.2 M sodium hydroxide with a 0.04 M acid solution composed of equimolar amounts of phosphoric acid, acetic acid, and boric acid combined. [5.1.1 - MED - TXT]
5.1.1. Talent prepares mixtures (TEXT: pH 2 - 95 mL Acid : 5 mL NaOH, pH 4 - 75 mL Acid : 25 mL NaOH, pH 6 - 56 mL Acid : 42 mL NaOH, pH 8 - 40 mL Acid : 60 mL NaOH, pH 10 - 22 mL Acid : 78 mL NaOH, pH 12 - 2 mL Acid : 98 mL NaOH).  (Video Editor: Make a table as shown instead of writing out all the text)
	pH
	0.04 M Acid (mL)
	0.2 M NaOH (mL)

	2
	95
	5

	4
	75
	25

	6
	56
	44

	8
	40
	60

	10
	22
	78

	12
	2
	98


5.2. Next, assemble the microfluidic pH meter by placing the modified screen printed electrode into the milled plastic holder [5.2.1 - CU] and then setting the microfluidic paper on top of the electrode so that one of the hydrophilic circles is located over the sensor. [5.2.2 - CU]
5.2.1. *Film as written

5.2.2. *Film as written

5.3. Using a pipette, add 60 µL of one of the pH samples directly onto the hydrophilic area of the micro-pad so that it will wick towards the sensor. [5.3.1 - CU] Then, begin to measure the voltage signal with an assembled portable digital pH meter.[5.3.2 - CU]
5.3.1. *Film as written

5.3.2. View of multi-meter as the value normalizes.
5.4. Once the open circuit potential becomes steady, record the value. [5.4.1 - MED] Repeat this process with each of the prepared standards to create a standard curve, using a new microfluidic paper each time.  [5.4.2 - CU]
5.4.1. Talent records the value in a notebook

5.4.2. Talent switches out the old paper for new and adds the next sample to the strip.
6. Results: Disposable Microfluidic pH Sensors Accurately Measure pH in Small Volumes
6.1. The electrochemical reduction technique described for the screen printed electrodes will produce a defect-free homogeneous reduced graphene oxide thin film, that is visible under scanning electron microscopy.[6.1.1 - LM]
6.1.1. Figure 2a
6.2. After iridium oxide modification, a uniform and smooth film should appear without any observable cracks. [6.2.1 - LM]
6.2.1. Figure 2c

6.3. Only a small amount of fluid is necessary as the sample wicks along the microfluidic paper towards the sensor. In the image shown here, a sample of blue dye solution was placed at the base of the paper and was drawn towards the sensor location, dying the entire piece of paper.

6.3.1. Figure 1b (Authors: Please submit a version of this image to your upload site with the arrows/numbers absent.  Alternatively, provide a similar set of images with the dye partially drawn towards the sensor and then one with the entire paper blue as shown in Figure 2b. )
6.4. The time required to reach a steady reading varies depending on the pH of the sample.  As with any pH meter, it is important to wait until a steady state is achieved before taking the reading. 

6.4.1. Figure 4a (Video Editor: Point to the dotted lines on the right with the word “steady state”)
6.5. Once steady state is reached, the values acquired with the microfluidic pH sensor linearly correlate with those acquired from a typical pH meter.

6.5.1. Figure 5a (Video Editor: Highlight the Y-axis label with the words “the microfluidic pH sensor” and the X-axis label with the words “typical pH meter.”)

7. Conclusion (said by authors on camera) 
7.1. Jiang Yang: The graphene-iridium oxide hybrid thin film pH meter has a small hysteresis, fast response time, reproducible performances, and good correlation with the commercial glass electrode of a laboratory-based pH meter. 
7.2. Jiang Yang: Once mastered, the pH measurements can be carried out rapidly, on the order of a few (seconds/minutes).  *Author: Please specify the time frame for measurement.
7.3. Prof. Chang: Following this procedure, other pH-responsive carbon-metal oxide hybrid thin films can be synthesized in a similar way on any conducting or semi-conducting substrates such as glassy carbon, indium tin oxide coated glass, and silica. 
7.4. Prof. Gunasekaran: The electrochemical microfluidic sensor combines advantages of both pH meters and pH strips and is a promising platform for future in-field or point-of-care pH measurements in limited-resource settings.
Provided Media

Authors, Please list all images, movie files, or 3-D rendered animations that can be included in the video per editor’s request.  The step in the script/video where these images will be inserted should be specified.   For example:

6.2 –  0123_PIname_Figure1.tif -  dual color imaging of tumor angiogenesis at 40X 

6.2 –  0123_PIname_Figure2.tif -  dual color imaging of tumor angiogenesis at 100X

Formats:  For static images we prefer .tiff, .eps, Illustrator, Powerpoint or Photoshop files at dimensions of at least 720X480 pixels and 300 dpi.  The higher resolution, the better.  Likewise any exported movie files should have at minimum these dimensions and be rendered to .mov, .mp4, or .avi files.  

Insert your media filenames here. We should record a video showing the flow of dye solution in the paper pads
General Preparation

It’s critical for a smooth and organized shoot that all reagents are accounted for, in advance.   

Any overnight or long incubation steps should be recognized and specimens/samples be prepared in advance so that prior steps can be recorded and shooting can continue with pre-prepared specimens/samples.  

All tubes/flasks should be pre-labeled neatly before we arrive.  

Ex. Luciferase assay done in 96 well plates should be labeled with negative/positive control wells and experimental samples are labeled accordingly.

You will receive more detailed preparation instructions are included in the email accompanying the finalized script.

( 2013, Journal of Visualized Experiments


