TITLE: Infrared videography coupled with smart phone LED light for the measurement of pupil light reflex

AUTHORS: Lily Y-L Chang 1, 3, 4, Jason Turuwhenua 1, 2, Tian Qu 1, Joanna Black 1, 3, 4, Monica L Acosta1, 3, 4

1 School of Optometry and Vision Science, Faculty of Medical and Health Sciences, The University of Auckland, New Zealand

2 Auckland Bioengineering Institute

3 Centre for Brain Research, The University of Auckland, New Zealand

4 New Zealand National Eye Centre, The University of Auckland, New Zealand

CORRESPONDING AUTHOR:  Dr Monica L Acosta m.acosta@auckland.ac.nz

KEYWORDS:

Optometry, Ophthalmology, Neuro-ophthalmology, Autonomic nervous system, Parasympathetic, Sympathetic, Pupil response, Constriction, Dilation, Smart phone, Infrared videography, Guinea pig

SHORT ABSTRACT:

Light triggers the pupillary reflex response which is a way of examining the integrity of the autonomic nervous system in visual function that could be affected due to various ocular conditions. In this current study we are validating and describing an infrared videography set-up suitable for examination of the pupil.

LONG ABSTRACT:

Clinical assessment of the pupil appearance and pupillary light reflex (PLR) informs us of the integrity of the autonomic nervous system, and is commonly used in optometric, ophthalmological and neurological examinations. Current clinical pupil assessment is however qualitative, and can only detect gross defects. High frequency infrared (IR) pupillography has therefore been explored for human studies in neurological and ophthalmological fields to provide more quantitative data.

In this study we validate and describe an IR videography set-up suitable for examination of the guinea pig pupil, which provides quantitative data similar to the quality of previously described studies in human. This set-up consists of a laptop computer connected to an IR camera, a smart phone with a strobe light application, an array of IR LEDs, and an animal stand. The guinea pig’s left pupil is illuminated by the smart phone LED, and right pupil is recorded by the IR camera.

The experimental set-up is validated by obtaining pupil diameters that are comparable to measurements from a commercially available pupillometer in humans. Following the validation, we report the resting pupil diameter, constriction velocity, percentage constriction ratio, re-dilatation velocity, and percentage re-dilatation ratio of the guinea pigs at three months (juvenile), and seven months of age (adult). The protocol allows for detection of subtle changes in PLR: at juvenile age male guinea pigs were found to have larger resting pupil size and more efficient pupil mobility than females. In female guinea pigs there was an age-related decline in the efficiency of pupil light reflex.

This technique is effective, accessible, and easy to assemble. It allows measurement of pupil diameter and parameters of pupil mobility that were not previously possible in animal studies. Such a set-up may provide a foundation for the further development of an integrated system useful for research and clinical applications.

INTRODUCTION:

The pupillary light reflex is a unique physiological mechanism that allows the examination of the function of the autonomic nervous system (ANS) gated by a light stimulus 1, 2. Pupil assessment has been proposed in the evaluation of cholinergic deficiency syndromes such as Alzheimer’s disease and Parkinson’s disease 3. An abnormal pupil light reflex may also be suggestive of lesions along the ANS. Examples of abnormal pupil reflex resulting from neurological problems including Adie’s tonic pupil, Argyll Robertson Pupil, and Horner’s syndrome 2. In a clinical setting pupil light reflex is examined by observation of the direct response elicited by light in the ipsilateral eye, and consensual response of the contralateral eye 2. Other relevant parameters examined include resting pupil size, and the relative afferent pupillary response 2. These observations are made with the use of a bright pen torch shone into the eyes, and rely heavily on the judgment of an experienced clinician. Such techniques are also empirical and provide qualitative information only. Pupil size measurement by means of imaging in animals has been described in the literature 4-7 but in those studies the parameters measured were limited by the experimental set up or by the way of data analysis. A more accurate and quantitative way of examining the pupils may be achieved by infrared videography, that allows the pupil constriction velocity, re-dilatation velocity, pre and post illumination pupil sizes to be calculated 3, 5, 8.

In recent years it has become increasingly common to utilise smart devices in ocular assessment due to their portability, ease of access, and versatility in the adjustment of colour and brightness display. Currently smart phones are used to assess colour vision, astigmatism, pupil size, macula integrity by the Amsler grid test, and fundus photography 9-11. Hence, an infrared videography set-up coupled with a smart phone application seems an accessible approach in the assessment of the pupil light reflex in the clinical setting 3, 12, 13. Such quantitative method is also more informative than the conventional technique using a pen torch, as the information could be used for future analysis and comparison.

In this investigation we describe the development of a pupillary light reflex measurement protocol that has potential to be used in both human and animal studies. For the purpose of this current study, resting pupil diameter is measured in human subjects, and pupil light reflex is studied in guinea pigs, using an experimental set-up not previously described in the literature. The infrared video pupillography was combined with a real-time light on-off stimulation protocol to allow evaluating constriction and re-dilatation response. We describe in the protocol that we are able to objectively measure resting pupil diameter, as well as parameters related to dynamic pupil response. We also describe in our representative results that there are gender-specific and age-specific differences in pupil dynamics in the guinea pigs.

PROTOCOL:

- Demonstration of experimental set-up. Measurement of resting pupil diameter in humans

  - Before beginning the experiment, position the patient’s head in front of the slit lamp biomicroscope and measure the horizontal visible iris diameter with the slit beam.

  - Dark adapt the patient for 2 minutes (in dark conditions the room light level should be ≤0.01 cd/m2).

  - The patient’s right pupil diameter is measured by the Neuroptics pupillometer (model: 59002, Neuroptics Inc, Irvine, CA, USA), and recorded (Fig. 1B).

  - The patient is further dark adapted for 2 minutes.

  - A firefly USB 2.0 infrared camera (Model #FMVU-03MTM-CS, 752x480 pixels, Point grey research, Canada) is attached to the eye piece of a slit lamp biomicroscope, with which the focus of the image can be adjusted by pushing the joystick back and forward.

  - The patient is again positioned appropriately in front of a slit lamp biomicroscope and instructed to fixate on a red light on the wall approximately 6 m away.

  - An infrared illumination system (5x 5 mm IR LEDs with peak spectral wavelength 940nm) is placed above the right eye to achieve high contrast filming conditions (Fig.1A).

  - The right pupil is recorded for 10 seconds in audio visual interleave format using Point grey FlyCapture program at 30 frames per second (fps).

  - A custom Matlab program (Mathworks; Natik, MA, USA) is used to define the pupil by a best-fit ellipse, formed by the points where the highest contrast is detected (Fig. 1C).

  - The known horizontal visible iris diameter measured in protocol 1.1 is used to set the scale in our custom Matlab program as a reference for pupil diameter measurement.

  - The centre of the pupil was marked with a white cross as a reference point to allow optimal detection of the pupil margin, and the corneal reflex was covered by a dark patch to avoid the program confusing the border of corneal reflex as pupil border.

  - The custom Matlab program outputs a spreadsheet with each frame number corresponding to a pupil diameter value. All measured values during the 10 s video are averaged to obtain the resting pupil diameter.

  - Data generated with the IR camera is compared with the pupilometer measurements in 1.3.

- Validation of experimental set up. Pupil light reflex measurement in guinea pigs

  - The recording station is set up prior to guinea pig handling. An animal stage is placed in the centre of the station, with the IR camera and IR illumination (as in 1.7) placed on the Right of the stage.

  - The smart device strobe light stimulus (spectrum 405-780 nm white light, oval-shaped, 2x4.5mm) is set to have 0.2 second light ON - 5 seconds light OFF using the iStrobe LED Flash app (Work Avoidance Ltd). It is placed on the Left of the animal stage (Fig. 2A).

  - The Right horizontal palpebral aperture (distance from temporal to nasal canthus) of the guinea pig is measured by a ruler (Fig.2B).

  - The guinea pig is gently held on an animal stage that allowed fine adjustment of lateral and vertical position.

  - The guinea pig is left on the stage for 5 minutes to familiarize and adapt to the surrounding and the researchers. This is followed by a 2 minute dark adaptation period.

  - The guinea pig is gently held and placed on the animal stage so its Left eye is 5 cm away from the smart device white LED light stimulus, and its Right eye is in front of the IR camera.

  - The IR camera is controlled by a laptop computer, and records for 10 s before the first white LED light flash stimulates the Left pupil. From then on, the 0.2 second light ON - 5 seconds light OFF cycle is repeated 7 times (Fig.2C).

  - The video is analyzed using the custom Matlab program as in protocol step 1.9-1.12.

  - The following parameters are measured and calculated:

2.9.1) Resting pupil diameter = Average across resting period* (see Fig. 3a)

2.9.2) Constriction velocity = Δ diameter/Δ time (see Fig. 3b)

2.9.3) % constriction ratio = (Constricted pupil diameter/Resting pupil diameter) x 100

2.9.4) Re-dilatation velocity = Δ diameter/Δ time (see Fig. 3c)

2.9.5) % re-dilatation ratio = (Re-dilated pupil diameter**/Resting pupil diameter***) x 100

*resting period: the amount of time elapsed between the end of dark adaptation to the time point just before the first cycle of pupil stimulation by white LED light.

**Re-dilated pupil diameter: the pupil diameter reached after a period of recovery (5 seconds light OFF) following a light stimulation cycle.

***Resting pupil diameter: the pupil diameter measured at the beginning of the videography recording, before stimulation of the pupil by the white LED light.

REPRESENTATIVE RESULTS:

- Comparison of  resting pupil diameter measurements using standard pupillometer and IR camera in human subjects

To determine whether using an IR illumination system (Fig.  1A) results in comparable results to standard pupillometers, we obtained repeated measurements of the pupil diameter using the NeurOpticsTM pupillometer (Fig. 1B) and compared the values with measurements obtained with the IR illumination system in three subjects (Fig. 1C). The results shown in figure 1D indicate that there are no significant differences with measurements from our experimental set-up (Student’s t-test, p>0.05).

- Pupil light reflex measurement  in guinea pigs shows age and gender differences

The IR videography set up was used to record pupillary light reflex (PLR) followed by video image analysis using a custom-made MATLAB program. The pupil reaction was analyzed at 2 times points in development (3 months and 7 months-old) and in female and males.

  - Gender differences at juvenile age

Male guinea pigs had larger resting pupil size than female guinea pigs (Male=7.42 ± 0.20mm vs. Female= 5.43 ± 0.32mm, P<0.05, Student’s t test). In terms of PLR kinetics, juvenile male guinea pigs had greater constriction velocity (M=-2.02 ± 0.18mm/s vs. F=-1.13 ± 0.10mm/s, P<0.01, Student’s t test), and also quicker PLR recovery kinetics, with greater re-dilatation velocity (M=0.41 ± 0.05mm vs. F=0.22 ± 0.02mm, P<0.001, Student’s t test). The juvenile male guinea pigs had both quicker pupil constriction and re-dilation velocity suggestive of a more readily mobile state of both the sphincter and dilator pupillae on stimulation of a PLR (Table 1).

  - Age specific effect in female guinea pigs

An age specific effect was detected in female guinea pigs when 3 months-old and 7 month-old data was compared. There were significant differences (Student’s t-test) in all parameters between the female guinea pigs at 3 months (n=5), and 7 months (n=4). Resting pupil size was larger in the adult female group than the juvenile female group (J=5.43 ± 0.32mm vs. A=6.93 ± 0.18mm, P<0.05, Student’s t test). Constriction velocity was slower (J=-1.13 ± 0.10mm/s vs. A=-0.88 ± 0.08mm/s, P<0.01, Student’s t test), and percentage constriction ratio was larger in the adult female group (J=85.13 ±0.89% vs. A=88.65 ± 0.77%, P<0.001, Student’s t test); (Table 2). This may be suggestive of an age-related decline in parasympathetic input as the mobility of sphincter pupillae was acting more slowly, and constricted to a lesser extent in the adult animals, and that the sympathetic input to the dilator pupillae muscle was relatively more unopposed. Such age-related effect in pupil constriction parameters were not observed in male guinea pigs.

DISCUSSION:

Pupillary Light Reflex is part of a routine examination for the integrity of the nervous system; it evaluates the state of the autonomic response to light, allowing us to see how efficiently the sphincter pupillae reacts, and how well the dilator pupillae counteracts the sphincter pupillae post illumination. Current pupil studies in animals utilize good quality commercial IR video cameras that are capable of recording at up to 60 fps 5, 14. The temporal frequencies of the video recording protocol across animal studies are however highly variable, with Taylor et al’s study recording at much lower temporal resolution - 1 frame every 2 seconds 14. The inclusion of pupil kinetic in our system - the rate of pupil change during the process of constriction and re-dilatation is not usually done in existing animal studies. This is a limitation of previous studies and is an advantage in our current set-up given that pupil studies involving human participants often take more kinetic parameters into account, such as PLR latency, constriction velocity, acceleration etc. Fotiou et al 3 concluded in their study that maximum pupil constriction velocity and acceleration were the best predicators in separating subjects with cholinergic deficiency from normal subjects, while baseline pupil diameter and minimum pupil diameter after pupil reaction to light were not significantly different from controls. This implies that parameters of iris muscle innervation are more informative than pupil diameter in evaluating PLR. Fotiou et al’s IR video camera recorded at up to 263 fps, and had high quality light sources, one IR source to illuminate the participant’s face (made of a 32 LED array), and one light stimulus stimulating the pupil with diffuse white light flashes 3. This set-up is superior to those described in animal studies due to its high sampling rate, good quality videos with high contrast due to the bright IR LED illumination, and a bright light stimulus.

The experimental set-up and analysis described in our study is more informative than the ones described in other animal studies 4-7, 14 as we have measured pupil mobility parameters as well as pupil diameter. Although we recorded our videos at a frame rate (30 fps) more similar to the other animal studies the sampling rate was every 33 ms, and is thought to be sufficient for calculating constriction velocity, and re-dilatation velocity. However, future studies with videography taken at frame rates comparable with human studies will give us insight in how high the frame rate might need to be to achieve optimal pupil mobility measures.

There are limitations to our techniques, and the two light sources in particular can be improved for future development. One potential drawback of our light stimulus for initiating PLR is that the white LED light of the smart device was a point source, so accurate positioning of the light was necessary to ensure the same amount of light reached the eye across all animals. A LED diffuser lens fitted in front of the smart device LED will ensure that the stimulus can illuminate the pupil more evenly, even if the position of the stimulus alters during experiment, especially with animal movement. And secondly, a brighter IR illumination system would allow better illumination of the guinea pig pupil for a higher-contrast video, which would enable our custom MATLAB software to better discriminate pupil margin for data analysis. The IR illumination could also be integrated with the camera - i.e. an array of ~20 IR LEDs can be soldered onto an electrical board and fitted around the camera lens by a collar.

The IR videography techniques described in this current study is an effective, accessible, and easy to assemble set-up, that allows measurement of pupil diameter as well as parameters of pupil mobility, that were not previously presented in animal studies. We analyzed our data by these parameters that are more comparable with human data 3, and may assist in transitional research looking at both animal models and human participants – for example, whether animal models of cholinergic deficits (PD and AD) also have deficits in PLR like in human. PLR is becoming an increasingly popular tool in neurological and eye research, contributing to the examination of the autonomic nervous system and retinal ganglion cells. Such experimental set-up described in this current paper may provide a foundation for further development of a more integrated system, which can be used in research as well as in clinical veterinary applications.

ACKNOWLEDGMENTS:

The authors would like to acknowledge the New Zealand Association of Optometrists for funding The Paul Dunlop Memorial Research Scholarship to assist in this research project. We would also like to acknowledge the Dick Roberts Community Trust through the Neurological Foundation of New Zealand (1229 SPG).

DISCLOSURES:

The authors have nothing to disclose.

REFERENCES

1. McLeod JG and T. Disorders of the autonomic nervous system: Part 2. Investigation and treatment. Ann Neurol 1987; 21: 519-529.

2. Friedman NJ, Kaiser PK, Pineda R, et al. The Massachusetts Eye and Ear Infirmary illustrated manual of ophthalmology. 3rd ed.. ed. Philadelphia, Pa.]: Philadelphia, Pa.], 2009.

3. Fotiou DF and S. Cholinergic deficiency in Alzheimer's and Parkinson's disease: Evaluation with pupillometry. International Journal of Psychophysiology 2009; 73: 143-149.

4. Dabisch P and H. Gender difference in the miotic potency of soman vapor in rats. Cutaneous And Ocular Toxicology 2008; 27: 123-133.

5. Mohan K, Kardon RH, Harper MM, et al. Retinal ganglion cell damage in an experimental rodent model of blast- mediated traumatic brain injury. Invest Ophthalmol Visual Sci 2013; 54: 3440-3450.

6. Ostrin L, Garcia M, Choh V, et al. Pharmacologically Stimulated Pupil and Accommodative Changes in Guinea Pigs. Invest Ophthalmol Vis Sci 2014; 55: 5456-5465.

7. Pennesi M and L. Extreme responsiveness of the pupil of the dark- adapted mouse to steady retinal illumination. Invest Ophthalmol Vis Sci 1998; 39: 2148-2156.

8. Miki A, Iijima A, Takagi M, et al. Pupillography of automated swinging flashlight test in amblyopia. Clinical Ophthalmology (Auckland, N.Z.) 2008; 2: 781-786.

9. Bastawrous A and C. iPhones for eye surgeons. Eye 2012; 26: 343-354.

10. Busis N, Busis N and Busis N. Mobile Phones to Improve the Practice of Neurology. Neurol Clin 2010; 28: 395-410.

11. Bastawrous A. Smartphone fundoscopy. Ophthalmology 2012; 119: 432-e2.

12. Fotiou D, Kosta V and Baloyannis S. Pupilometry as a useful functional procedure for differential diagnosis of Alzheimer's disease and vascular dementia. J Neurol Sci 2009; 283: 265-265.

13. Bergamin O, Kardon RH, Bergamin O, et al. Latency of the pupil light reflex: Sample rate, stimulus intensity, and variation in normal subjects. Invest Ophthalmol Visual Sci 2003; 44: 1546-1554.

14. Taylor JT, Davis E, Dabisch P, et al. Alterations in autonomic function in the guinea pig eye following exposure to dichlorvos vapor. Journal of Ocular Pharmacology and Therapeutics 2008; 24: 473-479.

Figure legends

Figure 1A. The IR videography set up for pupil diameter measurement in human. (a.) IR LED for pupil illumination (b.) slit lamp biomicroscope for visualisation of pupil (c.) IR camera attached to an observation tube showing what is seen through the slit lamp eyepiece (d.) laptop for video acquisition.

Figure 1B. (Left) Pupil diameter measurement using NeurOptics pupillometer. The display screen shows an image of the pupil and the diameter with standard deviation value at top left corner.

Figure 1C. (Above) Pupil diameter measurement in human using the experimental set-up described in this current study.(a.) the horizontal visible iris diameter, (b.) demarcation of pupil margin.

Figure 1D. Validation of experimental set-up by pupil diameter measurement in human 
The pupil diameter measured by Neuroptics pupillometer (light grey), and by our experimental set-up (dark grey) for subject A, B, C.

Figure 2A. The IR videography set up. 
The IR camera on the left (a.) is controlled by a laptop computer, a customized IR LED lighting source (b.) was placed close to the recorded eye of the animal on the animal stand (c.) and the flash was delivered by an IPhone flash light (d.) controlled by a strobelight app (I Strobe).

Figure 2B. Pupil diameter analisys by the custom Matlab program in guinea pig.
(a.) the horizontal palpebral aperture length, (b.) demarcation of the pupil margin is achieved by numerous blue crosses, indicating feature points detected by the program, and join up to be the best-fit ellipse. The white cross marks the pupil center. A hexagonal dark patch is created close to the pupil center to remove the corneal reflex from the video.

Figure 2C. A typical pupil diameter vs. frame number plot in a IR videography recording, showing (a.) pupil diameters at resting state (in black), (b.) negative pupil diameter change with light flash (in red), and (c.) positive pupil diameter change with cessation of light flash (in blue). Seven consecutive cycles are recorded for each video.

Table legends

Table 1. Male vs. female at juvenile age. Resting pupil diameter (D1), Constriction velocity (CV), Average constriction ratio (AvCR),  Re-dilatation velocity (DV), Average re-dilatation ratio (AvDR).

Table 2. Juvenile vs. adult female guinea pig. Resting pupil diameter (D1), Constriction velocity (CV), Average constriction ratio (AvCR),  Re-dilatation velocity (DV), Average re-dilatation ratio (AvDR).