Submission ID #: 53937

Editor: Steven Nilsen

Videographer: Kevin McRoberts

Film Date: Dec 16, 2015

Authors and Affiliations: 
James A. Fedchak, Julia Scherschligt and Makfir Sefa

Sensor Science Division, Thermodynamic Metrology Group

National Institute of Science and Technology

Gaithersburg, MD, USA
Title: How to Build a Vacuum Spring-Transport Package for Spinning Rotor Gauges
Corresponding Author: 
James.Fedchak@NIST.gov
Co-authors:
Julia.Scherschligt@NIST.gov
Makfir.Sefa@NIST.gov
1. Brief Questionnaire:
A.
i) Does your protocol involve microscopy, such as filming a complex dissection or a microinjection technique, which will need to be filmed by JoVE? (Y/N) _N__ 


B.
Does your protocol include detailed, step-by-step, descriptions of software usage? (Y/N) _N____ 
C.
Which steps of your protocol will viewers benefit most from having filmed?  


ITALCIZED
D.
What is the single most difficult aspect of this procedure and what do you do to ensure success?


UNDERLINED
E.
i) Will the filming need to take place in multiple locations? (Y/N) _N__

1. Introduction
A. Experimental Goal

The overall goal of this procedure is to build a spring-transport mechanism for the transport of spinning rotor gauges.  The design presented here has been tested for robustness and provides the user with better long-term stability of the calibration factor. (Intro)
B.  Interview
1.1. Fedchak: Spinning rotor gauges, or SRGs, are often used as transfer standards but damage to the rotor surface received during transport can change their calibration factor, also known as the accommodation coefficient.  
1.2. Fedchak: The advantage of using the spring transport mechanism is that the rotor is immobilized and under vacuum during transport, thus minimizing the possibility of changing the rotor surface.   
1.3. Sefa: Spring transport mechanisms have been used by a few national metrology institutes for more than 20 years, but these examples have been known to fail and their design details remain unpublished.    
1.4. Scherschligt: We realized that it is important both to build a robust spring-transport mechanism, and to properly package it in a way that minimizes shock during transport.
Protocol (read by voice talent at JoVE):
2. Assembly of Spring Transport Mechanism: Part 1
2.1. The text protocol reviews several precursor steps to preparing the spring transport mechanism, [2.1.1-WID] which include fabricating the parts for the mechanism [2.1.2-MED] and cleaning out the vacuum components. [2.1.3-MED]

2.1.1. representative action to fabricating rotor parts [Steps 2.1.1 and 2.1.2 could not be shot as these parts were fabricated outside of the lab.  We suggest panning over the mechanical drawings, Fig 2, fig. 3 and Fig 4]
2.1.2. 2nd representative action to fabricating rotor parts

2.1.3. representative action of cleaning vacuum components

2.2. For the assembly, wear nitrile, vinyl or latex gloves [2.2.1-MED] and never accidently handle the vacuum parts bare-handed.  Several tools are needed; [2.2.2-MED] be sure they are cleaned with ethanol or follow the procedure as described for vacuum parts. [2.2.3-MED]
2.2.1. talent putting on gloves

2.2.2. pan over the tools, talent can be setting them in place for washing, wearing gloves.  Tools include: needle nose pliers, socket head driver for 8-32 nut, wrench for adapter, ruler or micrometer caliper 
2.2.3. cleaning tools with ethanol
2.3. Additional parts include the valve, thimble, ball-bearing, [2.3.1-CU] 18-mm 8-32 studs, thread adaptors, nuts and lock washers, 3 mm Allen wrench and masking tape.  [2.3.2-CU] With everything in order, proceed with the assembly. [2.3.3-MED]

2.3.1. laying out valve, thimble and ball-bearing for use

2.3.2. laying out studs, thread adaptors, nuts and lock washers for use

2.3.3. pan over all the washed parts and smaller parts, ready for assembly
2.4. Begin the assembly by closing the right-angle valve. [CU]

2.4.1. turning the valve counter-clockwise until it stops.

2.5. Looking through the valve ports, identify the M6 stud centered on the valve seat. [CU]

2.5.1. view through valve ports, M6 studs visible in valve seat
2.6. Then, screw the adapter snugly onto the M6 stud. [2.6.1-CU] Use a UHV-clean wrench or pliers to get the adapter as snug as possible. [2.6.2-MED]

2.6.1. threading adapter onto stud.

2.6.2. Tightening the adapter onto the stud
2.7. Next, insert the 8-32 stud into the other end of the adapter [2.7.1-CU] and tighten it as snugly as possible.  The 8-32 stud should extend through the adapter [2.7.2-MED] and push against the M6 stud. [2.7.3-CU]

2.7.1. inserting stud into adapter

2.7.2. tightening stud in adapter

2.7.3. view of 8-32 stud contacting M6 stud, once tightened
2.8. Scherschligt:  By making certain that the 8-32 stud pushes firmly against the M6 stud on the valve seat, we ensure that the adapter will be held firmly in place. [MED/WID]

2.8.1. interview shot at bench
2.9. [2.8.1-MED] Place a lock-washer on the 8-32 stud and tighten a nut onto it using a clean socket driver or pliers. [2.8.2-CU]

2.9.1. film as written

2.9.2. film as written

2.10. Next, screw a second nut onto the 8-32 stud until it nearly touches the first. [CU]

2.10.1. film as written, be sure to view final position of nut

2.11. Then, place a lock washer on the 8-32 stud by the standoff followed by the spring/ball-holder [2.10.1-CU] and tighten the holder until it compresses the lock-washer. [2.10.2-MED]
2.11.1. attaching the lock washer and spring /ball holder

2.11.2. tightening down the lock washer
2.12. Now check the length adjustment of the spring assembly. [LM]
2.12.1. Figure 6

2.13. First, place the rotor in the custom thimble. [CU]

2.13.1. film as written

2.14. Then attach a magnet to the custom thimble exterior [2.13.1-CU] and fix it in place with masking tape. [2.13.2-CU]

2.14.1. putting magnet in place

2.14.2. taping down magnet

2.15. Next, slide the thimble over the spring/ball-holder until the ball-holder touches the ball [2.14.1-CU] and measure the gap between the valve flange and thimble flange. [2.14.2-MED]

2.15.1. film as written

2.15.2. using tool to measure gap

2.16. If the gap is in the range of 2 to 6 mm, it is acceptable.  [CU]

2.16.1. view of tool reading showing a value not quite 2-6 mm

2.17. Otherwise rotate the standoff to adjust the gap.  3 mm is ideal. [MED]

2.17.1. rotating standoff to adjust and taking a new measurement
2.18. Sefa: Making certain the gap is within the 2 to 6 mm range ensures the mechanism will function as designed and will hold the ball with sufficient force.

2.18.1. Interview shot at bench
2.19. Next, remove the thimble/flange assembly and set it aside. [MED]

2.19.1. film as written

2.20. Then, tighten the nut against the standoff and open the valve. [MED]

2.20.1. tightening nut and opening valve
3. Assembly of Spring Transport Mechanism: Part 2
3.1. Continue the assembly by attaching the two tines to the flange.  [MED]

3.1.1. film as written

3.2. Secure them with a lock-washer and bolt, provided by the commercial spinning rotor gauge.  [CU/TEXT]

3.2.1. film as written, TEXT: The lock-washer and bolt do not need cleaning.

3.3. Remove the rotor by removing the masking tape and magnet and carefully allowing the rotor to roll out of the thimble [3.3.0].  Now, rotate the tines so that they are square with one another [3.3.1-CU] and check their straightness by slipping the spinning rotor gauge head over the thimble. [3.3.2-CU]  The head should slide on freely. [3.3.3-CU]

3.3.0
added
3.3.1. film as written

3.3.2. film as written

3.3.3. spinning the head

3.4. Then, give the two bolts a final tightening and re-check for alignment. [MED]

3.4.1. tightening bolts and checking alignment

3.5. Place the rotor back into the thimble, and secure it with the magnet and masking tape [3.5.0].  Open the valve by turning counter-clockwise.  The next step is to place a copper or silver-plated copper gasket on the valve port. [ECU]

3.5.0
added
3.5.1. Fit the gasket onto the port

3.6. Place the rotor into the thimble if it’s not already there. [CU/MED]

3.6.1. film as written

3.7. Now, slide the flange/thimble assembly over the spring/ball-holder. [CU]

3.7.1. film as written

3.8. Orient the flange so that the set-screw in the tines will point downwards when the valve is attached to the chamber. [CU]

3.8.1. show rotation of tines with a view that illustrates that they will point down when the valve is attached 

3.9. Finally, using one-quarter-number-28 bolts and nuts, secure the flange to the valve. [CU]

3.9.1. film as written, tightening flange to valve

3.10. Then, close the valve [3.10.1-ECU] and remove the magnet and tape holding the ball. This completes the spring transport mechanism assembly. [3.10.2-CU]

3.10.1. film as written

3.10.2. removing magnet and tape holding the ball
4. Using the Spring-Transport Mechanism
4.1. To use the spring-transport mechanism, it must be properly mounted and the rotor must be properly suspended.  This should only be performed by those who know how to use the spinning-rotor gauge, the controller electronics, and are familiar with high-vacuum technology. [LM]

4.1.1. Fig 9

4.2. The first step is to unpack the spring-transport mechanism from its shipping container. [WID]

4.2.1. unpacking the transport mechanism 

4.3. Then, remove the end-cap and foil from the open port. [MED]

4.3.1. film as written 

4.4. Next, using a new copper or silver-plated copper gasket [4.4.1-CU] and a one-quarter-number-28 bolt set, attach the open port to a DN 40 port on the vacuum chamber. [4.4.2-MED/TEXT]  

4.4.1. attaching gasket to DN 40 port of vacuum chamber

4.4.2. attaching bolt and securing the mechanism to the vacuum chamber, TEXT: CF 2.75”

4.5. The orientation of the valve will depend on the orientation of the tines.  When oriented properly, the tine with the set screw points downward. [LM]

4.5.1. Figure 9 - proper orientation

4.6. Then, slip the head over the thimble [4.6.1-CU] and orient the thimble vertically to within two degrees of accuracy. [4.6.2-CU]

4.6.1. film as written

4.6.2. adjusting orientation and measuring check with a tool level 
4.7. Now, evacuate the vacuum chamber to less than one-one-thousandth Pascal. [CU]

4.7.1. pressure gauge showing it drop to < 1/1000 Pa

4.8. Then, slowly open [4.8.1-CU/ECU] the spring-transport mechanism valve.  [4.8.2-CU/ECU]

4.8.1. film as written

4.8.2. 2nd angle 

4.9. Finally attach the head to the controller, [4.9.1-MED] turn on the controller and suspend the ball. [4.9.2-MED]

4.9.1. attaching head

4.9.2. turning on controller and suspend the ball
5. Results: Proper Packing and Drop Tests 
5.1. After assembling the spring transport mechanism, it was packed in foam inserts. 

5.1.1. Figure 7 
5.2. The shipping container allows for 7.6 cm of foam surrounding the mechanism.  Ester foam was packed above and below and PE foam was packed on the sides.  

5.2.1. Figure 8 

5.3. In this configuration, the shock experienced by the spring-transport mechanism was measured at less than 100 G even when dropped from 152 cm.

5.3.1. Figure 10

5.4. If the foam was compressed during packing, even by a small amount, the shock increased by about 40%. In this case, the foam was compressed roughly 1 ¼ cm over a span of 22 ¾ cm. 

5.4.1. Figure 11

5.5. After performing over 180 drop tests of one prototype model, a neutron radiograph image was made of the mechanism.  The mechanism still worked – the ball remained secured to the thimble with the valve closed.  Note that the valve on this prototype is smaller than the one in the protocol and is no longer manufactured.

5.5.1. Figure 12
6. Conclusion (said by authors on camera)
6.1. Sefa: Once all the parts have been procured, the spring-transport mechanism can be assembled in just a few minutes.
6.2. Scherschligt: It is important to procure the custom thimble given in our design, the commercial thimbles do not have the proper length or interior geometry to completely immobilize the ball.  Step files of our custom parts are available as supplemental material.
6.3. Fedchak:  As described in our manuscript, it is critical to transport the spring-transport mechanism in a shipping container with at least 7.6 cm of foam packing surrounding it. This ensures that the rotor will remain immobilized during common shipping mishaps.
Provided Media
2.12 – Figure 6 – Setting the Distance.pdf
4.1 – Figure 9 – STM on Chamber.pdf
4.5 – Figure 9 – STM on Chamber.pdf
5.1 -- Figure 7 – Foam Assembly.pdf
5.2 – Figure 8- STM in package.pdf
5.3 – Figure 10 – Shock Versus Drop Height.pdf
5.3 – Figure 11 – Effect of Compressed Foam.pdf
5.4 – Figure 12 – Radiograph.pdf
General Preparation
It’s critical for a smooth and organized shoot that all reagents are accounted for, in advance.   

Any overnight or long incubation steps should be recognized and specimens/samples be prepared in advance so that prior steps can be recorded and shooting can continue with pre-prepared specimens/samples.  

All tubes/flasks should be pre-labeled neatly before we arrive.  

Ex. Luciferase assay done in 96 well plates should be labeled with negative/positive control wells and experimental samples are labeled accordingly.

You will receive more detailed preparation instructions are included in the email accompanying the finalized script.
 2015, Journal of Visualized Experiments


