Submission ID #: 55276
Editor Name: Michael Linnes
Videographer name: Nastasha Drobnica
Film Date: Nov. 7th, 2016
Authors and Affiliations: Kara Brower1,2,4*, Adam K. White1,2*, and Polly M. Fordyce1,2,3,4
1Department of Bioengineering 2Microfluidic Foundry 3Department of Genetics  4 Chem-H Institute, Stanford University, Stanford, CA, USA 
*These authors contributed equally to this work. 

Title: Multi-Step Variable Height Photolithography for Valved Multilayer Microfluidic Devices 
Corresponding Author: 

Polly M. Fordyce 

pfordyce@stanford.edu
Additional author email:

Kara Brower kbrower@stanford.edu
Adam K. White akwhite@stanford.edu
Questionnaire: 
A. Will you require JoVE to record video microscopy, such as filming a complex dissection or microinjection technique? (Y/N)_N  
(If you can record images/videos using your own camera/software, then mark No)
B.   Does your protocol include detailed, step-by-step, descriptions of software usage? (Y/N) N

C.  Which steps of your protocol will viewers benefit most from having filmed? Please list 4-6 individual steps using the step numbers listed in this document. (Please do not list entire sections.) 
2.3, 2.7, 2.8, 2.9, 3.5, 3.6, 3.9, 4.1, 4.3, 4.10
D.  What is the single most difficult aspect of this procedure and what do you do to ensure success?  Please list 1-2 individual steps using the step numbers listed in this document. (Please do not list entire sections.) 2.8, 2.9, 3.5, 3.6
E.  Will the filming need to take place in multiple locations? (Y/N) Y 
If yes, how far apart are the locations? Walking distance, 3 minute walk
1. Introduction (Experimental Goal and Author Interviews) – As the beginning of your video, the introduction should clearly present the goal of your method to the viewer and its significance.  Other information can be provided according to the various statements below, but the total introduction should not exceed 150 words. 
A. Experimental Goal (read by voice talent at JoVE): 
The overall goal of this video protocol is to demonstrate complete multi-step photolithography of microfluidic master molds with on-chip valves and multiple height features tunable for any application. (Intro)
B.  Required Interview Statements: (Said by you on camera. Don’t forget to smile!)   
1.1. Polly Fordyce: This method is a complete overview of how to fabricate master molds with complex geometries, including on-chip membrane valves, for microfluidic devices. 
1.2. Kara Brower: The main advantage of this technique is that it makes it possible to easily control flow on microfluidic devices, overcoming the barrier of entry to using microfluidic devices in biological research.  
C.  Optional Interview Statements: (Said by you on camera. Don’t forget to smile!)   
1.3. Adam White: Visual demonstration of this method is critical as the photolithography steps are often difficult to master for beginners, because proper alignment, development, and exposure all rely on visual cues and cleanroom experience.    
Protocol (read by voice talent at JoVE):  
2. Rounded Valve Fabrication

2.1. To begin, design your device and prepare the individual [1-MED Over the Shoulder] photomasks for the multilayered geometries.  [2-LM] Additionally, prepare about 4 wafers with a 5 micron layer of SU-8 2005 negative photoresist and flood expose as described in the accompanying text protocol. [3-MED-TXT]
2.1.1. Talent at computer working with feature design software. 

2.1.2. Figure 2 part 2 (photomasks) 
2.1.3. Talent performs the flood exposure on one of the wafers. (TEXT: *See accompanying text protocol for details)
2.2. Place the coated wafer on a spin coater and turn on vacuum to affix it to the spin chuck. [1-CU]  Use nitrogen or compressed air to blow away any dust from the surface. [2-CU]
2.2.1. *Film as written
2.2.2. *Film as written

2.3. Then, apply 2-3 mL of AZ50 XT positive photoresist to the center of the wafer. [1-CU] Spin coat the photoresist to create a 55 micron layer. [2-CU-TXT]
2.3.1. *Film as written 
2.3.2. *Film as written (TEXT: Spread: 200 rpm, 10s, 133 rpm/s acceleration; Cast: 12000 rpm, 40 s, 266 rpm/s acceleration; Snap spin: 3400 rpm, 1 s, 3400 rpm/s acceleration.)

2.4. Once coated, lay down the wafer carefully in a 5 inch petri dish and let relax for 20 minutes. [1-MED] Next, soft bake the wafer on a hotplate for 22 minutes while ramping the temperature from 65°C to 112°C at a rate of 450°C per hour.[2-CU-TXT] 
2.4.1. Talent sets wafer into petri dish.

2.4.2. Talent sets wafer on preheated hotplate. (TEXT: 65°C - 112°C, 450°C/hr for 22 min)
2.5. Then, remove the wafer and let it rest overnight at room temperature in a petri dish for ambient rehydration. [1-MED]
2.5.1. Talent removes the wafer and places it in a petri dish (and covers if applicable). 
2.6. Tape the flow round transparency mask to the 5” glass plate so that the print-side is closest to the wafer and load into the mask positioner of the UV mask aligner.[1-CU-LM] Expose the wafer to 930 mJ of UV in 6 cycles. [2-CU-TXT] 
2.6.1. *Film as written (Video Editor: Show the “Flow Round” mask in figure 2 as an insert briefly during the 2nd half of the sentence.) AUTHOR NOTE: “[became 2-OS-MED, use Stanford 70726 32K Jove v1.1 Mask File 1 checkplot 10x.pdf]”
2.6.2. Show wafer as it is being exposed to UV light. (TEXT: 6 cycles: 25 s at 6.2 mW/cm2 / 30 s wait time) 

2.7. Develop the wafer immediately by immersing it in a stirred bath of developer for 3 to 5 minutes [1-CU-TXT] or until bath turns purple and the features emerge. [2-CU] Once developed, remove the wafer [3-MED] and rinse it well with deionized water [4-MED]. 
2.7.1. Talent places wafer in bath. (TEXT: Bath: 25 mL of AZ500k 1:3 Developer)
2.7.2. Try to keep the same frame as step 2.7.1 and show the end point for a few seconds with the bath the color described.

2.7.3. *Film as written
2.7.4.  Added Medium shot of rinse
2.8. Then, hard bake the wafer to melt and round valve features.  Ramp the temperature from 65 to 190°C over the course of 15 hours at a rate of 10°C per hour.  [1-MED] Once finished, turn off the hotplate and let the wafer cool to room temperature [2-CU]. The features on the wafer are now rounded [3/4-CU]
2.8.1. Talent places wafer on hot plate and starts temp ramp.

2.8.2. *Film as written.  First close up.
2.8.3. Close up after bake. Use either or both this and 2.8.4.
2.8.4.  CU.  Wafer comparison.
2.9. Kara Brower: “This hardbake is critical to properly reflow rectangular valve features into a smooth rounded profiles necessary for fully closed valves in multilayer microfluidic devices. Shorter times may cause feature cracking or instability.”[1-INT]
2.9.1. Author saying the above statement interview style. 
3. Fabricating Variable Height Features in Tandem

3.1. In order to fabricate a device with variable height features, place the cleaned wafer on spin coater as previously shown.  [1-MED] Apply 1 to 2 mL of SU-8 2050 negative photoresist to center of the wafer and spin the photoresist over the developed valve features.[2-CU-TXT]
3.1.1. Talent places wafer on the spincoater, fixes it in place, and blows nitrogen on it. 
3.1.2. Talent applies photoresist to wafer, then starts spin. (TEXT: Spread: 500 rpm, 10s, 133 rpm/s acceleration; Cast: 3000 rpm, 40 s, 266 rpm/s acceleration) 

3.2. Then, carefully place the spun wafer in a 5 inch petri dish and let it relax for 20 minutes on a flat surface or until any streaking patterns fade. [1-MED]
3.2.1. *Film as written
3.3. Next, preheat two hotplates to 65°C and 95°C and then set the wafer on the 65°C plate for 2 minutes, the 95°C plate for 8 minutes, and the 65°C plate for 2 additional minutes to soft bake the wafer. [1-CU]
3.3.1. Talent places wafer on 65°C plate, waits a few seconds, moves it to the 95°C plate, waits a few seconds, then moves it back to the 65°C plate. (Video Editor: Add a wipe between the movements to show time lapsing.)  (Author Note: this was shot in three scenes, 3.3.1, 3.3.2, and 3.3.3.) Nicola – I don’t know exactly what each shot shows so, you’ll need to piece these together as it makes sense.
3.3.2. Scene in 3.3 was split into 3 shots.

3.3.3. Scene in 3.3 was split into 3 shots.
3.4. Once the wafer cools back to room temperature, tape the Flow Low transparency mask to a quartz 5” glass plate so that the print-side is closest to the wafer [1-CU-LM] and load it into the mask positioner of the UV mask aligner. [3-MED]
3.4.1. Talent tapes the mask onto the plate. (Video Editor: Add Figure 2 part 2 the image of the “Flow Low” mask as an insert once mentioned.) AUTHOR NOTE: use Stanford 70726 32K Jove v1.1 Mask File 2 revised checkplot 10x.pdf]
3.4.2. *Film as written. Putting mask in aligner.
3.4.3.  Put wafer in aligner. Not sure how this differs from 3.4.2, so use as you see appropriate.
3.5. Then, place the wafer in a UV mask aligner chuck [1-CU] and, using microscope eyepiece or camera, carefully align the new Flow Low layer features to the Flow Round valve layer features. [2-CU]
3.5.1. *Film as written.   Author Note: 3.5.1 is the same as 3.4.2 so use that here, and not there?
3.5.2. Talent looks through eyepiece at alignment.
3.6. Begin by aligning the horizontal, vertical and tilt axes of device borders to the device border features on mask… Next, align cross-hair features between the layers…[1-SCREEN]. Finally, confirm that the valve features intersect the Flow Low features where desired. [2-MED Over the Shoulder]
3.6.1. Talent aligns the items in order.
3.6.2. Talent makes small adjustments while looking at sample through optics (or on screen).
3.7. Next, expose the wafer to 170 mJ UV deposition. [1-MED-TXT] When finished, remove the wafer and bake it post-exposure by switching between the two hotplates set at 65°C and 95°C. [2-MED-TXT]
3.7.1. *Film as written (TEXT: 28 s at 6.2 mW/cm2)
3.7.2. Talent removes the wafer and sets it on the first of 2 hot plates. (TEXT: 65°C: 2 min, 95°C: 9 min, 65°C: 2 min)

3.8. Without developing the wafer, allow it to cool to RT and then sequentially add the Flow High layer and then the Chaotic Mixer Herringbone layer using SU-8 2025 as described in the accompanying text protocol.   [1-MED-LMx2]
3.8.1. Talent places the wafer back onto the spincoater and adds the SU-8 2025 (Video Editor: Add inserts showing the layer masks in Figure 2 part 2 for both the “Flow High” and “Herringbone” layers and label them as such.) Author note: use Stanford 70726 32K Jove v1.1 Mask File 3 checkplot 10x.pdf; use Stanford 70726 32K Jove 1.1 Mask File 4 checkplot 10x.pdf
3.9. After all layers have been completed, develop the features by immersing the wafer in a stirred bath containing 25 mL of SU-8 developer for 3.5 minutes or until the features clearly emerge. [1-MED] Use a stereoscope to verify [2-CU] that the features have clear, defined feature boundaries. [3-CU]
3.9.1. Talent adds the wafer into the developer and sets a timer for 3.5 minutes.

3.9.2. video of talent pulling out final wafer.
3.9.3. Close-up of final wafer showing defined features. 
3.10. Adam White: During development, be sure to check every 20 seconds to see if features have become fully defined and the resist has washed away. Overdevelopment may cause feature damage, especially in complex molds. [1-INT]
3.10.1. Author saying the above statement interview style. 
3.11. Then, hard bake the wafer to stabilize all of the photoresist features. [1-MED-TXT] Subsequently, fabricate the control layer as described in the accompanying text protocol. [2-LM]
3.11.1. *Film as written (TEXT: 65°C – 165°C, 2 h 30 min, 120 °C/h ramping speed) 
3.11.2. Author note: This is now just a still image of the control mask [LM]: use Stanford 70726 32K Jove v1.1 Mask File 5 checkplot 10x.pdf
4. Production of Hydrogel Beads from Droplets 

4.1. Fabricate multilayer microfluidic devices in a “push up” geometry on glass according to existing open-access protocols [1-WIDE] and use visual inspection to ensure that all of the valves are properly aligned to control lines and that all of the inlets are punched fully before proceeding [2-MED]. 
4.1.1. Wide establishment shot.

4.1.2. Talent lifts completed device from a petri dish and visually inspects it from all angles.  Then, checks that everything is fully punched. (One could simply view, or place under microscope and look through eyepieces.)

4.2. Connect Tygon tubing loaded with water to a flow control system such as a syringe pump, fluidic controllers, or an open-source solenoid valve array with reservoirs. [1-CU]
4.2.1. Talent connects water filled tubing to the control system. (Note: Take 1 is a wide shot.  Take 2 is wide, second attempt).
4.3. Next, connect metal pins to the tubing and the metal pins to the device ports at control line inlets. [1-CU] Then, set the flow control system to 25 psi for each line to pressurize the device control lines. [2-CU] Ensure that valves close and re-open by inspection under the microscope. [3-MED] 
4.3.1. *Film as written
4.3.2. *Film as written
4.3.3. *Film as written.  MED shot placing it under microscope and opening software. Author note: Not scope, screen was filmed,; have already sent in video labeled by step number
4.3.4.  Video of screen using software. Use at your discretion. I’m not sure how this differs from 4.3.3
4.4. In a microcentrifuge tube, suspend 3.9 mg of photoinitiator in 100 µL of DI water to prepare the photoinitiator solution used to polymerize droplets into hydrogel beads. [1-CU-TXT] Cover the solution to protect it from light. [2-CU] 
4.4.1. *Film as written (TEXT: Photoinitiator: LAP @ 39 mg/mL)
4.4.2. *Film as written 
4.5. In a second microcentrifuge tube, add 132 µL of deionized water… 172 µL of PEG diacrylate… 12 µL of the photoinitiator solution… and 85 µL of HEPES buffer to make the hydrogel droplet solution. [1-CU]
4.5.1. Talent adds items to the microcentrifuge tube in the order listed.
4.6. Transfer the hydrogel droplet solution to a customized cryogenic tube vessel. [1-CU-TXT] Then, connect the tubing of the cryogenic tube vessel to a controllable pressure source and connect the PEEK tubing to the device reagent inlet.[2-CU] 
4.6.1. *Film as written (TEXT: *See accompanying text protocol for details on vessel preparation) 
4.6.2. *Film as written
4.7. Next, insert the PEEK tubing at the device outlet in order to collect the droplets... [1-CU] Remove air bubbles from the device, repressurized the system, [2-CU] and then depressurize the Ro1 oil valve and set the oil pressure to 10 psi. [3-CU]
4.7.1. *Film as written
4.7.2. *Film as written (Author note: we filmed a screen capture but have also provided a video of the screen capture uploaded to the website).- so use whichever looks best
4.7.3. *Film as written
4.8. Next, set the PEG mixture pressure to 9 psi,[1-MED over the shoulder] depressure the upstream valves and adjust the pressure as necessary to produce droplets of the desired size. [2-MED over the shoulder] Determine the droplet size via microscopy using a camera with 50 fps or higher. [3-MED Over the Shoulder]
4.8.1. *Filmed over the shoulder
4.8.2. *Filmed screen.  (Author note: we filmed a screen capture but have also provided a video of the screen capture uploaded to the website).- so use whichever looks best
4.8.3. Talent looks at screen with image of droplets.
4.9. When the droplets have stabilized, position a UV light source over the polymerization region of the device and apply 100 mW/cm2 of 365 nm light from the source onto a 5 mm spot.[1-CU-TXT]
4.9.1. Talent moves the UV light into position and turns it on. (TEXT: 365 nm, 100 mW/cm2, 5 mm spot)

4.10. Pressurize the bead sieve valve [1-CU] to watch polymerized beads collect and ensure that droplets have hardened into beads.[2-SCREEN] [3-CU]
4.10.1. *Film as written
4.10.2. Bead sieve valve is pressurized and beads collect.

4.10.3. *Film as written

4.11. Finally, depressurize the bead sieve valve and collect beads into a tube through the PEEK outlet tubing. [1-CU]
4.11.1. *Film as written
5. Results: Production of PEG-diacrylate Hydrogel Droplets
5.1. This protocol starts by demonstrating a method for rounding flow valves. [1-LM] Here, a profilometer was used to determine the typical post-reflow valve rounding profile resulting from this method, showing a height of approximately 55 microns.[2-LM]
5.1.1. Figure 3, step 3

5.1.2. Figure 5a (Video Editor: Point to the peak of the graph with the word “height” through the end.)
5.2. In the image of the left, the valve is off and liquid can pass through the channels. [1-LM] Once activated, by pressurizing the valves, the flow through these valves is cut off. [2-LM]
5.2.1. Figure 5c-d (Video Editor: Place figure 5c on the left and 5d on the right.  Add arrows pointing to the two valve locations in figure 5c with the 2nd half of the sentence.)
5.2.2. Figure 5c-d (Video Editor: Place figure 5c on the left and 5d on the right.  Add arrows pointing to the two valve locations in figure 5d with the words “by pressurizing the valves”.)

5.3. Here, one can see the bead synthesizer device in operation, producing hydrogel droplets in an oil emulsion at the T junction droplet generator. [1-LM] By partially closing a downstream flow using a sieve valve, the fluid can continue to flow, but the droplets are trapped behind the valve.[2-LM]
5.3.1. 55276_video_5_3_1.avi (0:00-0:08)
5.3.2. 55276_video_5_3_2.avi (0:10-0:17)
5.4. The resulting beads produced using this process averaged 52.6 microns in diameter with a standard deviation of only 1.6 microns.[1-LM]  Out of nearly 3000 beads, less than 1% were off by more than 3 standard deviations.[2-LM]
5.4.1. Figure 6c

5.4.2. Figure 6d 
6. Conclusion (said by authors on camera) 
6.1. Adam White: Once mastered, this technique can be completed from design to testing in under 3 days if it is performed properly, allowing for fast device design iteration.

6.2. Polly Fordyce: Following this procedure, even researchers with little fabrication experience can build their own complex microfluidic devices and apply them towards their own biological question. 
6.3. Kara Brower: After watching this video, you should have a good understanding of how to perform the photolithography steps necessary to make microfluidic devices with any level of complexity including variable height features and valve structures.  
Provided Media

Authors, Please list all images, movie files, or 3-D rendered animations that can be included in the video per editor’s request.  The step in the script/video where these images will be inserted should be specified.   For example:

6.2 –  0123_PIname_Figure1.tif -  dual color imaging of tumor angiogenesis at 40X 

6.2 –  0123_PIname_Figure2.tif -  dual color imaging of tumor angiogenesis at 100X

Formats:  For static images we prefer .tiff, .eps, Illustrator, Powerpoint or Photoshop files at dimensions of at least 720X480 pixels and 300 dpi.  The higher resolution, the better.  Likewise any exported movie files should have at minimum these dimensions and be rendered to .mov, .mp4, or .avi files.  

(1) 55276_video_5_3_1.avi – Hydrogel Droplet Creation

(2) 55276_video_5_3_2.avi-Sieve Valve Bead Retention

General Preparation

It’s critical for a smooth and organized shoot that all reagents are accounted for, in advance.   

Any overnight or long incubation steps should be recognized and specimens/samples be prepared in advance so that prior steps can be recorded and shooting can continue with pre-prepared specimens/samples.  

All tubes/flasks should be pre-labeled neatly before we arrive.  

Ex. Luciferase assay done in 96 well plates should be labeled with negative/positive control wells and experimental samples are labeled accordingly.

You will receive more detailed preparation instructions are included in the email accompanying the finalized script.

( 2016, Journal of Visualized Experiments


