TITLE:

Merging Ion Concentration Polarization between Juxtaposed Ion Exchange Membranes for Blocking Propagation of the Polarization Zone

AUTHORS:

Minyoung Kim, Hyunjoon Rhee, Ji Yoon Kang, Tae Song Kim, Rhokyun Kwak

Kim, Minyoung1,2

1Center for BioMicrosystems

Korea Institute of Science and Technology (KIST)

Seoul, Republic of Korea

2Department of Mechanical Engineering,

Seoul National University

Seoul, Republic of Korea

t16012@kist.re.kr

Rhee, Hyunjoon1,3

1Center for BioMicrosystems

Korea Institute of Science and Technology (KIST)

Seoul, Republic of Korea

3Department of Industrial Engineering

University of Illinois Urbana-Champaign, Illinois

hyunjoonrhee@gmail.com

Kang, Ji Yoon1

1Center for BioMicrosystems

Korea Institute of Science and Technology (KIST)

Seoul, Republic of Korea

jykang@kist.re.kr

Kim, Tae Song1

1Center for BioMicrosystems

Korea Institute of Science and Technology (KIST)

Seoul, Republic of Korea

tskim@kist.re.kr

Kwak, Rhokyun1

1Center for BioMicrosystems

Korea Institute of Science and Technology (KIST)

Seoul, Republic of Korea

rhokyun@kist.re.kr

CORRESPONDING AUTHOR:

Kwak, Rhokyun (rhokyun@kist.re.kr; +82-2-958-6975)

KEYWORDS:

ion concentration polarization, preconcentration, ion exchange membrane, overlimiting current, electroosmotic flow, electro-osmotic instability

SHORT ABSTRACT:

We introduce an experimental protocol for the novel ion concentration polarization (ICP) platform that can stop the propagation of the ICP zone regardless of operating conditions. This unique ability of the platform lies in the use of merging ion depletion and ion enrichment, which are two polarities of the ICP phenomenon.

LONG ABSTRACT:

Ion concentration polarization (ICP) phenomenon is one of the most prevailing methods to preconcentrate low-abundance biological samples. The ICP induces a noninvasive region for charged biomolecules (i.e. ion depletion zone), and we can preconcentrate the targets on this region boundary. Despite the high preconcentration performances with the ICP, it is difficult to find operating conditions of non-propagating ion depletion zones. To overcome this narrow operating window, recently we developed a new platform for spatiotemporally fixed preconcentration. Unlike preceding methods of only using the ion depletion, this platform also utilizes the opposite polarity of the ICP (i.e. ion enrichment) to stop the propagation of the ion depletion zone. By confronting the enrichment zone with the depletion zone, the two zones merge together and stop. In this paper, we describe a detailed experimental protocol to build this spatiotemporally defined ICP platform, and characterize preconcentration dynamics of the new platform by comparing with that of the conventional device. Qualitative ion concentration profiles and current-time responses successfully capture the different dynamics between the merged ICP and the stand-alone ICP. In contrast to the conventional one that can fix the preconcentration location at only ~5V, the new platform can produce a target-condensed plug at a specific location in the broad ranges of operating conditions: voltage (0.5-100 V), ionic strength (1-100 mM), and pH (3.7-10.3).

INTRODUCTION:

Ion concentration polarization (ICP) refers to a phenomenon that occurs ion enrichment and ion depletion located on a permselective membrane, resulting in an additional potential drop with ion concentration gradients.1,2 This concentration gradient is linear, and it becomes steeper as a higher voltage is applied (Ohmic regime) until the ion concentration on the membrane approaches zero (limiting regime). At this diffusion-limited condition, the gradient (and corresponding ion flux) has been known to be maximized/saturated.1 Beyond this conventional understanding, when the voltage (or current) is increased further, overlimiting current is observed with flat depletion zones and very sharp concentration gradients at the zone boundary.1,3 The flat zone has a very low ion concentration, but surface conduction, electro-osmotic flow (EOF), and/or electro-osmotic instability promote ion flux, and induce overlimiting current.3-5 Interestingly, the flat depletion zone serves as an electrostatic barrier, which filters out6-9 and/or preconcentrates targets.10,11 Since there are insufficient amount of ions to screen the surface charges of charged particles (for satisfying electroneutrality), the particles cannot pass through this depletion zone and therefore line up at its boundary. This nonlinear ICP effect is a generic phenomenon in various types of membranes10-14 and geometries.6,15-21; this is why researchers have been able to develop various types of filtration6-9 and preconcentration10,11 devices using the nonlinear ICP.

Even with such high flexibility and robustness, it is still a practical challenge to clarify operating conditions for the nonlinear ICP devices. The nonlinear regime of the ICP quickly removes cations through a cation exchange membrane, which causes the displacement of anions to head toward the anode. As a result, the flat depletion zone propagates quickly, which is reminiscent of the shock propagation22. Mani et al. called this dynamics as the deionization (or depletion) shock.23 To preconcentrate targets at a designated sensing position, preventing the expansion of the ion depletion zone is necessary, for example, by applying EOF or pressure-driven flow against the zone expansion.24 Zangle et al.22 clarified the criteria for the ICP propagation in a one-dimensional model, and it highly depends on electrophoretic mobility,17 ionic strength,18 pH,25 and so on. This indicates that proper operating conditions will be changed according to sample conditions.

Here, we present a detailed design and experimental protocols of a novel ICP platform that preconcentrates targets within a spatiotemporally defined position.26 The expansion of the ion depletion zone is blocked by the ion enrichment zone, leaving a stationary preconcentration plug at an assigned position regardless of operating time, applied voltage, ionic strength, and pH. This detailed video protocol is intended to show the simplest method to integrate cation exchange membranes into microfluidic devices, and to demonstrate the preconcentration performance of the new ICP platform compared with the conventional one.

PROTOCOL:

1. Fabrication of Cation Exchange Membrane-Integrated Microfluidic Chips

1.1) Preparation of silicon masters

1.1.1) Design two kinds of silicon masters: one for patterning a cation exchange resin and the other for building a microchannel with Polydimethylsiloxane (PDMS).

Note: The detail geometry will be described in the Protocol 1.3.1 and 1.4.1.

1.1.2) Fabricate the silicon masters by using either conventional photolithography or deep reactive ion etching.27

1.1.3) Silanize the micropatterned silicon masters with trichlorosilane (~30 L) in a vacuum jar for 30 min.

CAUTION: Tricholorosilane is a pyrophoric liquid that is flammable and has an acute toxicity (inhalation, oral).

1.2) Preparation of PDMS molds

1.2.1) Mix a silicone elastomer base with a curing agent by the ratio of 10:1 and place the cup with this uncured PDMS (30-40 ml for replicating microstructures on a 4-inch silicon wafer) in a vacuum jar for 30 min to remove bubbles.

Note: the silicone base contains siloxane oligomers terminated with vinyl groups and a platinum-based catalyst. The curing agent contains crosslinking oligomers that have three silicon-hydride bonds.28

1.2.2) Pour the uncured PDMS on the silicon masters, remove bubbles with a blower, and cure the PDMS at 80 °C for 2 hr in a convection oven.

1.2.3) Detach the cured PDMS from the silicon masters, and properly shape the PDMS with a knife (squared shapes as shown in Figure 2a-b iv)).

1.3) Patterning of Cation Exchange Membranes

1.3.1) Cut the half of the PDMS mold perpendicularly to the two parallel microchannels and punch holes at the ends of the PDMS channels with a 2.0 mm biopsy punch.

Note: The PDMS mold for patterning cation selective membrane has two parallel microchannels (width: 100 μm; height: 50 μm; interchannel distance: 100 μm) (Figure 1a). We can imagine the original shape of the mold by mirroring the sliced mold along the cutting line. L-shaped microchannels are recommended for punching the two holes without overlapping.

1.3.2) Clean a glass slide and the PDMS mold with a tape and a blower, and put the mold onto the glass slide to create reversible attachment between them.

1.3.3) According to the microflow patterning technique,29 release ~10 μL of a cation exchange resin at the open-end of the channel that was sliced in step 1.3.1(Figure 1b). Next, place the syringe head on the punched holes and pull a plunger of the syringe (black arrows in Figure 1b). A gentle negative pressure pulls the cation exchange resin, and the resin fills the two channels.

Note: It is recommended that the height of the microchannel is greater than 15 μm because the high viscosity of the resin requires high pressure to fill the channels. On the other hand, the height is better not to exceed 100 μm because the patterned ion selective membrane becomes thicker than 1 μm; such thick membrane may create a gap between the membrane and the PDMS channel.13

1.3.4) Detach the PDMS mold without touching the patterned resin, and place the glass slide on the heater at 95 °C for 5 min to evaporate the solvent in the resin.

Note: the thickness of patterned membrane is usually less than < 1 μm.

Note: The mold is gently detached by hinging the mold to the open-end side (dotted line and arrow in Figure1b).

Note: We recommend the detachment time within one minute after filling the resin. If the mold is detached few minutes later, we obtain the thicker membranes, but it would has a concaved shape because of the capillary effect.

1.3.5) Peel off the unnecessary part of the patterned membrane with a razor blade, making the two separated line-patterns (Figure 1c).

Note: The cation exchange material used here has perfluorinated groups, meaning the pattern is not strongly bonded on the glass. Therefore, the simple blading method can remove the unnecessary part of the membrane easily.

1.4) Integration of the microchannel and the membrane-patterned substrate

1.4.1) Punch two holes at the ends of microchannels, and other two holes where the membrane patterns will be located after bonding the PDMS channel and the membrane-patterned substrate fabricated in Protocol 1.3.

Note: The PDMS microchannel has a one channel (width: 50-100 μm; height: 10 μm), but it is bended to neighbor the channel ends (Figure 1d).

1.4.2) Bond the PDMS microchannel with the membrane-patterned substrate immediately after oxygen plasma treatment for 40 s at 100 W and 50 mTorr.

Note: Place the patterned membrane perpendicularly on the middle of the microchannel.

2. ICP Preconcentration

2.1) Preparation for experiment

2.1.1) Prepare various test solutions including 1-100 mM KCl, 1 mM NaCl (pH ~ 7), the mixture of 1 mM NaCl and 0.2 mM HCl (pH ~ 3.7), the mixture of 1 mM NaCl and 0.2 mM NaOH (pH ~ 10.3), and 1x phosphate buffer saline.

2.1.2) Add a negatively charged fluorescent dye (~1.55 μM) into the test solutions.

Note: The concentration of the added dye should be much lower than that of the salt ions (< 10 μM), so that the charged dyes do not contribute to an electrical current.30,31

2.1.3) Load the sample solution in one reservoir of the channel and apply a negative pressure on the other reservoir to fill the solution in the channel. Next, connect the two reservoirs hydrodynamically by releasing a large droplet to eliminate the pressure gradient along the channel (Figure 2a).

2.1.4) Fill the two reservoirs, which are connected to the cation exchange patterns, with buffer solutions (1 M KCl or 1 M NaCl) by a syringe or a pipet to compensate the ICP effect in these reservoirs.

2.1.5) Place the wires at these reservoirs across the two patterned membranes (anode on the left reservoir and cathode on the right one), and connect them with a source measurement unit (Figure 2a).

2.2) Visualization of ICP phenomenon and ICP preconcentration.

2.2.1) Load the ICP device on an inverted epifluorescence microscope. Apply a voltage (0.5-100 V) and measure a current response with a source-measurement unit.

2.2.2) Capture fluorescent images with a charge-coupled device camera, and analyze fluorescent intensity using imaging software.32

REPRESENTATIVE RESULTS:

Schematic fabrication steps of a membrane-integrated microfluidic preconcentrator are shown in Figure 1. The detailed description of the fabrication is given in the Protocol. The designs and device images of the spatiotemporally defined preconcentrator26 are contrasted with those of conventional preconcentrator11 (Figure 2). ICP phenomenon in the spatiotemporally defined preconcentrator is investigated in terms of current-voltage-time responses and fluorescent intensity profiles (Figure 3-4). Similar to the ICP phenomenon with a single-membrane preconcentrator,3,11 three different regimes (Ohmic, limiting, and overlimiting) were observed in the current-voltage curve: 0.5-1 V (Ohmic and limiting) and 5 V (overlimiting). However, a nonconventional current recovery is detected in the current-time curve as the ion enrichment and the ion depletion zones are merged. Next, the ICP preconcentration has been tested in different times and voltages with the spatiotemporally defined preconcentrator (Figure 5) and the conventional one-membrane device (Figure 6). The preconcentration dynamics are quantified by fluorescence images, current-time responses, and fluorescent intensity graphs over different distances and times. Comparing the two platforms, the new ICP platform shows an advantage in collecting targets (fluorescent dyes) always between the two cation selective membrane patterns. In addition, we confirm the preconcentration plug remains the same in different ionic strength (1-100 mM NaCl), pH values (3.7-10.3), verifying high availability of the merging ICP preconcentrator in wide ranges of operating conditions (Figure 7). In Figure 8, 10,000-fold protein preconcentration is also demonstrated.

FIGURE LEGENDS:

Figure 1. Fabrication steps of a cation exchange membrane-integrated microfluidic chip. After a cation exchange resin is filled into a PDMS mold by the microflow patterning technique (a-c),29 the membrane-patterned glass substrate is bonded with a PDMS microchannel by oxygen plasma treatment (d).

Figure 2. Schematics of the spatiotemporally defined preconcentrator (a) and conventional preconcentrator (b). (a) In the new platform, between two membrane patterns (i), ion depletion/enrichment zones are developed and merged together with linear (Ohmic and limiting regime; ii) or nonlinear (overlimiting regime; iii) concentration profiles. In all three current regimes, the ion enrichment zone blocks the propagation of the depletion zone, and targets (hollow circles; i) are preconcentrated at the interface of ion depletion and enrichment zones (curved dotted line; i). The wall of the PDMS channel is negatively charged, and this generates electro-osmotic flow (EOF) between two cation exchange membranes under an electric field. The EOF delivers targets toward the interface of the depletion and enrichment zones continuously. (b) In the conventional platform, only the ion depletion zone is developed near the membrane with linear (Ohmic and limiting regime; ii) and nonlinear (overlimiting regime; iii) concentration gradients. As EOF delivers targets, the preconcentation also occurs at the depletion zone boundary, but this zone (and the preconcentrated plug) is moving away from the cation exchange membrane (black arrow; i). It is noted that there is no increase of the ion concentration here without the ion enrichment zone (ii-iii). In (a-b), the device images are shown in (iv). C0 represents the initial ion concentration. V+ and G indicate the anode and the cathode, respectively. Reprinted from Ref.26 with permission from The American Chemical Society.

Figure 3. Merged ICP phenomenon between two cation exchange membranes. (a) Current−voltage curve shows three distinct regimes (Ohmic, limiting, and overlimiting). The current response is measured by ramping up the voltage at discrete intervals of 0.25 V every 40 s, which is repeated three times. The error bar indicates the standard deviation of the current responses. (b, c) In the three regimes, fluorescence images (b) and intensity profiles along A−A′ at the middle of the channel (c) were obtained. Yellow dotted boxes indicate the location of the cation selective membranes. 1 mM KCl solution with a 1.55 μM (1 μg/mL) negatively charged fluorescent dye was used. Reprinted from Ref.26 with permission from The American Chemical Society.

Figure 4. Transient dynamics of merging ion depletion and enrichment zones. (a, b) In the Ohmic-limiting regimes, the linear concentration gradients grow (<1 s) from the both cation exchange membrane, and then overlap together (>1 s). (c) In the overlimiting regime, two ICP zones are merged more quickly (< 0.6 s) with the depletion shock (black arrow at 0.2 s). (d-f) The current-time responses show that the current is initially dropped due to the growth of the low concentrated depletion zone which corresponds to low electrical conductivity. The current drop is then recovered due to a convective transport by vortices confined between two membranes. Reprinted from Ref.26 with permission from The American Chemical Society.

Figure 5. Spatiotemporally fixed preconcentration at 5, 10, and 20 V. (a-c) Fluorescence images of the merged ICP and the current-time responses (d-f) over time (0-100 sec). Yellow dotted lines indicate the location of the cation exchange membranes. (g) Time-lapse fluorescent intensity profiles are plotted along the microchannel (A−A′). The peak intensities increase as time passes with fixed locations. (h) Peak intensity fold (i.e. how many times over the initial fluorescent intensity). At higher voltage, the faster EOF delivers targets toward the interface of the ion depletion and enrichment zones, so the preconcentration speed increases. A spike at 20 V is induced by the depletion shock (Figure 4c). The fluorescent dyes are accumulated on the shock boundary as it pushes the dyes. This initial accumulation is then somewhat dispersed when the depletion shock met the ion enrichment zone, creating the spike in the peak intensity curve. As can be seen in Figure 4c, the width of the peak at 0.8 sec was wider than that at 0.4 sec. This is probably because the left side of the left Nafion pattern (Figure 2a) was electrically floated, and the accumulated dyes could spread out. Reprinted from Ref.26 with permission from The American Chemical Society.

Figure 6. ICP phenomenon in the conventional ICP preconcentrator at 5, 10, and 20 V. (a-c) Fluorescence images of the ion depletion zone and the current-time response (d-f) over time (0-100 sec). The propagation of the depletion zone and the preconcentration plug is clearly visualized in the fluorescence images. Accordingly, the vortices are not confined, so the current recovery does not occur even in the overlimiting regime. Yellow dotted lines are marked for the location of the cation exchange membranes (g) Time-lapse fluorescent intensity profiles are plotted along the microchannel (A−A′). The peak intensities increase as time passes, but the location is moving away from the membrane. (h) Peak intensity fold of the conventional ICP device. In contrast to the merged ICP device (Figure 5h), there is no intensity spike without the confinement of ICP zones because the fluorescent intensity was keep increasing as the dyes was preconcentrated. The increase of the peak intensity fold is similar to that of the merged ICP device in the same time (at a given voltage). This indicates that how long we hold the preconcentrated plug in place is crucial for the preconcentration performance.

Figure 7. Spatiotemporally defined preconcentration at various ionic strengths (1−100 mM NaCl) and pH values (3.7−10.3). (a) Fluorescence images obtained after 100 sec of operation at 50 V. As can be seen, the locations of the preconcentration plugs are still between the two cation exchange membranes (yellow dotted lines), even though the intensity is weakened under high ionic strength and in strong acidic or basic solution. (b, c) Location of peak intensity and its intensity fold (i.e., how many times over the initial intensity) mapped under 10, 20, 50, and 100 V. For a single condition (1, 10, 100 mM and/or pH 3.7, 7, and 10), there are four data points corresponding to the four voltage conditions. At higher voltage, we have higher peak intensify fold in all cases. 100 V is not tested in 1 mM NaCl (pH 7) because the peak intensity already touches the highest values (due to the saturation of the camera) at 50 V. From the peak intensity profile, the peak region is also identified, with 1% below the peak intensity, which is represented by error bars (b, c). A higher voltage and a stronger EOF shift the peak location to the right, with a higher intensity fold and a sharper preconcentration plug. Gray boxes indicate the location of the cation exchange membranes. The 0 distance (a) represents the origin of the x-axis (b,c), which is on the right edge of the left cation exchange membrane. The origin of the distance is the right edge of the left membrane. Reprinted from Ref.26 with permission from The American Chemical Society.

Figure 8. Demonstration of spatiotemporally fixed protein preconcentration. FITC−albumin (1 μg/mL) in 1× phosphate buffer saline solution is used. 0.1% Tween 20 is also added to prevent nonspecific binding. Since the preconcentration is hardly achieved in higher ionic strength (Figure 7), we double the width of the Nafion pattern (200 μm) and use a narrower PDMS channel (50 μm). In this way, the performance of ICP preconcentration can be enhanced by broadening the ions pathway but by reducing the absolute amount of ions in the channel. At an applied voltage of 100 V, the peak and averaged fluorescence intensities are traced in the white dotted box, which is the region between the two cation exchange membranes. Within 10 min of operation, the proteins are preconcentrated up to 10 mg/mL (peak) and ∼0.1 mg/mL (averaged), indicating 10,000- and 100-fold preconcentration, respectively. The inset fluorescence images were obtained at 0, 10, 20 min. In this work, 20 min operation is enough to preconcentrate target molecules, so we did not cover longer operating times. Reprinted from Ref.26 with permission from The American Chemical Society.

DISCUSSION:

We have described the fabrication protocol and the performance of the spatiotemporally defined preconcentrator in the ranges of the applied voltage (0.5-100 V), ionic strength (1-100 mM), and pH (3.7-10.3) achieving 10,000-fold preconcentration of dyes and protein within 10 min. As like previous ICP devices, the preconcentration performance becomes better at higher voltage and at lower ionic strength. One additional parameter we can consider here is the distance between two cation exchange membranes. If we increase the inter-membrane distance, the electric field decreases under the same applied voltage, resulting the decrease of the preconcentration speed.26

The microflow patterning technique29 used in this work is a robust method for patterning cation exchange resins, so it has been one of the gold standard to integrate the ion exchange materials into the microfluidic systems. Nevertheless, it is required to pay special attention to fabricate two juxtaposed cation exchange membranes with a short intermembrane distance (< few hundred micrometers). In step 1.3.3-1.3.4, the cation exchange resin is in a liquid phase. Therefore, the resin in the two microchannels can be collapsed, and the remaining resin drop at the open-end of the channels also can flood during the mold detachment (step 1.3.4.). To build two cation exchange membranes with high pattern fidelity, we used the resin with a relatively high viscosity (20% of the cation exchange material in the solvents) and carefully set the detachment process with a designated detaching direction.

Even though the high operating flexibility of this platform was demonstrated, the reader might be concerned about which conditions are optimal among the wide range of its operating window. One of the representative trade-off is between the preconcentration speed and the stability of the ICP effect. As can be seen in Figure 5 in Kwak et al.,26 a high-applied voltage (> 50 V) can condense targets quickly; however, this also induces strong vortices in the depletion zone (1 mM / pH 7 in Figure 7a), which decrease the stability of the sample preconcentration. Accordingly, the preconcentration speed becomes not easy to be predicted.33 In the current stage, we recommend the experimental conditions under relatively low voltage (< 30 V) and ionic strength (< 10 mM) for a stable, predictable, and spatiotemporally fixed preconcentration. This trade-off between the preconcentration speed and the stability of the preconcentrated plug is also related to the sources of the nonlinear ICP (surface conduction, EOF, and electro-osmotic instability). The main source of the nonlinear ICP at a relatively small voltage (< 50 V) is EOF, creating a coherent vortex pair in a depletion zone (Figure 3b), which leads to a stable preconcentration. At a relatively high voltage (> 50 V), the main source of the nonlinear ICP is changed to electro-osmotic instability, resulting chaotic multiple vortexes, which decrease the stability of preconcentration.

Recently, paper-based ICP platforms have been developed by Phan et al.,34 Gong et al.,19 and Han et al.21. These paper devices with microporous structures can suppress electro-osmotic instability4,35, and alleviate the stability issue. However, the sizes of the paper channels are generally about 0.5~5 mm, which is much bigger than the conventional microfluidic channels. This wider paper channel with random fiber networks makes irregular motions of the preconcentrated plugs. This has been inevitable in paper-based ICP preconcentrators, because the minimum feature size of wax patterning and paper cutting (i.e. fabrication methods to build paper channels) is about few hundred micrometers.

ICP preconcentrator has been used in the wide range of biomicrofluidic platforms for preconcentrating various bio-agents, amplifying the signals of various assays, and finally detecting the targets such as therapeutic proteins,36 peptides,37 aptamers,17 and enzymes.38 As one can notice, these previous works targeted fluorescence-labeled biomolecules. This is because that we cannot specify the exact operating condition (e.g. voltage and flow rate) to maintain the preconcentration site, so we need to find the proper condition to preconcentrator targets first. Departing from the previous works, the merged ICP phenomenon allows us to fix the preconcentrated plugs always at broad range of operating conditions, while maintaining the ICP devices’ high flexibility. For example, we can modulate the merged ICP system with a tangential fluid flow, and operate it in the continuous-flow mode39. This indicates that now we can extend the applications of ICP preconcentrators for label-free detection techniques without visualization instruments and tracers. This unique advantage on the spatiotemporal controllability would provide a strong commercial opportunity to integrate the ICP device into generic benchtop platforms, such as for polymerase chain reactions and mass spectrometers.

ACKNOWLEDGEMENTS:

This work was supported by the internal fund of the Korea Institute of Science and Technology (2E26180), and the Next Generation Biomedical Device Platform program funded by the National Research Foundation of Korea (NRF-2015M3A9E202888).

DISCLOSURE:

The authors have nothing to disclose.

REFERENCES:

- Probstein, R. F. Physicochemical Hydrodynamics: An Introduction, 2nd ed. Wiley-Interscience. New York (2003).

- Strathmann, H. Ion-Exchange Membrane Separation Processes. Elsevier. Amsterdam (2004).

- Dydek, E.V., et al. Overlimiting Current in a Microchannel. Phys. Rev. Lett. 107 (11), 118301, doi:10.1103/PhysRevLett.107.118301 (2011).

- Kwak, R., Pham, V.S., Lim, K.M., & Han, J. Shear flow of an electrically charged fluid by ion concentration polarization: scaling laws for electroconvective vortices. Phys. Rev. Lett. 110 (11), 114501, doi:10.1103/PhysRevLett.110.114501 (2013).

- Rubinstein, I., & B. Zaltzman, B. Electro-osmotically induced convection at a permselective membrane. Phys. Rev. E 62 (2), 2238-2251, doi:10.1103/PhysRevE.62.2238 (2000).

- Kwak, R., Kim, S., & Han, J. Continuous-flow biomolecule and cell concentrator by ion concentration polarization. Anal. Chem. 83 (19), 7348-7355, doi:10.1021/ac2012619 (2011).

- Jeon, H., Lee, H., Kang, K.H., & Lim, G. Ion concentration polarization-based continuous separation device using electrical repulsion in the depletion region. Sci. Rep. 3, 3483, doi:10.1038/srep03483 (2013).

- Kim, S.J., Ko, S.H., Kang K.H., & Han, J. Direct seawater desalination by ion concentration polarization. Nat. Nanotechnol. 5 (4), 297-301, doi:10.1038/nnano.2010.34 (2010).

- MacDonald, B.D., Gong, M.M., Zhang, P., & Sinton, D. Out-of-plane ion concentration polarizationfor scalable water desalination. Lab Chip 14 (4), 681-685, doi:10.1039/C3LC51255J (2014).

- Schoch, R.B., Han, J., & Renaud, P. Transport phenomena in nanofluidics. Rev. Mod. Phys. 80 (3), 839-883, doi:10.1103/RevModPhys.80.839 (2008).

- Kim, S.J., Song, Y.A., & Han, J. Nanofluidic concentration devices for biomolecules utilizing ion concentration polarization: theory, fabrication, and applications. Chem. Soc. Rev. 39 (3), 912-922, doi:10.1039/b822556g (2010).

- Mai, J.Y., Miller, H., & Hatch, A.V. Spatiotemporal mapping of concentration polarization Induced pH changes at nanoconstrictions. ACS Nano 6 (11), 10206-10215, doi:10.1021/nn304005p (2012).

- Kim, B., et al. Tunable ionic transport for a triangular nanochannel in a polymeric nanofludic system. ACS Nano 7 (1), 740-747, doi:10.1021/nn3050424 (2013).

- Syed, A. Mangano, L. Mao, P.Han J., & Song, Y.-A. Creating sub-50 nm nanofluidic junctions in a PDMS microchip via self-assembly process of colloidal silica beads for electrokinetic concentration of biomolecules. Lab Chip 14, 4455-4460, doi:10.1039/C4LC00895B (2014).

- Wang Y.C., Stevens, A.L., & Han, J., Million-fold preconcentration of proteins and peptides by nanofluidic filter. Anal. Chem. 77 (14), 4293-4299, doi:10.1021/ac050321z (2005).

- Lee, J.H., Cosgrove, B.D., Lauffenburger, D.A., & Han, J. Microfludic concentration-enhanced cellular kinase activity assay. J. Am. Chem. Soc. 131 (30), 10340-10341, doi:10.1021/ja902594f (2009).

- Cheow, L.F. & Han, J. Continuous signal enhancement for sensitive aptamer affinity probe electrophoresis assay using electrokinetic concentration. Anal. Chem. 83 (18), 7086-7093, doi:10.1021/ac201307d (2011).

- Ko, S.H., et al. Nanofluidic preconcentration device in a straight microchannel using ion concentration polarization. Lab Chip 12 (21), 4472-4482, doi:10.1039/C2LC21238B (2012).

- Gong, M.M., Nosrati, R., Gabriel, M.C.S., Zini, M., & Sinton, D. Direct DNA analysis with paper-based ion concentration polarization. J. Am. Chem. Soc. 137 (43), 13913–13919, doi:10.1021/jacs.5b08523 (2015).

- Hong, S., Kwak, R., & Kim, W. Paper-based flow fractionation system applicable to preconcentration and field-flow separation. Anal. Chem. 88 (3), 1682–1687, doi:10.1021/acs.analchem.5b03682 (2016).

- Han, S.I., Hwang, K.S., Kwak, R., & Lee, J.H. Microfluidic paper-based biomolecule preconcentrator based on ion concentration polarization. Lab Chip 16, 2219-2227, doi:10.1039/C6LC00499G (2016).

- Zangle, T.A., Mani, A., & Santiago, J.G. Theory and experiments of concentration polarization and ion focusing at microchannel and nanochannel interfaces. Chem. Soc. Rev. 39 (3), 1014-1035, doi:10.1039/B902074H (2010).

- Mani, A., & Bazant, M.Z. Deionization shocks in microstructures. Phys. Rev. E 84, 061504, doi:10.1103/PhysRevE.84.061504 (2011).

- Slouka, Z., Senapati, S., & Chang, H.C. Microfluidic systems with ion-selective membranes. Annu. Rev. Anal. Chem. 7, 317-335, doi:10.1146/annurev-anchem-071213-020155 (2014).

- Kirby, B.J., & Hasselbrink, E.F. Zeta potential of microfluidic substrates: 1. Theory, experimental techniques, and effects on separations. Electrophoresis 25 (2), 187-202, doi:10.1002/elps.200305754 (2004).

- Kwak, R., Kang, J.Y., & Kim, T.S. Spatiotemporally defining biomolecule preconcentration by merging ion concentration polarization. Anal. Chem. 88 (1), 988−996, doi:10.1021/acs.analchem.5b03855 (2016).

- Duffy, D.C., McDonald, J.C., Schueller, O.J.A., & Whitesides, G.M. Rapid prototyping of microfluidic systems in poly(dimethylsiloxane). Anal. Chem. 70 (23), 4974-4984, doi:10.1021/ac980656z (1998).

- Campbell, D.J., et al. Replication and compression of surface structures with polydimethylsiloxane elastomer. J. Chem. Educ. 76(4), 537-541, doi:10.1021/ed076p537 (1999).

- Lee, J.H., Song, Y.A., & Han, J. Multiplexed proteomic sample preconcentration device using surface-patterned ion-selective membrane. Lab Chip 8 (4), 596-601, doi:10.1039/b717900f (2008).

- Kwak, R., Guan, G., Peng, W.K., & Han, J. Microscale electrodialysis: concentration profiling and vortex visualization. Desalination 308, 138–146, doi:10.1016/j.desal.2012.07.017 (2013).

- Chambers, R.D. & Santiago, J.G. Imaging and quantification of isotachophoresis zones using nonfocusing fluorescent tracers. Anal. Chem. 81, 3022–3028, doi:10.1021/ac802698a (2009).

- Rasband, W.S., ImageJ, U. S. National Institutes of Health, Bethesda, Maryland, USA, http://imagej.nih.gov/ij/, 1997-2016 (2016).

- Minerick, A.R., Ostafin, A.E., & Chang, H.C. Electrokinetic transport of red blood cells in microcapillaries. Electrophoresis 23 (14), 2165-2173, doi:10.1002/1522-2683(200207)23:14<2165::AID-ELPS2165>3.0.CO;2-# (2002).

- Phan, D-.T., Shaegh, S.A.M, Yang, C., & Nguyen, N-.T. Sample concentration in a microfluidic paper-based analytical device using ion concentration polarization. Sens. Actuators B 222, 735-740 doi: 10.1016/j.snb.2015.08.127 (2016).

- Rubinstein, S.M., Direct observation of a nonequilibrium electro-osmotic instability. Phys. Rev. Lett. 101, 236101 doi: 10.1103/PhysRevLett.101.236101 (2008).

- Ouyang, W., et al. Microfluidic platform for assessment of therapeutic proteins using molecular charge modulation enhanced electrokinetic concentration assays. Anal. Chem. 88, 9669-9677, doi:10.1021/acs.analchem.6b02517 (2016).

- Cheow, L.F., Sarkar, A., Kolitz, S., Lauffenburger, D., & Han, J. Detecting kinase activities from single cell lysate using concentration-enhanced mobility shift assay. Anal. Chem. 86, 7455-7462, doi:10.1021/ac502185v (2014).

- Chen, C-.H., et al., Enhancing protease activity assay in droplet-based microfluidics using a biomolecule concentrator. J. Am. Chem. Soc. 133, 10368-10371, doi:10.1021/ja2036628 (2011).

- Kwak, R., Pham, V.S., Kim, B., Lan C., & Han, J. Enhanced salt removal by unipolar ion conduction in ion concentration polarization desalination. Sci. Rep. 6, 25349, doi:10.1038/srep25349 (2016).