# Submission ID #: 5634

# Scriptwriter Name: Amy Manocchi

# Videographer name: ?

# Filming Date: May 8, 2015

# JoVE Science Education Series: Analytical Chemistry

# Title: Introduction to Mass Spectrometry

- Overview

  - Mass spectrometry is an analytical technique that enables the identification and quantification of unknown compounds within a sample, and the determination of their structure.

    - Title Slide.

  - In mass spectrometry, gas phase ions are generated from the atoms or molecules in a sample. (1.2.1.) The ions are then separated based on their mass-to-charge ratio, symbolized by m-over-z. (1.2.2.)

    - See storyboard

  - This separation enables the determination of quantitative and qualitative information about a sample, including the quantity and mass of each compound, as well as their structure.

    - See storyboard.

  - This video will introduce the basic concepts and instrumentation of mass spectrometry, and demonstrate element quantification in the laboratory using a mass spectrometer.

    - Use shot 4.1.1.

- Principles of Mass Spectrometry

  - A mass spectrometer is composed of an ionization source, a mass analyzer, and a detector. (2.1.1.) At the ionization source, the compounds are ionized, usually to a single positive charge.   (2.1.2.)

    - See storyboard.

  - Ions can be generated using various techniques, each resulting in a range of fragmentations.  These methods are loosely grouped into “hard” and “soft” ionization. (2.2.1.) Hard ionization techniques cause more extensive fragmentation, resulting in more fragments of lower mass.  Hard ionization, such as Inductively Coupled Plasma, or ICP, is typically used for gases or volatile organic molecules. (2.2.2.)

    - See storyboard.

  - Soft ionization techniques result in less, or almost no, fragmentation with a high molecular mass range.  (2.3.1.) Soft ionization, such as Matrix Assisted Laser Desorption Ionization, or MALDI (pronounced mall-dee), is typically used in the analysis of large molecules such as proteins. (2.3.2.)

    - See storyboard.

  - The fragmentation of a compound into a number of smaller ions can aid in the determination of a molecules structure.  However, if the fragmentation is too great, valuable structure information can be lost.  Thus, the selection of an ionization method depends on the sample of interest, the desired degree of fragmentation, and how the sample is introduced to the mass spectrometer.

    - See storyboard

  - The ions are then accelerated in an electric field at a certain voltage, where ions with small mass-to-charge ratios travel at a higher acceleration than those with large mass-to-charge ratios. (2.5.1.) After the ions have been accelerated, they enter the mass analyzer.  (2.5.2.)

    - See storyboard.

  - The most basic mass analyzer is a magnetic sector, which is composed of a curved magnet that produces a homogeneous magnetic field.  (2.6.1.) The attractive force of the magnet, plus the centrifugal force of the accelerating ions causes them to travel in a circular path.  (2.6.2.)

    - See storyboard.

  - The radius of the ions circular path depends on the accelerating voltage, the applied magnetic field, and the mass-to-charge ratio. When voltage and magnetic field strength are held constant, the radius depends only on the mass-to-charge ratio of the ion.

    - See storyboard.

  - The voltage and magnetic field can then be selected to only allow certain mass-to-charge ratio species through the curved path.  Other ions crash into the sides of the magnetic pathway and are lost. (2.8.1.) By scanning the magnetic field strength, desired ions reach the detector at different times, thereby identifying each species precisely. (2.8.2.)

    - See storyboard.

  - The most common type of mass analyzer is the quadrupole mass filter. The quadrupole consists of two pairs of parallel metal rods, with each pair of opposing rods electrically connected.

    - See storyboard.

  - An oscillating radio frequency voltage and a constant direct current voltage is applied to the rod pairs.  The potentials of the rod pairs are always out of phase with the other.

    - See storyboard.

  - The ion beam is then directed through the center of the four rods.  Ions travel in a corkscrew-like path, due to the constant attraction and repulsion from the rods. (2.11.1.) Depending on the ions mass-to-charge ratio, the ion will either travel the full path of the quadrupole and reach the detector, or will crash into the rods.  (2.11.2.)

    - See storyboard.

  - This video will demonstrate the use of a mass spectrometer in the laboratory.  The mass spectrometer used in this experiment is an ICP ionizer, with a quadrupole analyzer.  (2.12.1.) ICP is a method of hard ionization that uses coupled argon plasma at about 10,000 °C that is electrically conductive to ionize the sample molecules. (2.12.2.). ICP-MS will be demonstrated to detect and quantify a metal component in a sample, in this case iron.  (2.12.3.)

    - -2.  See storyboard

2.12.3.  Use shot: 5.5.1.

- Sample Preparation and Digestion

  - To begin the experiment, fill all polypropylene tubes with 5 ml of 0.1 M hydrochloric acid in order to remove any contaminating trace of iron.

    - WIDE: Talent brings materials to the bench.

    - MED: Talent fills polycarbonate tubes with HCl.

  - Place the tubes in a water bath for one hour at 50 °C.

    - MED: Talent puts tubes in water bath.

  - After incubation, wash the tubes with 5 ml of deionized water, and dry the tubes in an oven (TEXT: 3 h at 80 °C) or chemical hood.

    - MED (over the shoulder): Talent pours HCl into waste bottle.

    - MED: Talent rinses tubes with deionized water.

    - WIDE: Talent places them in the oven to dry.

  - In the clean tubes, add 1.8 ml of concentrated nitric acid (TEXT: 65% v/v Nitric acid) and 200 l of sample containing the isotope of interest. Follow safety precautions when using concentrated acid.

    - MED: Talent adds nitric acid to tube.

    - CU: Sample to be measured.

    - MED: Talent adds sample to tube.

  - Place the tubes in a water bath overnight (TEXT: 50 °C). The temperature can be increased to shorten digestion time, if necessary.

    - MED: Talent caps tubes

    - WIDE: Talent puts tubes in water bath.

  - After the sample has been digested, let the tubes cool to room temperature.

    - MED: Talent removes tubes from water bath.

  - Next, add 8 ml of deionized water to dilute the samples, and to obtain a nitric acid concentration below 20%.  The final dilution of the sample is 1/50. The ideal concentration for ICP/MS is in the parts-per-billion range. Centrifuge the tubes (TEXT: 3,000 x g, 10 min) to pellet any remaining macroscopic residues.

    - MED: Talent adds water to tubes, and caps them.

    - WIDE: Talent puts tubes in centrifuge.

- Preparation of the Instrument

  - Begin the instrument set up by inspecting the torch. The torch is made of three concentric quartz tubes, which directs the plasma and injects the sample into the heart of the plasma. If the torch is dirty, place it in 5% nitric acid, and ultra-sonicate for 15 minutes to clean.

    - MED: Talent shows the Torch

    - Talent checks that the torch is clean

  - Next, inspect the sampler and skimmer cones, and if dirty, remove them.  These cones enable the sampling of only the inner portion of the ion beam generated by the ICP torch and act as a barrier to the high vacuum of the mass spectrometer. If necessary, wipe the sampler and skimmer cones with 5% nitric acid.

    - MED: Talent shows the cones

    - CU: Talent checks that cones are clean

    - Talent close the door of the ICP-MS machine

  - Next, check and, if necessary, change the peristaltic tubing and check the pump oil level.

    - MED: Talent check tubing

    - WIDE: Talent checks the pump oil level

  - Check the argon pressure and start the chiller.   Start the plasma and liquid flow into system. Wait 20 minutes for system to warm up fully. (TEXT: 20 min).

    - WIDE: Talent check that argon stream is on

    - a: MED: Talent starts the chiller

b: Talent starts the plasma and camera shows the plasma

  - Next, aspirate the test solution containing a range of elemental masses.  For example, magnesium, indium and uranium are three convenient test elements that cover wide mass range.

    - MED: Talent aspirates the test solution

  - When the solution flow is established, check and adjust the torch alignment with cones, especially if torch and cones have been disturbed. Optimize the torch alignment using the instrument software.

    - MED (Over the shoulder): Talent checks to see that flow is established.

    - MED: Talent adjusts torch alignment with cones.

    - See storyboard: Author Image 5634 image 1.png

  - Then, adjust the nebulizer gas flow, which optimizes the flow of argon through the nebulizer, to give the highest sensitivity without excessive oxide formation.

    - MED (Over the shoulder): Talent adjusts the nebulizer gas flow.

    - See storyboard: Author Image 5634 image 2.png

  - Next, optimize the electrostatic lens voltages. This lens prevents neutral particles and photons from entering, and collimates the ion beam.

    - MED (Over the shoulder): Talent optimizes the electrostatic lens voltages.

    - See storyboard: Author image 5634 image 3.png

  - Using the instrument software, run a standard mode performance check by measuring the same test solutions. Oxide formation, double charge formation, and background level are also checked.

    - MED (Over the shoulder): Talent runs a performance check.

    - See storyboard: Author image 5634 image 4 and 5.png

  - Next, run a dynamic reaction cell or DRC check. The dynamic reaction cell is a chamber located before the mass filter that is filled with ammonia to avoid polyatomic interferences.  To do so, set the flow rate of ammonia to 0.6 ml/min. Run the DRC check using the instrument software.

    - MED (Over the shoulder): Talent sets flow rate and runs DRC.

    - See storyboard: Author image 5634 image 6 and 7.png

- Instrument Operation

  - To run the instrument, first select the elements and isotopes of interest.  Then set the scan mode to peak hopping.

    - MED (Over the shoulder): Talent inputs elements, and isotopes of interest.  Sets scan mode to peak hopping

    - See storyboard: Author Image 5634 image 8.png

  - Select five replicates per measurement. Set each replicate to contain 40 measurement sweeps, each sweep with a dwell time of 50 milliseconds.  The total integration time is 2,000 milliseconds per replicate. If the sample is limited (TEXT: 0.5-1 ml), consider reducing the dwell time, number of sweeps, replicates and/or dilute the samples.

    - MED (over the shoulder): Talent inputs parameters into software.

    - See Storyboard: Author Image 5634 image 8.png

  - Prepare a calibration curve for the elements of choice using pre-prepared standard solutions.  In this case, three standards were prepared containing different concentrations of iron-56. Place all calibrants, blanks, and samples on the auto sampler.

    - CU: Standard solutions.

    - MED: Talent loads samples into auto sampler.

  - Ensure that intensity increases with increasing concentration during calibration measurements. (Comments: called 5.5 in video slating)

    - MED (over the shoulder): Talent shows calibration curve.

    - See storyboard: Author image 5634 image 9.png

    - See storyboard: Author image 5634 image 11.png

  - Next, run the sample. (Comments: called 5.4 in video slating)

    - Camera shows the samples running

    - See storyboard: Author image 5634 image 10.png

- Applications

  - Mass spectrometry is used in a wide range of applications using various ionization and mass analysis techniques.

    - Title Slide.

  - (Lower third: Application #1- Analysis of Proteins using MALDI-TOF) In this example, Matrix Assisted Laser Desorption Ionization coupled with a Time-of-Flight analyzer, called MALDI-TOF, (TEXT: MALDI-TOF) was used to analyze high molecular weight proteins. With MALDI, molecules are co-crystalized with a matrix, to improve stability by absorbing energy during laser desorption.

    - 50635 @ 00:14-00:32: Animation of sample deposition/analysis

    - 50635 @ 00:33-00:37: If needed to cover VO.

  - The matrix solution was prepared, which co-crystallized the proteins. (6.3.1.) The protein solution and matrix were both spotted on the clean MALDI plate, (6.3.2.) and dried. (6.3.3.)

    - 50635 @ 03:24-03:32: Preparation of matrix solutions.

    - 50635 @ 04:20-04:35: Ddeposition of samples on plate.

    - 50635 @ 04:44: Close up shot of dried protein/matrix spot.

  - The MALDI plate was inserted into the instrument, (6.4.1.) and the sample analyzed.  (6.4.2.)

    - 50635 @ 04:50-04:52: Insertion of MALDI plate in instrument.

    - 50635 @ 05:09-05:13: Obtaining spectrum of protein samples.

  - (Lower third: Application #2- Analysis of Volatiles) The analysis of volatile and oxidation sensitive compounds was measured using electron impact mass spectrometry, a hard ionization technique, with a custom inlet port, as shown in this example.

    - 51858 @ 00:40-00:58: Schematic of custom gas inlet port.

  - First, a lockable tube system was designed in order to enable full evacuation of the tube, (6.6.1.) followed by loading of the sample under cooling by liquid nitrogen. (6.6.2.)

    - 51858 @ 01:48-01:53: Evacuate lockable test tube.

    - 51858 @ 02:02-02:06: Immerse evacuated tube in liquid nitrogen.

  - The sample tube connected to the inlet port, (6.7.1.) and the sample loaded into the instrument. (6.7.2.) The mass spectrum of the sample in this case tris(trifluoromethyl) phosphate, was then analyzed. (6.7.3.)

    - 51858 @ 02:53-02:58: Connect sample tube to the mass spec inlet.

    - 51858 @ 03:24-03:28: Open tube to let sample flow to MS.

    - 51858 @ Figure 3: Spectrum of tris(trifluoromethyl)phosphate.

  - (Lower third: Application #3- Molecular Beam Mass Spectrometry)  A molecular beam coupled to an ultraviolet photoionization mass spectrometer was used to explore the electronic structure of gas phase molecules and clusters. (6.8.1.) The molecular beam integrated with synchrotron radiation provided a selective ionization method to probe molecules in the gas phase. (6.8.2.)

    - 50164 @ Figure 1

    - 50164 @ 03:44-03:50: BROLL of instrument.

  - The sample was loaded into the nozzle, the nozzle reloaded into the instrument, (6.9.1.), and the photon beam allowed to enter the chamber.  (6.9.2.)

    - 50164 @ 02:06-02:16:  Loading powder into sample holder

    - 50164 @ 03:31-03:35:  Open the chamber to allow the beam in.

  - The mass spectrum was then collected and compared to photoionization efficiency data (6.10.1.) in order to determine the electronic structure of molecules.  (6.10.2.)

    - 50164 @ Figure 7

    - 50164 @ 00:47: Molecular electronic structure data.

- Summary

  - You’ve just watched JoVE’s introduction to mass spectrometry.  You should now understand the basic instrumentation of mass spectrometry, and how to run a basic mass spectrometry based analysis.  Thanks for watching!