Photoconversion movie (time written in min:sec)

<cut 0:00 to 3.27>

Step 3.2 of script

3:26 Software finishes loading and opens

3.48 Select sequential and load pre-saved settings

4:01 Go to configuration and set laser power

4:08 Return to Acquire tab. Click live and adjust display settings.

4:30 Select bidirectional setting. Change zoom factor to 1.5

4:35 Click live and go through z stack to see signal in red channel

4:57 Select Seq 2 and check signal in green channel

Step 3.3 of script

5:30 Select FRAP mode

5:45 Load pre-saved settings and click live

5:58 Change to Airy 1 pinhole

Step 3.4 of script

6:02 select slice you want to photoconvert

6:10 Go to Bleach tab

6:14 Clear all existing ROIs

6:19 Turn on UV laser, set to 25%

6:20 Select ROI for bleaching

6:29 Go to TimeCourse tab and bleach. Bleach again if necessary (not shown).

<cut 6:45 to 7:03>

Step 3.5 of script

7:02 Return to TCS SP8 mode, select xyz acquisition mode and load pre-saved settings. Click live and visualise red channel.

7:40 Adjust 561 laser power

7:51 Select beginning of z stack

8.15 Select end of z stack

8:20 Visualize the signal in the green and brightfield channels, adjust laser power for green channel and gain for BF channel

8:58 Start image acquisition

11:04 Check acquired images

<cut 11:29 to end>