
TITLE:
Database-guided Flow-cytometry for Evaluation of Bone Marrow Myeloid Cell Maturation
AUTHORS & AFFILIATIONS:
Carmen Mariana Aanei1, Marie Christine Jacob2, Richard Veyrat-Masson3, Tiphanie Picot1, Maria Alessandra Rossenthal-Allieri4, Anne-Catherine Lhoumeau5, Michel Ticchioni4, Florent Dumezy6, Lydia Campos Catfal1
1Department of Hematology, University Hospital of Saint-Etienne, Saint-Etienne, France

2Department of Immunology, University Hospital of Grenoble, Grenoble, France

3Department Hematology, University Hospital of Clermont-Ferrand, Clermont-Ferrand, France 
4Department of Immunology, University Hospital of Nice, Nice, France

5Department of Biopathology, Institute Paoli-Calmettes, Marseille, France

6Department of Hematology, University Hospital of Lille, Lille, France 

KEYWORDS:
Multiparameter flow cytometry, MFC, myelodysplastic syndromes, MDS, myeloid dysplasia, maturation database, bone marrow myeloid cells, database-guided analysis in MFC

SHORT ABSTRACT 

The MDS diagnosis is difficult in the absence of morphological criteria or non-informative cytogenetics. MFC could help refine the MDS diagnostic process. To become useful for clinical practice, the MFC analysis must be based on parameters with sufficient specificity and sensitivity, and data should be reproducible between different operators.

LONG ABSTRACT 
A working group initiated within the French Cytometry Association (AFC) was developed in order to harmonize the application of multiparameter flow cytometry (MFC) for myeloid disease diagnosis in France. The protocol presented here was agreed-upon and applied between September 2013 and November 2015 in six French diagnostic laboratories (University Hospitals of Saint-Etienne, Grenoble, Clermont-Ferrand, Nice, and Lille and Institut Paoli-Calmettes in Marseille) and allowed the standardization of bone marrow sample preparation and data acquisition. Three maturation databases were developed for neutrophil, monocytic, and erythroid lineages with bone marrow from “healthy” donor individuals (individuals without any evidence of a hematopoietic disease). A robust method of analysis for each myeloid lineage should be applicable for routine diagnostic use. New cases can be analyzed in the same manner and compared against the usual databases. Thus, quantitative and qualitative phenotypic abnormalities can be identified and those above 2SD compared with data of normal bone marrow samples should be considered indicative of pathology. The major limitation is the higher variability between the data achieved using the monoclonal antibodies obtained with methods based on hybridoma technologies and currently used in clinical diagnosis. Setting criteria for technical validation of the data acquired may help improve the utility of MFC for MDS diagnostics. Establishment of these criteria requires analysis against a database. The reduction of investigator subjectivity in data analysis is an important advantage of this method.
INTRODUCTION:
In the absence of phenotypic markers specific to the dysplastic changes occurring in myeloid cells during MDS initiation and progression, a new approach has been proposed in recent years based on evaluation of the maturation pathways (altered expression of myeloid antigens during the production of mature myeloid cells) or of the abnormal distribution of different cell types within bone marrow (BM) cell compartments1,2.
This article presents a new method for standardized application of MFC in order to detect dysplastic changes in BM myeloid cell compartments related to myelodysplastic syndromes (MDS) or other myeloid hematological diseases. This study also shows the utility of using maturation databases for MFC data analysis.
Standardization of the sample preparation procedure, data acquisition, and analysis using the databases would allow the identification of the most relevant phenotypic abnormalities related to dysplastic changes in BM myeloid cells. Therefore, statistically selected subsets based on well-labeled and well-recognized formats (Automatic Population Separator (APS) diagrams, histograms, and dot plots) are required for developing an analysis strategy that can be used in subsequent analysis rounds. The discovery of robust phenotypic abnormalities in MDS would ease the diagnosis in cases with or without minimal morphological dysplasia and without cytogenetic aberrancies. Identification of discriminatory parameters allowing for the reduction of immunophenotypic panels may simplify the current scores2, permitting their applicability in normal laboratories.

This method limits the subjective interpretations of cytometry data, as have been signaled by MDS diagnosis3. This step is a prerequisite for the development of automated tools for processing and analyzing flow data4.
MDS comprises a heterogeneous group of clonal hematopoietic stem cell (HSC) disorders in which the spliceosome mutations cooperate with specific epigenetic modifiers to yield the MDS phenotype. It is now known that, along with HSC mutations, other mechanisms are involved in MDS pathophysiology, such as aberrant immune-mediated inflammation and interactions between malignant HSCs and the stromal microenvironment of the BM. However, these mechanisms remain poorly understood. The wide clinical and biological heterogeneity of MDS makes diagnosis and selection of the optimal therapy a challenge. In the last decade, multiple studies have shown that MFC is often more sensitive in detecting dysplasia2 than morphology, but technical and economic constraints make this technique difficult to standardize, with results often depending on the experience of the interpreter3. In addition, it is unclear how MFC can tip the balance toward MDS in cases with or without minimal morphological dysplasia and the absence of cytogenetic anomalies, or in borderline cases such as hypocellular MDS, with a low blast count, from other non-clonal BM disorders such as bone marrow failure (i.e., aplastic anemia). It also remains difficult to differentiate borderline cases of MDS with an excess of blasts from acute myeloid leukemia (AML). For all these reasons, the clinical guidelines do not integrate MFC testing into the MDS final diagnosis. In 2011, the US National Comprehensive Cancer Network (NCCN) recommended MFC for estimation of the percentage of CD34+ cells, detection of paroxysmal nocturnal hemoglobinuria clones, and presence of cytotoxic T-cell clones in hypocellular MDS5. These two latter situations also involve a therapeutic goal because clinical data have shown a good response of these patients to immunosuppressive therapy6. The 2017 NCCN guidelines, citing the International Working Group (IWG) recommendations, listed aberrant immunophenotyping detection by MFC among the co-criteria for MDS diagnosis, but without making any specifications6. In addition, the recently published WHO classification stipulates that MFC findings alone are not sufficient to establish a primary diagnosis of MDS in the absence of conclusive morphological and/or cytogenetic data7. However, MFC can be used as an additional test showing the dysregulation of myeloid cell maturation patterns and quantifying the “distance from normal” for a patient at a specific time in the disease course.
This method is applicable at clinical laboratories interested in the evaluation of dysplasia in BM myeloid cells using MFC immunophenotyping, in order to refine the diagnosis in MDS or other myeloid disorders with dysplastic abnormalities.

PROTOCOL
The protocol listed below has been approved by the "Comité de Protection des Personnes" (Independant Ethics Committee) Sud-Est 1 from University Hospital of Saint-Etienne, France. 

1.  Cytometer Settings 
Note: The cytometer settings were performed according to France Flow recommendations, in accordance with EuroFlow Procedure “EuroFlow Standard Operating Protocol (SOP) for Instrument Setup and Compensation (https://www.euroflow.org/usr/pub/protocols.php).

1.1 Monthly instrument setup 
1.1.1 Turning on the cytometer. Check that all fluid levels are appropriate and open Diva 6.1.3. Perform fluidics startup: in the menu bar, select Cytometer | Fluidics Startup. Click OK when prompted. Allow the cytometer to warm up for at least 30 min. 

1.1.2  Performance check - CST beads
Note: For this step, prepare 12 ( 75 mm polystyrene tube, CST beads and sheath fluid (see Table of Materials).
1.1.2.1  Label a 12 ( 75 mm polystyrene tube CST. Mix the provided bead vial by gentle inversion or very gentle vortexing. Add to the labeled tube: 0.35 mL of Sheath Fluid and 1 drop of CST beads. Vortex the tube gently and proceed to acquisition. Store the tube for up to 8 h at 2-25 °C in the dark if not acquiring immediately.

1.1.2.2 Perform the performance check: in the menu bar, select Cytometer | CST.
1.1.2.3  In the Setup tab of the CST module: confirm the Canto II as in the default 4-2H-2V configuration and also confirm that a baseline created using the current lot of CST beads exists for this configuration. If a baseline does not exist, refer to the CST Beads IFU. Confirm that this baseline has not expired: under Setup Control, select Check Performance from the drop-down menu. 
1.1.2.4  Check Load tube manually and click Run. Confirm the lot number displayed. Gently vortex the diluted beads prepared above and when prompted, load the diluted beads and click OK. 
1.1.2.5  When the performance check is complete, verify that Cytometer Performance passed. Click View Report. Re-run the performance check if the results did not pass. Save the Report in  PDF format: Performance Tracking Report Date. 
1.1.3  Adjust Fluorescent PMT voltages with Rainbow Beads (Table of Materials). 
Note: The target values are stipulated in the Euroflow Standard Operating Procedure (SOP) entitled “20180302_7th_Peak_Target_Values_Rainbow_Beads”. This SOP is available on Euroflow site (www.euroflow.org; in the public area; Protocols tab). 
1.1.3.1  Create a new experiment: Monthly Instrument Setup Date | Specimen. 

1.1.3.2  In Cytometer Settings, choose the optical parameters and fluorochromes corresponding to the tubes concerned (FITC, PE, PerCPCy5.5, PE-Cy7, APC, APC-H7, V450, V500), and check the desired acquisition parameters (Log, A, H and / or W). Apply current CST settings (right click on Cytometer Settings of the Experiment). Set the threshold for FSC parameter at 10,000. 
1.1.3.3  On the cytometer, turn compensation off while setting fluorescence PMT voltages for Target MFI setting; for this purpose, go in the Inspector, navigate to the Compensation tab. Disable compensation by unchecking Enable Compensation option. 
1.1.3.4  Create a worksheet Target MFI with all necessary dot plots (n = 2; FSC versus SSC, FITC versus PE), histograms (n = 8; one histogram for each fluorescence detector) and statistics showing the reference peak values (MFI and CV) for each fluorescence channel. 
1.1.3.5  Dilute 1 drop of 8-peak Rainbow beads calibration particles in 1 mL of distilled water and vortex before use. Acquire without recording the 8-peak Rainbow beads solution at LOW flow rate. Store the tube for up to 8 h at 2-25 ˚C in the dark if not acquiring immediately.
1.1.3.6  Gate singlet beads Population P1 in the FSC versus SSC bivariate dot plot and the 8th or 7th peak in the FITC versus PE bivariate dot plot (the brightest peak or the next one downwards, as is stipulate in the Euroflow document entitled “20180302_7th_Peak_Target_Values_Rainbow_Beads”) and name this gate Population P2. 
1.1.3.7  Continue the acquisition of the 8-peaks Rainbow bead suspension and adjust PMT voltages in all fluorescence channels to reach target MFI values according the Euroflow document “20180302_7th_Peak_Target_Values_Rainbow_Beads”. 
1.1.3.8  Once Target MFI values for the 8th or 7th Peak are reached, record the Target MFI achieved and the corresponding final PMT values. Acquire 5,000 events and record data. 
Note: That PMT values must be used below in step 1.2.3. Performance Check – Confirmation of PMT Values with Rainbow Beads. 
1.1.3.9  Record these values making a print screen with Worksheet Target MFI and instrument settings and save as .jpg picture. 
Note: When a “New” tube is created, sometimes the PMT values vary unexpectedly. For this reason, a double checking of PMT is essential. Compare each time the PMT values to your notes, double check that Target MFI values and PMT values are correct! 
1.1.3.10 Save the Application Settings. In the Browser, right-click on Cytometer Settings. From the drop-down menu, select Application Settings, and save. Click OK. If prompted, click Yes to maintain the threshold values. 
Note: Save the Application Settings using the default name. Do NOT rename the settings.
1.1.4  Adjust the FCS and SSC voltages with lysed washed blood (LWB). 
Note: For this step, 50 μL of a peripheral blood (PB) sample from a healthy volunteer, lysing solution (Table of Materials) and washing buffer are needed.
1.1.4.1  Pipets 50 μL of PB into a tube. Add 2 mL of freshly diluted lysing solution. Mix gently and incubate for 10 min at RT. 

1.1.4.2  Centrifuge for 5 min at 540 g.
1.1.4.3  Aspirate the supernatant without disturbing the cell pellet, leaving approximately 50 μL residual volume in the tube. Mix gently and add 2 mL of filtered wash solution. 
1.1.4.4  Centrifuge for 5 min at 540 x g.
1.1.4.5  Repeat one more time the steps 1.1.4.3-1.1.4.4.
1.1.4.6  Add 250 µL of filtered washing buffer and mix gently. 
1.1.4.7  In the Experiment created for Rainbow Beads acquisition, create a new Specimen | New Worksheet: draw a bi-parametric SSC-A / FSC-A graph. 
1.1.4.8 Acquire the cells, gate lymphocytes in a FSC versus SSC bivariate dot plot and adjust FSC and SSC voltages to reach the following mean target values for the gated lymphocyte population: FSC: 55,000 (range 50,000 – 60,000) and SSC: 13,000 (range 11,000 – 15,000).
1.1.4.9  Acquire and record the data with about 10,000 events. Verify the mean FSC and SSC target values for gated lymphocytes. Readjust FSC and SSC voltage if necessary. 
1.1.4.10  Print screen and store the print of the target channel values that are obtained. 
1.1.5 Fluorescence compensation settings. 
Note: The single-stained compensation controls must be set after the Target MFI settings and FSC/SSC settings have been established. For this step, a PB from a healthy volunteer, Compensation Particles (Table of Materials), lysing solution and washing buffer are needed. The list of fluorochrome-conjugated antibody reagents used to setup the fluorescence compensation matrices and their reference populations are listed in the Table 1. 
1.1.5.1  Label one tube per reagent to be used in setting up fluorescence compensation (FITC, PE, PerCPCy5.5, APC, V450, V500, PECy7 – CD117, APC-H7 – CD10, APC-H7 – CD14 and APC-H7 – CD71) and a “blank/unstained” tube. 

1.1.5.2  Pipet 50 μL of PB into each tube or 1 drop of “negative” Compensation Particles + 1 drop of “positive” Compensation Particles in the compensation control tubes indicated above in Table 1. 
1.1.5.3  Add appropriate amount of the antibody reagent to the tube. Add filtered washing buffer to reach a final volume of 100 µL per tube and mix gently. Incubate for 15 min at RT and protected from light.
1.1.5.4  Add 2 mL of freshly diluted lysing solution only in the tubes with the cells and mix gently. Incubate for 10 min at RT, protected from light. 
1.1.5.5  Centrifuge for 5 min at 540 x g. 
1.1.5.6  Aspirate the supernatant without disturbing the cell pellet leaving approximately 50 μL residual volume in each tube. Mix gently. Add 2 mL of filtered washing buffer. 
1.1.5.7  Centrifuge 5 min at 540 x g. 
1.1.5.8  Aspirate the supernatant without disturbing the cell pellet leaving approximately 50 μL residual volume in each tube. Add filtered washing buffer to reach a final volume of 250 µL per tube and mix gently. 
1.1.5.9  Create Compensation Controls. 

1.1.5.9.1  From the menu bar, select Experiment created for Rainbow Beads acquisition. Create a new Specimen | Compensation Setup | Create Compensation Controls. 

1.1.5.9.2  In the resulting dialog box, select the Include separate unstained control tube/well checkbox. Create generic (not label-specific) compensation controls for FITC, PE, PerCPCy5.5, APC, HV450, HV500. Create label-specific compensation controls for PE-Cy7– CD117, APC-H7 – CD10, APC-H7 – CD14 and APC-H7 – CD71. Click OK.
1.1.5.9.3  In a new worksheet, create a bi-parametric SSC-A / FSC-A graph and draw a gate on lymphocytes (P1) and the histogram corresponding to the fluorochrome that will be detected in each tube and draw a P2 gate for the positive peak. Display the hierarchy (right click on a graph and select Show Population Hierarchy) to visualize the number of events in P2, except for the Unstained Control tube.

1.1.5.9.4 In the Browser, expand the Compensation Controls specimen. 

1.1.5.9.5 Vortex the unstained cells, prepared above, for 3-5 s. Install the prepared unstained cells on the cytometer. Adjust the flow rate to Medium and click Acquire Data. In the FSC-A vs SSC-A dot plot, adjust the P1 gate to fully encompass the lymphocyte population. 
Right-click on the P1 gate. Select Apply to all compensation control. 

1.1.5.9.6  From the Acquisition Dashboard, click Record Data to acquire 5,000 events. For all the single-color stained control cells, verify that the P2 interval gate encompasses the positive population. 
1.1.5.9.7 For the PE-Cy7 and APCH7 tubes, add a P3 interval gate to the histogram and ensure that it encompasses the negative population, and that the P2 encompasses the positive population.
1.1.5.9.8 Calculate Compensation. From the menu bar, select Experiment | Compensation Setup | Calculate Compensation. Name the compensation matrix: Compensations date. Select Link and Save. 
1.1.5.9.9 Save the compensation matrix in the Catalog Application Settings: click on Cytometer Settings | Application settings | Save, name compensation matrix Compensation date and click OK.
1.1.5.9.10 In the Browser, click on Cytometer Settings. In the Inspector, navigate to the Compensation tab. Click Print in the lower right corner. 
Note: This information can also be retrieved from the catalog.
1.1.5.9.11 Control of the compensation matrix.
1.1.5.9.11.1 Mix in one tube all the single stained tube (APCH7 of choice). 
1.1.5.9.11.2 Create a new Experiment named Compensation verification date, add new Specimen and click right on the Cytometer Settings of this experiment and choose Link | Unlink | Application Setting saved in the step 1.1.3.10. Acquire 50,000 events from this tube with the new settings.
1.1.5.9.11.3 Apply a new Global Worksheet. Create 1 dot plot FSC-A/SSC-A and draw a gate to visualize the lymphocytes and n x (n-1)/ 2 other plots focused on the lymphocytes gate to visualize two-by-two parameters. 
1.2 Daily Instrument setup 
1.2.1 Turning on the cytometer. Check that all fluid levels are appropriate and open Diva 6.1.3. Perform fluidics startup: in the menu bar, select Cytometer | Fluidics Startup. Click OK when prompted. Allow cytometer to warm up for at least 30 min. 

1.2.2 Performance Check - CST Beads: repeat the steps 1.1.2.1 – 1.1.2.5. 
1.2.3 Performance Check - Confirmation of PMT Values with Rainbow Beads. 
1.2.3.1 Label a polystyrene tube as Rainbow Beads and check that the lot number is the one in use. Thoroughly mix the Rainbow Bead vial. Prepare the Rainbow Beads, add 1 drop of Rainbow Beads to 1 mL of deionized or distilled water. Protect from light. 
Note: Proceed to acquisition or store the tube at 2-8 ˚C until acquisition. 
1.2.3.2 Create a new Experiment: Rainbow Beads Date. 
1.2.3.3 Link the compensations: right click on Cytometer Settings, select Link Setup, select the appropriate compensation matrix created in step 1.1.5.9.9 and select Overwrite. 
1.2.3.4 Unlink compensation: right click on Cytometer Settings, select Unlink from the previously linked setup and click OK. 
1.2.3.5 Apply Application Settings: right click on Cytometer Settings, select Application Settings, apply the setting created in step 1.1.3.10 during the Monthly Setup and select Keep the compensation value. 
1.2.3.6 Deselect Enable compensation. 
1.2.3.7 Create new Specimen in the Experiment with worksheet template for Rainbow Beads. Acquire the tube in LOW acquisition. 
1.2.3.8 During the beads acquisition, adjust the P1 gate to include only the singlet bead population. Adjust the P2 gate on the FITC-A / PE-A dot plot to include only the singlet bead population. Record 10,000 events. 
1.2.3.9 Check that the MFI and the CV values obtained for P2 population are in the pre-defined targets of the protocol. Otherwise, wash the cytometer and start the operation again. Save the report as PDF format. 
2. BM Sample Preparation 
Note: Perform the cell washing protocol just before the staining procedure.

2.1. Pipette 600 μL of primary sample into a 15 mL centrifuge tube. 
2.2. Add 10 mL of washing buffer (PBS + 0.5% BSA [>98% pure BSA] + 0.09% NaN3 filtered solution) with pH 7.4. Mix the cell suspension well using a pipette. 
2.3. Centrifuge for 5 min at 540 x g (wash 1). Discard the supernatant without disturbing the cell pellet. 

2.4. Repeat steps 2.2–2.3 (wash 2). 
2.5. Suspend the cell pellet in 400 μL of washing buffer. 
2.6. Staining of backbone markers. Transfer the entire volume of the backbone antibodies to a polypropylene tube for FACS analysis, identified with the patient data and the "backbone". Add 350 μL of washed sample (the volume of the washed sample required to fill all the tubes on the panel). Mix well using a pipette. 

Note: Calculate the total volume of backbone antibodies for surface membrane staining (as shown in Table 2).
2.7. Pipette equal amounts of the sample‐backbone mix into 3 polypropylene tubes for FACS analysis, identified with the patient data and “tube number 1” to “tube number 3”. If necessary, use washing buffer to reach a final volume of 200 μL per tube. 

CAUTION: Be careful not to leave any trace of the sample on the walls of the tubes; otherwise, these cells will not be stained. If necessary, vortex the cells and centrifuge the tube.
2.8. In each tube, add the appropriate volume of antibodies directed against cell surface markers (except for the backbone markers), as specified in Table 2. Mix well using a pipette. Incubate for 30 min at RT protected from light.
2.9.  Add 2 mL of lysing solution. Mix well using a pipette. Incubate for 10 min at RT protected from light.
2.10.  Centrifuge for 5 min at 540 x g. Discard the supernatant without disturbing the cell pellet, leaving approximately 50 μL residual volume in each tube. Mix well using a pipette.
2.11.  Add 2 mL of washing buffer to the cell pellet. Mix well using a pipette. 
2.12.  Repeat steps 2.10-2.11 (wash 2).
2.13.  Centrifuge for 5 min at 540 x g. Discard the supernatant without disturbing the cell pellet and re-suspend the cell pellet in 200 μL of PBS. Mix well using a pipette.
2.14. Acquire the cells, preferably, immediately after staining or store at 4 °C, protected from light, for no more than 1 h until measured in the flow cytometer. 
3. Data Acquisition
3.1. Open a New Experiment in Diva software and rename it according to the name, type of sample and date. 
3.2. Create a new Specimen containing 3 tubes. Specify in the Experiment Layout the antibodies used in each tube. 
3.3. Click right on the Cytometer Settings, choose Application Settings and apply the values obtained in Monthly Setup (step 1.1.3.10). 
3.4. Open a new Global Worksheet and create the dot plots: SSC-A/ FSC-A, SSC-A/ CD45-HV500-A, SSC-A/ FITC-A, SSC-A/ PE-A, SSC-A/ CD34-PerCP-Cy5.5, SSC-A/ CD117-PECy7-A, SSC-A/ APC-A, SSC-A/ APCH7-A, SSC-A/ HLA-DR-HV450-A. In the SSC-A/FSC-A dot plot, create a gate to select singlet cells. In the SSC-A/CD45-BV500-A dot plot, create gates to select 4 populations: granulocytes, monocytes, blasts, and lymphocytes. Project these populations in the dot plots created previously. 
3.5. Create a new Global Worksheet for compensation control as described in step 1.1.5.9.11.3. 
3.6. Acquire the tube in MEDIUM acquisition and record 500,000 events/tube. After technical validation (evaluation of compensation and the proper staining), export the data as FCS3.0 files.
4.  Data Analysis
Note: To construct the normal BM databases were used files from healthy donors and from individuals without any evidence for a hematopoietic disease as follows 11 from 18 files for the Neutrophils_NM database, 10 from 18 files for Monocytes_NM database and 14 from 18 files for NRC_NM database. The files discarded showed various technical problems, as presented in the Representative Results section. The files were individually analyzed using the Infinicyt software (Table of Materials), conforming to the various strategies depicted in Figures 1A.1-3 for neutrophil lineage (Profile Neutrophils_Maturation.inp), Figure 2A for monocyte lineage (Profile Monocytes_Maturation.inp), and Figure 3A.1-2 for erythroid cell lineage (Profile NRC_Maturation.inp).
4.1 Strategy of analysis for neutrophils
4.1.1 Identify CD34+ neutrophil-committed blasts using an intersection of seven gates, allowing the selection of CD34+ CD117+ HLADR+low CD10- CD13+ CD11b- events (Figure 1A.1). Assign these events to the Neutrophil tab in the Population Hierarchy Tree and thereafter uncheck this tab in order to remove these cells (depicted in blue) from the display of the remaining events (gray).
4.1.2 Isolate the CD117+ CD34- CD13+ CD11b- HLADR+low neutrophil precursors using an intersection of six gates as depicted in Figure 1 A.2 and assign them to the “Neutrophil” tab.
4.1.3 Identify more mature neutrophils using an intersection of four gates, allowing for discrimination of CD45dim SSCint-hi CD117- HLADR-cells and their assignment to the Neutrophil tab.
4.1.4 Uncheck the remaining events, keeping visible only the neutrophils, and then export this population by clicking File | Export and verifying that all the required parameters are checked and save data as FCS files). 

4.1.5 In a merged file consisting of all exported FCS files, perform a quality check by evaluating the intensity of expression of markers for each subpopulation. Using APS plots with medians for each file and SD curves for each subpopulation shown, remove cases outside the 2SD curves (see details in the “Representative Results” section) (Figure 1B). 
4.1.6 In the resulting composed file with “Neutrophils” visible, draw the Maturation Pathway on an APS diagram (Figure 1C left) and save as a .cyt file. 
Note: A comparison Maturation Diagram allows visualization of all parameters from all files included in the Neutrophils_NM database represented against the normalized database. The diagram presented in Figure 2C, right side, shows that all files included in the Neutrophils_NM database (n=11) fit in 2 SD compared with median of the group.

4.2 Strategy of analysis for monocytes

4.2.1 Identify monocytic lineage cells (CD117+/- CD64+hi HLADR+hi) using an intersection of four gates (Figure 2. A). Assign these events to the Monocytic tab in the Population Hierarchy Tree.
4.2.2 Uncheck the remaining events, keeping visible only the monocytic cells, and then export this population by clicking File | Export and verifying that all the required parameters are checked and save data as FCS files. 

4.2.3 In a merged file consisting of all exported FCS files, perform a quality check by evaluating the intensity of expression of markers for each subpopulation, then remove cases outside the 2SD curves (details in “Representative Results” section) (Figure 2B).

4.2.4 In the resulting file with “Monocytic” cells visible, draw the Maturation Pathway on the APS diagram (Figure 2C left) and save this as a .cyt file. 
Note: A comparison Maturation Diagram allows visualization of all parameters from all files included in the Monocytes_NM database represented against the normalized database. The diagram presented in Figure 2 C, right side, shows that all files included in the Monocytes_NM database (n=10) fit in 2 SD compared with median of the group.

4.3 Strategy of analysis for nucleated red cells (NRCs)

4.3.1 Identify CD34+ erythroid committed blasts using an intersection of seven gates that allow selection of CD34+ CD117+ HLADR+low CD105+ CD33- CD36+ CD71+ events (Figure 3A.1). Assign these events to the “NRC” tab in the Population Hierarchy Tree and then uncheck this tab in order to remove these cells (depicted in red) from the display of the remaining events (gray).

4.3.2 Identify more mature NRCs using an intersection of four gates that allow discrimination of CD45-/+dim SSClow CD36+hi CD71+hi CD105+/-cells. Assign these events to the “NRC” tab (Figure 3A.2). The platelets (CD36+hi SSClow cells) must be removed from the NRC population (Figure 3A.2).

4.3.3 Uncheck the remaining events, keeping visible only the NRC cells, and then export this population by clicking File| Export and verifying that all the required parameters are checked and save data as FCS files.

4.3.4 In a merged file consisting of all exported FCS files, perform a quality check by evaluating the intensity of expression of markers for each subpopulation, followed by removal of cases outside the 2SD curves (see details in “Representative Results” section) (Figure 3B).
4.3.5 In the resulting file with “NRC” cells visible, draw the Maturation Pathway on the APS diagram (Figure 3C left) and save this as a .cyt file. 
Note: A comparison Maturation Diagram allows visualization of all parameters from all files included in the NRC_NM database represented against the normalized database. The diagram presented in Figure 3 C, right side, shows that all files included in the NRC_NM database (n=14) fit in 2 SD compared with median of the group.

4.4  Evaluation of maturation in BM myeloid compartments using the Maturation Databases 
4.4.1  Open the .cyt file corresponding to the lineage of interest (i.e., Neutrophils_NM_GMFF.cyt for neutrophil lineage, Monocytes_NM_GMFF.cyt for monocytic lineage, and NRCs_NM_GMFF.cyt for erythroid lineage). 

4.4.2  Right-click on Maturation tab below the tab corresponding to the lineage of interest (Neutrophils | Monocytic| NRC) and save the Maturation to Maturation Database. 

4.4.3  Open a new FCS file and perform analysis as explained previously (step 4.1.1–4.1.3 for Neutrophils, step 4.2.1 for Monocytic cells, and step 4.3.1–4.3.2 for NRC). 

4.4.4  Draw the maturation pathway for the population of interest.
4.4.5  Open the corresponding database in the Tools tab (Database Analysis) and compare the population to be analyzed with the corresponding Maturation Database. Check data for compatibility with the available database: complete compatibility (green triangle); partial compatibility, in most cases discrepancies in the name of the parameters (yellow triangle); and incompatibility (red triangle).
4.4.6  If data are compatibles or partial compatibles, the software creates the Normalized Maturation Differences diagram. To visualize the Parameter Band Maturation Differences, open a new diagram in Diagram tab, click Maturation and choose how many parameters to be displayed, click OK and the diagram appear. With right click in the diagram; changes can be made in Data Visualization, Database Visualization and Maturation Diagram Visualization. 
4.4.7  To visualize the significance of the differences between the new file and data included in the database, configure a zoom (right-click on Normalized Maturation Differences and apply Zoom).
REPRESENTATIVE RESULTS 
The 54 BM samples harvested in K-EDTA anticoagulant were included in the study. The MFC data were analyzed in the absence of any information about the patients. Retrospective study showed that the BM samples were from 7 healthy donors (5 males and 2 females with a median age of 47.4 [35-48 years], 11 individuals with no evidence of a hematopoietic disease (8 males and 3 females with a median age of 57.9 years [35-72 years]) and 36 cases with various pathological conditions: 1 case with anemia and low creatinine level, 3 cases with anemia and high creatinine level, 8 cases with anemia of inflammation, 1 case with anemia caused by vitamin B12 deficiency, 4 cases with anemia ± other cytopenias caused by liver damage, 3 cases with autoimmune hemolytic anemia, 5 cases with idiopathic thrombocytopenic purpura, 1 case with macrophage activation syndrome, 3 cases with complete remission after lymphoma treatment (minimal residual disease < 0.01%) and 7 cases carrying hematological disorders (4 MDS, 1 AML, 1 chronic myelomonocytic leukemia and 1 myeloproliferative syndrome JAK2 positive). This last category included 18 males and 18 females with a mean of age 60.3 [17-100 years]. All samples were from Caucasian individuals.
The building databases allowed identifying and fixing several issues related to sample preparation and acquisition. In our case, the final databases included 11/18 files for neutrophil maturation, 10/18 files for monocyte maturation and 14/18 files for NRC maturation. The files that were not included in databases posed various technical problems. The most common staining issues were observed for CD11b and CD13 in neutrophils, CD300e and HLADR in monocytes, and CD71 and HLADR in NRCs. Exporting data can be a source of errors; the most frequent in our dataset was the absence of FSC-H in exported files and the inversions of FSC-W with FSC-H that sometimes occur in exporting DIVA files8.
The evaluation of new cases of cytopenias suspected of being MDS against the Myeloid Normal Maturation Databases allow for identification of abnormal expression of maturation antigens of neutrophils lineage (Figure 4), monocytes (Figure 5), and NRC (Figure 6) even in cases without cytological or cytogenetic abnormalities (Figure 7). Otherwise, using routine acquisition software, such as Diva, these would be difficult or impossible to realize. Please insert here the figures 4, 5, 6 and 7.
Moreover, interpretation of the significance of differences between antigen expression on pathological settings and normal counterparts allows for “quantification” of these abnormalities, with identification of those that are discriminant for a group of cases. This may also make it possible to rank them based on importance for the purposes of a follow-up.
In this study, all 7 cases of hematological disorders with myeloid dysplasia features (4 MDS, 1 AML, 1 chronic myelomonocytic leukemia and 1 myeloproliferative syndrome JAK2 positive) were classified as abnormal when compared with the NBM databases. In addition, evaluation against the Neutrophils_NM database eliminated suspicion of myeloid dysplasia in 7 cases with toxic, inflammatory, autoimmune hemolytic anemia and anemia from chronic kidney disease. The evaluation against the Monocytes_NM database eliminated suspicion of myeloid dysplasia in 4 cases with idiopathic thrombocytopenic purpura (ITP), while evaluation against NRC_NM allowed for differentiation in 3 cases, 2 of inflammatory anemia and 1 of ITP. A representative ITP case is presented in Figure 8. The BM aspirates from patients in complete remission after lymphoma or solid tumor treatments were within 2 SD from median of the normal databases for all three lineages (Figure 9), and thus these samples may be alternatives to healthy donor BM samples.
FIGURES AND TABLES LEGENDS
Figure 1. Analysis strategies for neutrophil cell lineage. Selection of CD34+ CD117+ HLADR+low CD10- CD13+ CD11b- neutrophil progenitors was realized by intersection of seven gates as depicted in Figure 1 A.1. The next step of maturation, CD117+ CD34- CD13+ CD11b- HLADR+low neutrophil precursors was chosen using an intersection of six gates as depicted in Figure 1 A.2. The more mature neutrophils are finally identified using an intersection of four gates that allow discrimination of CD45dim SSCint-hi CD117- HLADR- cells (Figure 1 A 3). The most discriminant markers for neutrophil lineage were CD34, CD117, CD11b, and CD16. Along with CD13, these parameters allow identification of five subpopulations: CD34+ progenitors (CD34+ CD117+ CD13+low CD11b- CD16-) (dark blue); CD34- CD117+ CD13+hi CD11b- CD16- neutrophil precursors (blue); CD34- CD117- CD13+low CD11b- CD16- neutrophils (3rd step of maturation; light blue); CD34- CD117- CD13+low CD11b+ CD16+low neutrophils (4th step of maturation; pink); mature neutrophils (CD34- CD117- CD13+hi CD11b+hi CD16+hi; violet). Each colored circle represents the median of a subpopulation for the marker of interest from one sample (Figure 1B). Figure 1C displays the homogeneous distribution of the tested parameters during the neutrophil cell maturation when compared with 2 SD of the normal maturation database (n=11).
Figure 2. Analysis strategy for monocytes. Identification of monocytic lineage cells (CD117+/- CD64+hi HLADR+hi) is performed using an intersection of four gates (Figure 2. A). The most discriminant markers for monocytic lineage were CD14, CD300e (IREM2), CD35, and HLA-DR. Along with CD117, these parameters allow identification of three subpopulations: CD34- CD117+low/- HLADR+hi CD35- CD14- CD300e- (red); CD34- CD117- HLADR+med CD35+med CD14+med CD300e- (orange); mature monocytes CD34- CD117- HLADR+med CD35+hi CD14+hi CD300e+ (green) (Figure 2B). Figure 2C displays the homogeneous distribution of the tested parameters during the monocytic cell maturation when compared with 2 SD of the normal maturation database (n=10).
Figure 3. Analysis strategies for NRCs. The identification of CD34+ erythroid committed blasts was realized using an intersection of seven gates that allow selection of CD34+ CD117+ erythroid progenitors (Figure 3 A.1). The more mature NRCs are identified using an intersection of five gates (Figure 3 A.2). Exclusion of the platelets (CD36+hi SSClow cells) from NRC population is required (Figure 3 A.2). The most discriminant markers for erythroid cell lineage were CD34, CD117, CD71, and CD105. Along with CD33, these parameters allow identification of three subpopulations: CD45+low CD34+ CD117+ HLADR+med CD71+med CD36+med CD105+hi (red); CD34- CD117+ HLADR+low CD71+hi CD36+hi CD105+hi NRCs (2nd step of maturation; pink); CD34- CD117- HLADR-/+low CD71+hi CD36+med CD105+low/- NRCs (more mature NRCs; pink salmon) (Figure 3B). Figure 3C displays the homogeneous distribution of the tested parameters during the NRC maturation when compared with 2 SD of the normal maturation database (n=14).
Figure 4. Representative flow cytometry analysis of neutrophils in a del(7) MDS case. A. In the Normalized Maturation Differences diagram, the gray area corresponds to the normal expression of antigens resulting from the analysis of the 11 normal files included in the neutrophil database (median± 2SD). The abnormal expression of several antigens was observed when compared the del(7) MDS case against Neutrophils_NM database (n=1 versus n=11 normal controls samples). B. Parameter Band Maturation Diagrams allow comparisons between the median intensity of expression of each marker at different stages of maturation (continuous full lines) and 2SD curves calculated for the 11 normal BM cases included in the database (continuous dashed lines). Phenotypic abnormalities of neutrophils observed in a del(7) MDS case as compared with the Neutrophil Normal Maturation Database are as follows: overall increased expression of CD34, slight increase of CD11b expression on immature neutrophils (stages 2), diminished expression of CD13 on immature neutrophils (stages 1-3) and slight increase of CD16 expression on the first two stages of neutrophil maturation followed by moderate increases of CD16 expression on the mature neutrophils (stages 4-5). The panel C shows the corresponding dot plots as visualized in DIVA software. Image evaluation using the DIVA software allows for identification of a small number of phenotypic abnormalities in this MDS case: an increased percentage of CD34+ CD13+ precursors and the downregulation of CD13 expression. The abnormal expression of CD16 is difficult to observe in this type of representation.
Figure 5. Representative flow cytometry analysis of monocytes in a del(7) MDS case. A. The Normalized Maturation Differences diagram offers an overall view of antigen expression at various stages of monocytic cell lineage maturation. The gray area corresponds to the normal expression of antigens resulting from the analysis of 10 normal files included in the monocytic database (median±2SD). The abnormal expression of several antigens exceeds 2SD, but for some markers, such as CD35 and HLADR, it exceeds 4SD. B. In the Parameter Band Maturation Diagrams, the phenotypic abnormalities of monocytic cells observed in a SMD del(7) case when compared with Monocyte Normal Maturation Database (n=1 versus n=10 normal controls samples) are as follows: diminished expression of CD45 during stages 1–4 of monocytic cells, of CD117 at stages 1-2 of monocytic precursors, and of HLA-DR overall. The increased expression of CD35 in the first 3 stages of maturation of the monocytic cells was the most significant abnormality for this lineage. C. Panel C shows the DIVA dot plots that allow evaluation of monocytic cells in a single normal BM. D. Panel D shows the DIVA dot plots that allow evaluation of monocytic cells in del(7) MDS case. Image evaluation using the DIVA software allows for identification of two phenotypic abnormalities in this MDS case: the diminished expression of HLA-DR and the increased expression of CD35 in a part of monocytic precursors (red population).
Figure 6. Representative flow cytometry analysis of nucleated red cells in a del(7) MDS case. A. The Normalized Maturation Differences diagram offers an overall view of antigen expression at different stages of NRC lineage maturation. The gray area corresponds to the normal expression of antigens resulting from the analysis of the 14 normal files included in the NRC Normal Maturation Database (median±2SD). The abnormal expression of several antigens exceeds 2SD, but for some markers, such as CD34, CD117, CD71, and CD105, it exceeds 4SD. B. In the Parameter Band Maturation Diagrams, the phenotypic abnormalities observed in a SMD del(7) case when compared with the NRC Normal Maturation Database (n=1 versus n=14 normal controls samples) are as follows: diminished expression of CD34 in the first two stages of maturation, of CD117, CD105, CD71 and HLA-DR on stage 1 of erythroid lineage precursors. Image evaluation using the DIVA software allows for identification of CD71 abnormal expression on erythroid cells in this MDS case, but other modifications are not detectable.
Figure 7. Representative flow cytometry analysis of neutrophils, monocytes and NRCs in an early stage MDS case without morphological dysplasia and without cytogenetic aberrancies.  A. Panel A shows the distribution of immature and mature cells within different maturation stages for the three myeloid lineages: neutrophils, monocytes and NRCs. An increased number of mature monocytes (stage 5) and the slightly diminution of monocytic immature cells (stage 2 of maturation) was observed. B. The Normalized Maturation Differences diagram shows that the phenotypic abnormalities exceed 2SD, but for CD34, CD117, CD11b, CD16 and HLADR, they are higher at 4SD. Phenotypic abnormalities observed for neutrophils are as follows: slightly increased SSC values in neutrophil immature precursors (stage 1), diminution of expression of CD34 (stage 1), of CD117 (stage 1), of CD16 (stage 1-2) and of HLADR (stages 1-3 of maturation). An increased expression of CD11b was observed on neutrophil immature precursors (stage 1). C. The Normalized Maturation Differences diagram shows that the most important phenotypic abnormalities observed for the monocytic lineage (exceeding 10SD) are increased expression of CD35 and early acquisition of CD300e on the immature monocytic precursors. Other phenotypic abnormalities observed on the monocytic lineage are as follows: diminution of FSC and SSC on more mature monocytes (stages 3–5) and slight diminution of expression of CD14 on stages 2-3 of monocytic cells. D. The Normalized Maturation Differences diagram and the Parameter Band Maturation Diagrams shows that the most important phenotypic abnormalities observed for the NRC are: increased expression of CD117 (exceeding 10SD) on stages 2-3 of NRC maturation, increased expression of CD34 (exceeding 5SD) during the second stage of maturation of NRC and decreased expression of CD36 on the early erythroid precursors (stages 1–3). In addition, the SSC on immature NRC is lower than on the normal counterparts.

Figure 8. Representative example of myeloid cell maturation in an idiopathic thrombocytopenic purpura case when compared with the normal maturation databases

Figure 8 displays the distribution of the tested parameters within 2 SD of the Neutrophils Normal Maturation Database (n=1 versus n=11; Figure 8 top), Monocyte Normal Maturation Database (n=1 versus n=10; Figure 8 middle), and NRC Normal Maturation Database (n=1 versus n=14; Figure 8 bottom).

Figure 9. Representative example of myeloid cell maturation in a bone marrow sample from a donor in complete remission after lymphoma treatment when compared with the normal maturation databases

Figure 9 displays the distribution of the tested parameters within 2 SD of the Neutrophils Normal Maturation Database (n=1 versus n=11; Figure 9 top), Monocyte Normal Maturation Database (n=1 versus n=10; Figure 9 middle), and NRC Normal Maturation Database (n=1 versus n=14; Figure 9 bottom).
Table 1. List of fluorochrome-conjugated antibody reagents used to setup the fluorescence compensation matrices with their reference populations

Table 2. Combination and quantities of antibodies used for evaluation of maturation of neutrophil, monocytes and erythroid cells in BM

DISCUSSION

The quality of BM aspirate could impact on the final results. The hemodilution of the BM aspirate could distort the distribution of cells in different stages of maturation due to the absence of progenitors or precursors cells. Probably employing a bulk lysing method may help in normalization of BM aspirates for hemodilution in flow cytometric analyses. In addition, the critical steps for evaluation of BM myeloid dysplasia by flow-cytometry are the sample processing and staining, data acquisition, and interpretation2,3. The sample processing and staining should be performed up to 72 hours after BM harvesting. The data should be acquired, preferably, immediately after staining or store at 4°C, protected from light, for no more than 1 hour until measured in the flow cytometer. The database-guided interpretation avoids the subjective evaluation of BM myeloid cell maturation.

The antibody combinations and the sample preparation were broadly conformed to the EuroFlow procedures9,10. The difference compared with the EuroFlow panel was that all antibodies except CD300e were provided by BD Bioscience. The flow cytometry standardization made it possible to obtain data with high levels of reproducibility11, but this entails joining a group working on standardization of the panels in question. In our group, the standardization of SSC remains an area to be improved. The major issue regarding MFC databases building is the staining procedure. The initial distribution of the non-backbone antibodies in FACS tubes before adding the equal amounts of the sample‐backbone mix avoids repeating the backbone staining in case of eventual errors in distribution of other non-backbone antibodies. In addition, to overcome the staining problems there are two possibilities: either increase the incubation time of antibodies from 15 to 30 minutes, or use the recombinant antibodies, which provide greater reproducibility, according to the manufacturer’s specifications12. The analysis was performed in conformity to recently published data13-15. Concerning monocytes analysis, we frequently observed the absence of CD34+ CD117- monocytic precursors, as has been previously reported by Shen et al.16. For this reason, we eliminated from the analysis strategy a supplementary gate for isolation of this particular population. The intersection of gates allowing isolation of CD64+F HLADR+F/int and CD117+/- includes the immature precursors (CD117+ CD34+low/-) and the more mature monocytes.

The evaluation of new FCS files compared to a database contributes to a more objective, standardized analysis and data interpretation and avoids misinterpretations of the so-called patterns recognition, which is difficult to standardize3. The quantification of the amplitude of phenotypic abnormalities by comparison with normal databases makes it possible to rank these abnormalities, which may help improve MFC scores for MDS diagnosis. Moreover, this method allows precise identification of the phenotypic abnormalities on immature CD34+ and/or CD117+ precursors and in more mature cells, which are not evident when using the current analysis strategies in acquisition software. This method is relevant in the diagnosis of MDS cases that do or do not have evident morphologic abnormalities, or that do or do not carry cytogenetic recurrent aberrancies.
Improvement of antibody staining is needed, and the addition of new normal BM data in databases is required in order to increase the robustness, reliability, and sensitivity of analysis. 

This method could be applied, with further investigation, to cytopenias from other causes, or to clonal hematopoiesis of indeterminate potential (CHIP), in order to identify the phenotypic changes reliably related to dysplastic processes.
ACKNOWLEDGMENTS 
The antibodies used in this study were provided by BD Biosciences. The authors would like to thank their colleague, Dr. Pascale Flandrin-Gresta, from the Department of Molecular Biology, Hematology laboratory, University Hospital of Saint-Etienne, France, who provided expertise for interpretation of NGS data for the second MDS case. The authors are thankful for the clinician hematologists for their interest and involvement in this study and for the patients and healthy donors for their agreement to participate in this study. The authors would also like to thank the “Les Amis de Rémi” Foundation for financial support for publication.
DISCLOSURES 
The authors declare that they have no competing financial interests. The Flow-Cytometry Laboratory at the University Hospital of Saint-Etienne is a member of EuroFlow Consortium.

REFERENCES
1. Orfao, A., Ortuño F., de Santiago, M., Lopez, A. & San Miguel, J. Immunophenotyping of Acute Leukemias and Myelodysplastic Syndromes. Cytometry Part A. (58), 62–71, doi: 10.1002/cyto.a.10104 (2004).
2. Porwit, A. et al. Revisiting guidelines for integration of flow cytometry results in the WHO classification of Myelodysplastic Syndromes - proposal from the International/European LeukemiaNet Working Group for Flow Cytometry in MDS (IMDSFlow). Leukemia. 28(9), 1793-1798, doi: 10.1038/leu.2014.191 (2014).
3. Westers, M. T. et al. Implementation of flow cytometry in the diagnostic work-up of myelodysplastic syndromes in a multicenter approach: Report from the Dutch Working Party on Flow Cytometry in MDS. Leukemia Research. (36), 422-430, doi: 10.1016/j.leukres.2011.09.015 (2012).
4. Meehan, S. et al. AutoGate: automating analysis of flow cytometry data, Immunologic Research (58) 218–223, doi: doi: 10.1007/s12026-014-8519-y (2014).
5. Greenberg, P. L. et al. NCCN Clinical Practice Guidelines in Oncology: myelodysplastic syndromes. Journal of the National Comprehensive Cancer Network. (9), 30-56, PMID: 21233243 (2011).
6. Greenberg, P. L. et al. Myelodysplastic Syndromes, Version 2.2017, Clinical Practice Guidelines in Oncology. Journal of the National Comprehensive Cancer Network. (15-1), PMID: 28040720 (2017).
7. Swerdlow, S. H. et al. eds. WHO Classification of Tumours of Haematopoietic and Lymphoid Tissues. Revised 4th Edition ed. Lyon, France. IARC Press: 2017, (2), 86-128 (2017).
8. https://lists.purdue.edu/pipermail/cytometry/2009-March/036889.html

9. Composition of EuroFlow panels and technical information on reagents Version 1.5 (2017).
10. EuroFlow Standard Operating Procedure (SOP) for sample preparation and staining Version 1.2.1 11 (2015).
11. Kalina, T. et al. EuroFlow standardization of flow cytometer instrument settings and immunophenotyping protocols. Leukemia. (26), 1986–2010, doi: 10.1038/leu.2012.122 (2012).
12. http://www.miltenyibiotec.com/en/products-and-services/macs-flow-cytometry/reagents/antibodies-and-dyes/recombinant-antibodies.aspx 

13. van Dongen, J. J. M. et al. EuroFlow antibody panels for standardized n-dimensional flow cytometric immunophenotyping of normal, reactive and malignant leukocytes. Leukemia. (26), 1908–1975, doi: 10.1038/leu.2012.120 (2012).
14. Matarraz, S. Introduction to the Diagnosis and Classification of Monocytic-Lineage Leukemias by Flow Cytometry. Cytometry Part B (Clinical Cytometry). 92(3):218-227. doi: 10.1002/cyto.b.21219 (2015).
15. Westers, M. T. Immunophenotypic analysis of erythroid dysplasia in myelodysplastic syndromes. A report from the IMDSFlow working group. Haematologica. (102-2), 308-319, doi: 10.3324/haematol.2016.147835 (2017).
16. Shen, Q. et al. Flow cytometry immunophenotypic findings in chronic myelomonocytic leukemia and its utility in monitoring treatment response. European Journal of Haematology. (95-2), 168-76, doi: 10.1111/ejh.12477 (2015).

