TITLE:

Analyzing Protein Architectures and Protein-Ligand Complexes by Integrative Structural Mass Spectrometry

AUTHOR:

Zainab Ahdash1, Andy M. Lau1, Chloe Martens1, Argyris Politis1

1Department of Chemistry, King’s College London, London, United Kingdom

zainab.ahdash@kcl.ac.uk

andy.lau@kcl.ac.uk

chloe.martens@kcl.ac.uk

argyris.politis@kcl.ac.uk

Corresponding Author:

Argyris Politis

argyris.politis@kcl.ac.uk

Keywords:

Mass spectrometry (MS), native MS, ion-mobility, protein complexes, non-covalent interactions, nucleotide binding, DNA binding, molecular dynamics, modelling.

Summary:

Mass spectrometry (MS) has emerged as an important tool for the investigation of structure and dynamics of macromolecular assemblies. Here, we integrate MS-based approaches to interrogate protein complex formation and ligand binding.

Abstract:

Proteins are an important class of biological macromolecules that play many key roles in cellular functions including gene expression, catalyzing metabolic reactions, DNA repair and replication. Therefore, a detailed understanding of these processes provides critical information on how cells function. Integrative structural MS methods offer structural and dynamical information on protein complex assembly, complex connectivity, subunit stoichiometry, protein oligomerization and ligand binding. Recent advances in integrative structural MS have allowed for the characterization of challenging biological systems including large DNA binding proteins and membrane proteins. This protocol describes how to integrate diverse MS data such as native MS and ion mobility-mass spectrometry (IM-MS) with molecular dynamics simulations to gain insights into a helicase-nuclease DNA repair protein complex. The resulting approach provides a framework for detailed studies of ligand binding to other protein complexes involved in important biological processes.

Introduction:

Native mass spectrometric analysis of intact proteins and their complexes is carried out using electrospray and nano-electrospray ionization (nESI), which preserve protein folding and non-covalent interactions during the ionization process1,2. In native MS, the structure of proteins and their complexes are retained in a near-native state in the gas-phase3,4. Native MS detects multiple charged protein ions, which are separated according to their mass to charge ratio (m/z) allowing the mass of the protein or protein-ligand complex to be calculated. This information enables the determination of an intact protein’s stoichiometry, subunit composition, ligand binding, and interaction networks3-6. Native MS has several advantages compared to other techniques such X-ray crystallography and nuclear magnetic resonance spectroscopy5. Firstly, native MS is a rapid and highly sensitive technique, requiring only a few microliters (2-3 μL) of sample at relatively low final complex concentrations in the high nM to low μM range6. Secondly, native MS can be used to interrogate heterogeneous protein samples making it possible to analyze multiple proteins and oligomeric states simultaneously. Thirdly, native MS does not require protein samples to be modified before analysis by chemical crosslinking or protein labelling. These advantages have made structural MS a powerful tool for the structural investigation of protein complexes.

Native MS can be combined with ion mobility (IM), a technique that measures the time a protein ion takes to travel through an electric field, enabling the collisional cross section (CCS) to be determined. The CCS provides low-resolution structural information, which enables topology and conformational heterogeneity information of proteins to be obtained. Furthermore, it allows the examination of protein structural models generated by computational approaches.

Protein gas-phase stability can be investigated using collision induced unfolding (CIU) measured by IM-MS. During the CIU process, protein ions are accelerated and activated through increased accelerating collisions with an inert buffer gas within a mass spectrometer7-9. This collisional activation process causes the protein to partially unfold, which translates into an increase in CCS. This change in CCS and the energy required to unfold the protein can be measured by IM-MS. Using this approach, the effect of ligand binding on protein stability can be measured10. Subcomplexes can be generated in solution using in-solution disruption methods such as the addition of organic solvents to monitor native-like topologies of protein complexes. Disruption of the protein complexes is mainly due to the disruption of intra non-covalent interactions. The sub-complexes maintain native-like topologies and, upon MS detection, reveal information about inter-subunit connectivity.

Integrative approaches in structural biology combine diverse methods to study the structure and dynamics of proteins and their complexes3-6. Native MS and IM-MS have been used to uncover the molecular details of challenging biological systems. There have been several examples of applications including the study of protein assembly pathways11-14, studying protein-protein interaction networks15-17, membrane proteins6,18-21, and protein-ligand interactions such as nucleic acids22-24.

However, native MS also has its limitations. Native MS measurements are often performed in volatile buffers such as aqueous ammonium acetate in which some proteins will not retain their folded native state3,25. Nevertheless, recent work has shown that this limitation can be overcome by optimization of spraying needle tip diameter (0.5 mm tips) such that protein and protein complex ions can be formed directly from non-volatile buffers with high-ionic-strength that better mimic the physiological environment26. Additionally, native MS uses electrospray to ionize and transfer non-covalent assemblies from solution to the gas phase; therefore, the relative abundance of detected complexes may not wholly represent that in solution5,27. Moreover, in comparison to in solution, the gas phase hydrophobic interactions become weaker and electrostatic interactions become stronger and hence favored3,28.

In this article, we provide protocols, data analysis, and interpretation for protein identification and ligand binding using native MS, IM-MS, CIU, in-solution disruption, and modelling. The DNA repair complex, HerA-NurA, is used as a model system. DNA double-stranded breaks (DSBs) are one of the most cytotoxic and deleterious forms of DNA damage, resulting in genetic instability and the eventual development of cancer in humans. Homologous recombination is the repair mechanism which eradicates DSBs, a process which is orchestrated by the ATP dependent helicase-nuclease complex, HerA-NurA22.

Combining native MS and IM-MS with functional assays and modelling allowed the investigation of: i) the role of NurA in the assembly, conformation, and stability of the complex, ii) the interaction between dsDNA and the complex and its influence on the overall stability of the complex, and iii) the stoichiometry and impact of ATP binding on the assembly22. Overall, this work led to an improved understanding of the molecular basis of the HerA-NurA complex by linking protein complex conformational changes and stability with nucleotide binding. This protocol is generic for any protein complex(es) which interacts with one or several ligand(s) types.

Protocol

- Sample Preparation for Native MS of Protein and Protein-Ligand Complexes

Note: To gain an understanding of the molecular basis of a protein complex and ligand binding using native MS, suitable sample preparation is key. The aim of this section is to highlight the essential sample preparation steps prior to MS analysis using the HerA-NurA complex which binds DNA and nucleotides as an example.

  - Prepare 20 µL aliquots of concentrated purified protein (typically 15 – 30 µM) in a 1.5 mL tube.

  - For ATP or ADP binding analysis

Note: Add increasing concentrations of the non-hydrolyzable ATP analog adenosine 5′-O-(3′thiotriphosphate), tetralithium salt (ATP-γ-S) or adenosine 5′-diphosphate (ADP). Non-hydrolyzable ATP derivatives generate a stable complex which would enable the ATP-bound protein to be captured. Other non-hydrolysable ATP analogues that could be tested include AMP-PNP and ATP-γ-S-Mg2+.

    - For HerA-NurA studies, mix 5 μM of purified protein with ATP-γ-S and ADP at concentrations ranging from 0-1 mM.

    - To capture simultaneous ATP-γ-S and ADP binding, add both nucleotides at the same or varying concentrations.

    - Add 2 mM MgCl2 and incubate at 25 °C in a dry bath incubator for 1 h.

Note: Analysis of nucleotide binding using nESI native MS can result in artefactual binding at high concentrations, therefore non-specific binding must be taken into account29. To investigate non-specific binding, add a higher concentration of nucleotides between 2-5 mM).

  - For DNA binding analysis

    - Mix the protein and DNA at a molar ratio that allows for protein-DNA complex formation. For HerA and HerA-NurA, mix 5 μM of purified protein with DNA at a 1:1 ratio.

    - Incubate the HerA-NurA or HerA - DNA mixture for 30 min at 25 °C in a dry bath incubator until it reaches equilibrium. The duration and temperature of incubation may vary depending on the protein under investigation.

    - Buffer exchange protein samples to MS compatible buffers. Commonly, aqueous ammonium acetate solution between 5 mM-1 M at pH 7-8 are used. Other MS compatible buffers include ethylenediammonium diacetate (EDDA) and Triethylammonium acetate (TEAA)30. For HerA-NurA studies, use 200 mM ammonium acetate pH 7.

Note: There are several methods for buffer exchange prior to analysis by MS such as using a spin concentrator or chromatography columns. Native MS is mostly limited by the quality of the sample such as buffers and adducts used during purification. Therefore, it is essential to perform sufficient desalting to obtain resolved peaks.

    - For HerA-NurA ligand binding studies, buffer exchange samples 6–8 times into 200 mM ammonium acetate using a concentrator. Although this method is more time consuming, it ensures that resolved peaks are achieved and allows for accurate mass determination of ATP/ADP bound species.

- Native MS Acquisition and Analysis for Investigating Protein Complexes and Protein-Ligand Complexes

Note: MS conditions should be optimized to achieve highly resolved peaks to enable accurate mass measurements. This section details optimized parameters on a Q-ToF mass spectrometer with a 32k upper limit m/z quadrupole.

  - Prepare in-house capillaries for nano-electrospray and perform instrument mass calibration for accurate mass measurements as detailed by Kirshenbaum et al.1.

  - Select sensitivity, positive ion acquisition and mobility TOF modes.

  - Turn on the Trap, API and IMS gases. For IM separation, use nitrogen (60 mL/min) and argon (8.4 mL/min for the trap region) as starting points and then adjust.

  - Set an appropriate m/z acquisition range. For an unknown protein, initial optimization steps should use a wide range such as 500-32000 m/z.

  - Load 2-3 µL of the protein complex solution to be analyzed into a gold coated capillary and insert it into a capillary holder.

  - Gently tighten the capillary and place the capillary in the electrospray source stage and slide the stage into position to start acquiring data.

  - Apply low nano-flow gas pressure (0.00-0.05 Bar) until a drop is formed at the tip of the capillary. The nano-flow pressure can then be dropped until the spray is maintained.

  - Adjust the capillary with respect to the cone by moving the capillary in x, y, z positions and monitor the ion current to achieve a stable ion current. Apply a capillary voltage in the range of 0.9-1.6 kV.

  - Set the sampling cone (50-120 V), source offset (60.0), source temperature (25 °C) and cone gas flow (0.0 L/h). These suggested initial conditions can be adjusted.

  - To acquire a well resolved mass spectrum and to maximize ion transmission, adjust MS parameters and monitor the resulting change in the spectra. These include adjusting the gas flow in the Trap (2-8 mL/min), He Cell (180 mL/min) and IMS cell (90 mL/min) to achieve best separation at maximum transmission.

  - Adjust the trap collision energies if voltage offsets are insufficient. An optimal starting point is between 10-50 V.

Note: Increasing the trap energy can remove non-covalently bound adducts. However, take care to avoid collision induced dissociation and unfolding of the protein-ligand complex. Perform ion mobility measurements to check if instrument conditions retain the protein in the native folded state (Step 3).

  - Improve desolvation by optimizing the trap bias voltage. An optimal starting point is 20-45 V.

  - Optimize the wave velocity and wave height to achieve best mobility separation. A detailed explanation and protocol can be found here31. For the HerA-NurA studies, use wave velocity of 40 (m/s) and wave height of 550-650 (V).

  - Use all other parameters as instrument default values.

  - Prepare a ligand-free sample for analysis as a control for each run (Figure 1). For ligand binding experiments, perform at least three independent measurements.

  - Use the Masslynx software to measure masses of generated species and identify the ligand binding, such as ATP and ADP binding and oligomeric states (Figure 2 and 3). Other software available include UniDec32, PULSAR33 and Amphitrite34.

  - To quantify the relative abundance of species, use the corresponding ion intensities observed in the raw ESI-MS spectra (for example ligand bound, different oligomers, etc.). Alternatively, perform quantification using specialized software like UniDec and Massign35 (Figure 1 and 3).

- Acquiring and Analyzing IM-MS

Note: IM-MS separates ions in the gas-phase based on their size (mass), shape and charge. Every feature resolved in m/z spectrum is associated with a drift time distribution. IM-MS measures the drift-time of an ion which can be used to calculate the collision cross section (CCS). Drift time values measured from IM-MS data acquired using a drift-tube can be linearly correlated to CCS values36. For travelling wave IM-MS (TWIMS) measurements, calculating CCS values requires a calibration curve obtained from protein standards with known CCS values37. Compact structures travel faster than extended or elongated structures due to reduced interactions with buffer gas in the mobility cell38. Therefore, IM-MS can be used to detect if the native folded structure has been retained in the gas phase39,40. This section outlines how to measure IM-MS and calculate the CCS of protein using TWIMS.

  - After optimizing instrument conditions for stable transmission (Step 2), reduce the collisional energy and sampling cone as low as possible whilst retaining good spectra quality.

  - Use the optimized wave velocity and wave height to acquire IM-MS (Step 2).

  - Measure the ion drift time with IM-MS at three different wave velocities (e.g., 550, 600 and 650 m/s) whilst maintaining the same wave height (e.g., 40 V).

  - To determine the protein ions CCS, measure protein calibrants under the same instrument conditions used for protein under investigation. Optimal drift-time calibration requires measurement of proteins with known CCS.

    - Select four calibrants, two with a mass above and two with a mass below that of the protein under investigation37. Most importantly, make sure that the wave height and wave velocity are the same as those recorded for the protein under investigation.

  - Calculate CCS manually41 or using a specialized software such as PULSAR33 and Amphitrite34 (Figure 5).

  - To check whether the protein is native-like in the gas-phase, compare experimental CCS to theoretical CCS obtained from high resolution structures. For HerA-NurA, calculate theoretical CCS using the Projection Approximation (PA) method used in MOBCAL42. Other methods include trajectory method (TM)42 and exact hard sphere scattering (EHS)43.

- In-Solution Disruption of Protein Complexes for Native MS and IM-MS Led Structure Determination

Note: Protein sub-complexes can in some cases be identified from the same solution as the intact complex. However, further structural information such as inter-subunit connectivity and complex assembly can be attained from disrupting protein interactions in solution, to form sub-complexes. This can be achieved in several ways such as the addition of organic solvent, increasing the ionic strength or manipulating the pH. To gain insight into the HerA-NurA complex subunit connectivity and complex assembly, sub-complexes were generated in solution by adding solvents which perturb subunit interactions.

  - Prepare the protein sample and buffer exchange into ammonium acetate as described in Step 1.

  - Add 10-40% of solvent in 10% increments. Solvents typically used are Methanol (MeOH), dimethyl sulfoxide (DMSO) or acetonitrile (ACN).

Note: This can be performed within a polypropylene microcentrifuge tube.

  - Incubate the mixture on ice for 1 h.

  - Acquire an IM-MS spectrum for each condition (Steps 2 and 3) (Figure 4).

  - Use the SUMMIT software44 to assign protein sub-complexes and generate protein interaction networks. Alternatively, manually generate a list of theoretical masses of the expected species.

  - To ensure subcomplexes are folded, calculate the experimental CCS values for the sub-complexes and compare to theoretical CCS as explained in Step 3 (Table 1, Figure 5 and Figure 6).

- Investigating Protein Complex Stability using Collision Induced Unfolding (CIU)

Note: CIU can be used probe the structural stability of proteins and their complexes upon ligand binding. Specialist software packages such as PULSAR33, Amphitrite34 and CIU suite9 can then be used to model the gas-phase unfolding of the protein under investigation with and without ligand. As an example, this section outlines procedure for monitoring gas-phase unfolding trajectories and investigating the stabilizing effect of DNA and ATP binding on the HerA-NurA complex.

  - Record IM-MS data whilst increasing the trap acceleration voltage from 10 V to 200 V in 2-10 V increments to progressively unfold the protein in the gas-phase.

Note: Recording smaller increments results in more data files to process, however this approach provides more resolved unfolding plot, which is important for analyzing the transition points between folded/unfolded species.

  - Analyze the data acquired using PULSAR33, Amphitrite34 or CIU suite9 and generate two-dimensional unfolding plots in units of CCS as a function of accelerating voltage (Step 3). For each charge state, this is created by stacking the intensity- normalized CCS distributions at each accelerating voltage (Figure 7 A-Bi).

  - Generate a theoretical unfolding plot using one of the software packages. The data will be fitted to an unfolding model. This makes it possible to quantify the collisional energy at which unfolding transitions occur and determine the stability of proteins with and without bound ligands33. An unfolding transition is when a species transitions from one state (based on their experimental CCS values) to another state with a larger CCS.

  - To quantitate the transitions, calculate the transitional midpoint(s) between states using algorithms and software such as PULSAR33. This is commonly reported as CV50, which is the collision (trap) voltage value at which 50% of a specific state is depleted.

  - Using the CV50 value, calculate the total internal energy of an ion using the center-of-mass collision energy (KECOM)45. KECOM is defined by the total internal energy available for the unfolding transition of an ion and is calculated from the kinetic energy and masses of the collision partners (protein ion and neutral gas) as described in equation (1)10:

KEcom (eV) = CV50 x ZMNMN+MION (equation 1).

Where Z is the ion charge, MN is the mass of the neutral gas and MION is the mass of the protein ion.

Note: This is because CIU of proteins is charge dependent46,47. It is recommended to perform the KECOM analysis to more than one charge state (Figure 7Aii).

- Modelling Procedures for Differential Molecular Dynamics Simulations used in Integrative MS

Note: Using models of protein subunits or complexes such as from crystal structures, differential MD simulations (protein complex with and without ligand) can be used to determine effects of for example ligand presence on protein structure and dynamics. This section details a workflow and tools needed for modelling procedures necessary to set-up differential molecular dynamics simulations.

  - Identify the subunits which compose the complex (Figure 8A, in Steps 2 and 3). Source existing models of subunits, e.g., crystal structures from the RCSB databank (https://www.rcsb.org). The UniProt entry of the protein will contain a list of known crystallographic/NMR structures (http://www.uniprot.org). If these are not available, the theoretical sequence can be input to BLAST to identify suitable templates for homology modelling (https://blast.ncbi.nlm.nih.gov/).

  - Assemble the complex in the correct topology (Figure 8A-ii). This can be done through various methods. The individual subunits can be fitted into available electron microscopy maps found on the EMDB to assemble the intact complex (https://www.ebi.ac.uk/pdbe/emdb/). A tutorial for fitting PDBs into EM maps using Molecular Dynamics Flexible Fitting (MDFF) can be found here: http://www.ks.uiuc.edu/Training/Tutorials/science/mdff/tutorial_mdff-html/.

  - Identify missing regions of the complex (Figure 8A-iii). Perform multiple sequence alignment (MSA) between the PDB and theoretical sequence to identify residues which may be unfitted into crystal structures, or any mutations inherited from crystallographic experiments. MSA can be performed using the webservers such as T-Coffee (http://tcoffee.crg.cat/apps/tcoffee/do:regular).

  - Regenerate missing residues via homology modelling (Figure 8A-iv). Missing residues of the protein complex can be built using the MODELLER program (https://salilab.org/modeller/). MODELLER can output an ensemble of n models in different regenerated configurations. Good models can be identified based on their Discrete Optimised Protein Energy (DOPE) score. A comprehensive tutorial is provided on the software website (https://salilab.org/modeller/tutorial/).

  - Perform differential molecular dynamics (MD) simulations of the protein complex (Figure 8B) to identify regions of proteins which respond to a particular environmental change, e.g., presence of a ligand. In such simulations, behavioural parameters from Simulation A (protein only) which acts as a reference, is subtracted from Simulation B (protein+ligand). The differential root mean square fluctuation (RMSF) calculated between Simulations A and B can inform on regions of the protein which increase or decrease in flexibility in a ligand dependent manner.

    - Perform MD simulations and downstream analysis using GROMACS (http://www.gromacs.org). A tutorial can be found at: http://www.bevanlab.biochem.vt.edu/Pages/Personal/justin/gmx-tutorials/Lysozyme/index.html. To eliminate model bias, the structure of the ligand-bound complex should be generated first. The protein is then copied from this without the ligand, to yield a protein model identical to the ligand-bound complex.

Representative results:

Native MS results revealed the oligomeric state, composition and topology of the HerA-NurA complex (Figure 1). As non-covalent interactions are preserved in the gas-phase, native MS of ATP-γ-S and ADP titrations experiments determined the pairwise nucleotide binding to HerA-NurA (Figure 2) and that increasing the ATP-γ-S concentration increases the relative intensity of hexameric HerA (Figure 3). Structural information regarding subunit interactions were obtained from in-solution disruption followed by native MS and were in agreement with and theoretical masses (Figure 4 and Table 1).

The experimental CCS values of proteins and their complexes was derived from IM-MS experiments (Figure 5). These values are rotationally averaged gas-phase cross-sectional calculations of the molecular shape, and describe the dimensional state of the protein. CCS values are compared to theoretical measurements from x-ray crystallography and a good agreement infers that the native shape in retained in the gas-phase (Table 1). This validates using CCS values for building low-resolution models of the protein assembly48.

Experimental CCSs can be calculated for each charge state ion. A native-like protein conformer may give rise to charge state ions with similar CCS values. However, higher charge state ions increased coulombic repulsions which may lead to protein gas-phase unfolding and larger CCS values compared to the theoretical CCSs. The CCS value of the lowest charge state ions are therefore usually used49. For HerA-NurA, in-solution disruption experiments on HerA and HerA-NurA with and without DNA prompted the generation of an assembly pathway starting with monomers then forming the entire hexameric HerA (HerA6)-dimeric NurA (NurA2) complex with DNA (Figure 6).

Differences in the CIU unfolding plots between the apo (ligand-free) and ligand bound define the change in complex stability upon ligand binding. A higher CV50 or KECOM value implies a more stabilized ion in the gas-phase. CIU and KECOM analysis revealed DNA-bound HerA-NurA is more stable than the DNA-free complex (Figure 7Aii). From CIU-MS analysis in the respective ATP-binding states, the four-ATP-γ-S bound state reduced complex stability in the gas-phase and the six -ATP-γ-S bound state where are all sites are occupied was the most stable (Figure 7Bii). Native MS can reveal the discrete nucleotide binding states of HerA; however, it cannot distinguish which HerA subunits are binding ATP and where this binding takes place. This information can be derived from explicit solvent MD simulations on the hexameric HerA and the HerA-NurA following the summarised Workflow (Figure 8).

Figure 1. Interrogating the oligomeric state, composition and topology of the HerA-NurA non-covalent complex. (A) Mass spectra of HerA, HerA-NurA and HerA-NurA in the presence of DNA (15.4 kDa 25 bp double- stranded DNA). The HerA sub-complex exists as both a hexamer and a heptamer. NurA dimer binds and to a HerA hexamer imposing oligomeric conversion. The DNA binds the formed HerA– NurA complex (Results adapted from Z. Ahdash et al., 201722). (B) Relative intensities of the identified species are calculated using UniDec32.

Figure 2. Native ESI-MS reveals the mechanism of nucleotide binding to HerA-NurA. Mass spectra (A) HerA-NurA and (B) HerA-NurA-DNA with increasing concentrations of (i) ATP-γ-S and (ii) ADP. Measured masses are compared to theoretical masses and the amount of ATP-γ-S or ADP bound are determined. Measured masses and number of bound nucleotides are shown on the spectra. ATP-γ-S and ADP titrations experiments determined the pairwise nucleotide binding to HerA-NurA alone and when in complex with DNA indicating a cyclical reaction mechanism (Results adapted from Z. Ahdash et al., 201722).

Figure 3. Measuring the effect of increasing ATP-γ-S concentrations on HerA oligomeric state. (A) Mass spectra of HerA at increasing concentrations of ATP-γ-S. (B) Graph showing the relative intensities of different species from native MS calculated using UniDec deconvolution software32. As the ATP-γ-S concentration increases, the relative intensity of hexameric HerA also increases. Number of ATP-γ-S molecules bound is shown on the spectra (Results adapted from Z.Ahdash et al., 201722).

Figure 4. Mass spectra and sub-complex dissociation products of (A) HerA and (B) HerA-NurA (i) alone and in the presence of (ii) DNA following of in-solution disruption. In-solution disruption experiments were performed using 10-40% of Acetonitrile, Methanol (MeOH) or dimethyl sulfoxide (DMSO) and resulted in the formation of various subcomplexes.

Figure 5. Ion mobility arrival time distributions shown on a CCS axis for complexes and generated sub-complexes. Icon for each sub-complex correlate with those annotated on the spectra in Figure 4. Experimental and calculated masses and CCS values of Sub-complexes are listed in Table 1 all of which showed an agreement between experimental and calculated values (after considering the typical uncertainty in the resolution of travelling wave ion mobility mass spectrometry of ±5–8%37.

Figure 6. Assembly pathway of the HerA-NurA complex generated from in-solution disruption native MS and IM-MS. Colored circles indicate conditions where each sub-complex was observed: native (green, prior to disruption), Methanol (yellow), DMSO (purple) or Acetonitrile (red).

Figure 7. Investigating the stabilizing effect of (A) DNA on HerA-NurA and (B) ATP binding to HerA-NurA-DNA. (i) Gas-phase CIU-MS plots and (ii) center-of-mass collision energies (KEcom) calculation show that that the presence of dsDNA stabilizes the HerA-NurA complex and that the six ATP-γ-S bound state is the most stable. Stabilization for different charge states is shown. Plots were generated using PULSAR 33. Results from (A) adapted from Z. Ahdash et al., 201722.

Figure 8. Workflow for the modelling procedures for differential molecular dynamics simulations. (A) Generating the complex under investigation by building the topology from existing subunits and rebuilding missing residues. (B) Workflow for running molecular dynamics simulations on the protein complex with and without ligand. Molecular dynamics simulations are ran for the protein only which acts as a reference and which is subtracted from simulation of protein plus ligand. This is followed by calculating the differential root mean square fluctuation (RMSF) between the simulations and determining the effect of ligand binding.

Table 1. Experimental and calculated masses and CCS values of HerA-NurA and its sub-complexes generated form in-solution disruption studies.

Discussion:

MS is playing an increasingly important role in characterizing the stoichiometry, interactions and subunit architecture of protein complexes. IM-MS data can be used to define topological arrangements of subunits within multicomponent complexes. Compared to other existing structural biology methods, MS has several advantages. Native MS is a rapid and highly sensitive technique and can be used to probe heterogeneous protein samples. When coupled with in-solution disruption experiments, dissociation pathways of protein assemblies can be monitored. Together with crystal structures or homology models, the information offered by structural MS offers a tool for investigating protein-ligand interactions and provide near-native models and assembly pathways11.

Here, we describe the necessary experimental procedures for analyzing the stoichiometry and composition of protein-ligand interactions, with one or more ligands, using integrative MS. This includes MS sample preparation, data acquisition, data analysis, and the integration of MS data using computational tools. To do this, we used the DNA-resection HerA-NurA hetero-oligomeric protein complex, bound to three ligands (DNA, ATP, and ADP), as our model system. The protocol shows the use of the currently available software to aid data analysis and presentation.

Acquiring high quality spectra is important for ligand binding analysis, therefore, careful sample preparation steps are critical, including protein purification, ligand titration, and buffer exchange. One limitation of nESI native MS when studying ligand binding is non-specific binding. Non-specific binding occurs during droplet desolvation throughout the electrospray process. This increases the ligand concentrations and therefore alters the protein/Ligand ratio29. The binding of nucleotides results in a relatively small mass difference between apo and nucleotide-bound protein which does not alter the ionization efficiency50,51.

We used the Synapt G2-Si MS system for our work, but the protocols are applicable for different investigations of other protein-ligand complexes using other commercially available nano-electrospray mass spectrometers. Integrative structural MS is increasingly playing an important role in addressing biological problems of greater complexity. The workflow and techniques described here are well-suited for understanding the structural consequences and building mechanisms of protein complex and protein-ligand formation which are otherwise difficult to study using conventional structural techniques.

Disclosures:

No conflicts of interest declared.

Acknowledgments:

We would like to thank Karl-Peter Hopfner and Robert Thomas Byrne for kindly providing HerA and HerA-NurA protein samples and for their help with experimental design. We also thank Dr. Eamonn Reading for his review of the manuscript. We gratefully acknowledge our funding bodies: Wellcome Trust [109854/Z/15/Z] and Royal Society [RG150216 to A.P.].

References:

1	Wilm, M. & Mann, M. Analytical Properties of the Nanoelectrospray Ion Source. Analytical Chemistry. 68 (1), 1-8, (1996).

2	Fenn, J.B., Mann, M., Meng, C.K., Wong, S.F. & Whitehouse, C.M. Electrospray ionization for mass spectrometry of large biomolecule. Science. 246 (4926), 64-71, (1989).

3	Hernandez, H. & Robinson, C. V. Determining the stoichiometry and interactions of macromolecular assemblies from mass spectrometry. Nature Protocols. 2 (3), 715-726, (2007).

4	Heck, A. J. & Van Den Heuvel, R. H. Investigation of intact protein complexes by mass spectrometry. Mass Spectrom Review. 23 (5), 368-389, (2004).

5	Boeri Erba, E. & Petosa, C. The emerging role of native mass spectrometry in characterizing the structure and dynamics of macromolecular complexes. Protein Science. 24 (8), 1176-1192, (2015).

6	Arthur Laganowsky, E. R., Jonathan T S Hopper & Carol V Robinson. Mass spectrometry of intact membrane protein complexes. Nature. 8 (4), 639–651, (2013).

7	Benesch, J. L. Collisional activation of protein complexes: picking up the pieces. Journal of the American Society for Mass Spectrometry. 20 (3), 341-348, (2009).

8	Tian, Y., Han, L., Buckner, A. C. & Ruotolo, B. T. Collision Induced Unfolding of Intact Antibodies: Rapid Characterization of Disulfide Bonding Patterns, Glycosylation, and Structures. Analytical Chemistry. 87 (22), 11509-11515, (2015).

9	Joseph D. Eschweiler, J. N. R.-G., Yuwei Tian, and Brandon T. Ruotolo. CIUSuite: A Quantitative Analysis Package for Collision Induced Unfolding Measurements of Gas-Phase Protein Ions. Analytical Chemistry. 87 (22), 11516–11522, (2015).

10	Hopper, J. T. & Oldham, N. J. Collision induced unfolding of protein ions in the gas phase studied by ion mobility-mass spectrometry: the effect of ligand binding on conformational stability. Journal of the American Society for Mass Spectrometry. 20 (10), 1851-1858, (2009).

11	Politis, A., et al. Topological models of heteromeric protein assemblies from mass spectrometry: application to the yeast eIF3:eIF5 complex. Chemical Biology. 22 (1), 117-128, (2015).

12	Morgner, N., et al. Hsp70 forms antiparallel dimers stabilized by post-translational modifications to position clients for transfer to Hsp90. Cell Reports. 11 (5), 759-769, (2015).

13	Lai, Y.T., et al. Structure of a designed protein cage that self-assembles into a highly porous cube. Nature Chemistry. 6 (12), 1065-1071, (2014).

14	Levy, E. D., Boeri Erba, E., Robinson, C. V. & Teichmann, S. A. Assembly reflects evolution of protein complexes. Nature. 453 (7199), 1262-1265, (2008).

15	Sharon, M., et al. Symmetrical modularity of the COP9 signalosome complex suggests its multifunctionality. Structure. 17 (1), 31-40, (2009).

16	Jurneczko, E. & Barran, P. E. How useful is ion mobility mass spectrometry for structural biology? The relationship between protein crystal structures and their collision cross sections in the gas phase. Analyst. 136 (1), 20-28, (2011).

17	Politis, A. & Schmidt, C. Structural characterisation of medically relevant protein assemblies by integrating mass spectrometry with computational modelling. Journal of Proteomics. 175, 34-41, (2017).

18	Zhou, M., et al. Mass spectrometry of intact V-type ATPases reveals bound lipids and the effects of nucleotide binding. Science. 334 (6054), 380-385, (2011).

19	Barrera, N.P., Booth, P.J. & Robinson, C.V. Micelles Protect Membrane Complexes from Solution to Vacuum. Science. 321 (5886), 243-246, (2008).

20	Konijnenberg A., et al. Global structural changes of an ion channel during its gating are followed by ion mobility mass spectrometry. Proceedings of the National Academy of Sciences of the United States of America. 111 (48), 17170-17175, (2014).

21	Schmidt, C., Beilsten-Edmands, V., Mohammed, S. & Robinson, C. V. Acetylation and phosphorylation control both local and global stability of the chloroplast F1 ATP synthase. Scientific Reports. 7, 44068, (2017).

22	Ahdash, Z. et al. Mechanistic insight into the assembly of the HerA-NurA helicase-nuclease DNA end resection complex. Nucleic Acids Research. 45 (20), 12025-12038, (2017).

23	Pagel, K., Natan, E., Hall, Z., Fersht, A. R. & Robinson, C. V. Intrinsically disordered p53 and its complexes populate compact conformations in the gas phase. Angewandte Chemie. 52 (1), 361-365, (2013).

24	Afonso, J.P., et al. Insights into the structure and assembly of the Bacillus subtilis clamp-loader complex and its interaction with the replicative helicase. Nucleic Acids Research. 41 (9), 5115-5126, (2013).

25	van Duijn, E. Current limitations in native mass spectrometry based structural biology. Journal of the American Society for Mass Spectrometry. 21 (6), 971-978, (2010).

26	Susa, A. C., Xia, Z. & Williams, E. R. Native Mass Spectrometry from Common Buffers with Salts That Mimic the Extracellular Environment. Angewandte Chemie. 56 (27), 7912-7915, (2017).

27	Breukera, K., and McLafferty, F.W. Stepwise evolution of protein native structure with electrospray into the gas phase, 10 to 10 s. Proceedings of the National Academy of Sciences of the United States of America. 105 (47), 18145–18152, (2008).

28	Robinson, C.V., et al. Probing the Nature of Noncovalent Interactions by Mass Spectrometry. A Study of Protein-CoA Ligand Binding and Assembly. Journal of the American Chemical Society. 118 (36), 8646–8653, (1996).

29	Shimon, L., Sharon, M. & Horovitz, A. A method for removing effects of nonspecific binding on the distribution of binding stoichiometries: application to mass spectroscopy data. Biophysical Journal. 99 (5), 1645-1649, (2010).

30	Dyachenko, A., Gruber, R., Shimon, L., Horovitz A., Sharon, M. Allosteric mechanisms can be distinguished using structural mass spectrometry. Proceedings of the National Academy of Sciences of the United States of America. 110 (18), 7235–7239, (2013).

31	Kirshenbaum, N., Michaelevski, I. & Sharon, M. Analyzing large protein complexes by structural mass spectrometry. Journal of Visualized Experiments. 19 (40), (2010).

32	Marty, M.T., et al. Bayesian deconvolution of mass and ion mobility spectra: from binary interactions to polydisperse ensembles. Analytical Chemistry. 87 (8), 4370-4376, (2015).

33	Allison, T.M., et al. Quantifying the stabilizing effects of protein-ligand interactions in the gas phase. Nature Communications. 6, 8551-8561, (2015).

34	Sivalingam, G.N., Yan,J., Sahota, H. & Thalassinos, K. Amphitrite: A program for processing travelling wave ion mobility mass spectrometry data. International Journal of Mass Spectrometry. 345-347, 54-62, (2013).

35	Morgner, N. & Robinson, C. V. Massign: an assignment strategy for maximizing information from the mass spectra of heterogeneous protein assemblies. Analytical Chemistry. 84 (6), 2939-2948, (2012).

36	Matthew, F., et al. Collision Cross Sections of Proteins and Their Complexes: A Calibration Framework and Database for Gas-Phase Structural Biology. Analytical Chemistry. 82, 9557–9565, (2010).

37	Ruotolo, B. T., Benesch, J. L., Sandercock, A. M., Hyung, S. J. & Robinson, C. V. Ion mobility-mass spectrometry analysis of large protein complexes. Nature Protocols. 3 (7), 1139-1152, (2008).

38	Lanucara, F., Holman, S. W., Gray, C. J. & Eyers, C. E. The power of ion mobility-mass spectrometry for structural characterization and the study of conformational dynamics. Nature Chemistry. 6 (4), 281-294, (2014).

39	Wyttenbach, T. & Bowers, M. T. Structural stability from solution to the gas phase: native solution structure of ubiquitin survives analysis in a solvent-free ion mobility-mass spectrometry environment. Journal of Physical Chemistry B. 115 (42), 12266-12275, (2011).

40	Marcoux, J. & Robinson, C. V. Twenty years of gas phase structural biology. Structure. 21 (9), 1541-1550, (2013).

41	Michaelevski, I., Kirshenbaum, N. & Sharon, M. T-wave ion mobility-mass spectrometry: basic experimental procedures for protein complex analysis. Journal of Visualized Experiments. 41, (2010).

42	Mesleh, M. F., Hunter, J.M., Shvartsburg, A.A., Schatz, G.C. & Jarrold, M.F. Structural Information from Ion Mobility Measurements: Effects of the Long-Range. Potential. Journal of Physical Chemistry A. 100 16082-16086, (1996).

43	Shvartsburg, A.A. & Jarrold, M. F. An exact hard-spheres scattering model for the mobilities of polyatomic ions. Chemical Physics Letters. 261 86-91, (1996).

44	Taverner, T., Hernández, H., Sharon, M., Ruotolo, B.T., Matak-Vinković, D., Devos, D., Russell, R.B & Robinson, C.V. Subunit Architecture of Intact Protein Complexes from Mass Spectrometry and Homology Modeling. Accounts of Chemical Research. 41 (5), 617–627, (2008).

45	Wysocki, V. H., Joyce, K. E., Jones, C. M. & Beardsley, R. L. Surface-induced dissociation of small molecules, peptides, and non-covalent protein complexes. Journal of the American Society for Mass Spectrometry. 19 (2), 190-208, (2008).

46	Hall, Z., Politis, A., Bush, M. F., Smith, L. J. & Robinson, C. V. Charge-state dependent compaction and dissociation of protein complexes: insights from ion mobility and molecular dynamics. Journal of the American Chemical Society. 134 (7), 3429-3438, (2012).

47	Hyung, S.J., Robinsons, C. V. & Ruotolo, B.T. Gas-Phase Unfolding and Disassembly Reveals Stability Differences in Ligand-Bound Multiprotein Complexes. Chemistry & Biology. 16, 382–390, (2009).

48	Hall, Z., Politis, A. & Robinson, C. V. Structural modeling of heteromeric protein complexes from disassembly pathways and ion mobility-mass spectrometry. Structure. 20 (9), 1596-1609, (2012).

49	Woods, L. A., Radford, S. E. & Ashcroft, A. E. Advances in ion mobility spectrometry-mass spectrometry reveal key insights into amyloid assembly. Biochimica et Biophysica Acta. 1834 (6), 1257-1268, (2013).

50	Daubenfeld, T., Bouin, A. P. & van der Rest, G. A deconvolution method for the separation of specific versus nonspecific interactions in noncovalent protein-ligand complexes analyzed by ESI-FT-ICR mass spectrometry. J Am Soc Mass Spectrom. 17 (9), 1239-1248, (2006).

51	Loo, J. A. Studying noncovalent protein complexes by electrospray ionization mass spectrometry. Mass Spectrometry Reviews. 16 1–23, (1997).