- Submission ID #: 58168

- Editor Name: William Hoston

- Videographer name:

- Film Date:

- Link: http://www.jove.com/files_upload.php?src=17741288

- Authors and Affiliations: Robert P. Thoerner, Jonathan D. King, Marnie M. Saunders

Department of Biomedical Engineering, University of Akron, USA

- Title: Application of Design Aspects in Uniaxial Loading Machine Development

- Corresponding Author: Marnie Saunders (mms129@uakron.edu)

- Co-authors: Robby Thoemer, Jonathan King

A.  Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? N

B.   Software Usage: Does your protocol include detailed, step-by-step, descriptions of software usage? N

C.  Which steps of your protocol will viewers benefit most from having filmed?

D.  What is the single most difficult aspect of this procedure and what do you do to ensure success?

E.  Will the filming need to take place in multiple locations? N

1. Introduction (Experimental Goal and Author Interviews) – As the beginning of your video, the introduction should clearly present the goal of your method to the viewer and its significance.  Other information can be provided according to the various statements below, but the total introduction should not exceed 150 words.

A. Required Interview Statements: (Said by you on camera. Don’t forget to smile!)

  - Marnie Saunders: This method can help researchers develop loading machines that perform motions that otherwise may not be possible in their labs.

  - Marnie Saunders: Reproducibility and accuracy are important and we demonstrate reproducible assembly and accurate alignment techniques.

B. Optional Interview Statements: N/A

- C. Introduction of Demonstrator: (Said by you on camera. Don’t forget to smile!)

  - Marnie Saunders: Demonstrating the procedure will be Robby Thoerner, a recent bachelor’s degree recipient from my laboratory.

    - Interview style: Author saying the above

    - The named technician, post doc, student looks up from workbench or desk or microscope and acknowledges the camera.

D. Ethics title card: N/A

- Protocol: (read by voice talent at JoVE)

- The Device and Its Component Parts

  - The goal is to assemble the testing device. [1-WIDE] Here is an example of an assembled device, which has both a front and a rear actuator, load cells, and a frame to support them. [2-CU]

    - Talent at bench with assembled device.

    - Detail of the assembled device to reveal the named parts

  - Gather the parts to construct the device. [1-MED] Use two programmable actuators, each with 30 millimeters travel, 60 milliliters when working together. [2-CU-TXT] Daisy-chain the actuators to sync them for equal extension and retraction. [3-CU]

    - Talent with parts arranged for easy display (actuators, load cells, carriages, rail)

    - Actuators (at least one).  [TEXT: See text protocol for other actuator features]

    - Daisy-chained actuators

  - Next, prepare two load cells suitable for confined spaces. [1-CU-TXT] To support the actuators, prepare one carriage for each of them, along with a rail. [2-CU]

    - Load cells [TEXT: Max. force capacity: 44.5 N]

    - The two carriages and a rail, not attached

- Frame Construction: Base and Side Plate Assembly and Attaching the Rail/Carriage Assembly

  - Create the base from aluminum stock. [1-WIDE] Cut a piece to the appropriate size and use a mill to machine the plate according to specifications. [2-CU-TXT]

    - Talent at bench with base

    - The completed base [TEXT: Dimensions: 64 cm x 15 cm x 1.3 cm; See specifications in supplemental files]

  - Machine two side plates according to specifications. [1-MED-TXT]  The two should mirror one another, and each should have a slot and drilled holes to accommodate the rail. [2-CU] Drill and tap the bottom faces of the side plates. [3-CU]

    - Talent getting side plates [TEXT: See supplemental files.]

    - Detail of the two side plates, side-by side, as they would be on the base.

    - The bottom face of a side plate to demonstrate the holes there

  - Mount the side plates in the track on the base. [1-MED] Be sure the mount points for the rail are near each other. [2-CU] Fasten the side plates to the base plate from underneath. [3-MED]

    - Talent mounting the side plates on the base

    - Detail of the side plates to show their proper orientation

    - Talent fastening the side plates

  - Next, fasten the rail to the track using its clearance holes and the drilled holes in each side plate. [1-CU] When done, the rail connects the two side plates. [2-CU] (Authors: It is not clear when you install the carriages. Are they already on the rail? Are they added immediately after the rail?)

    - The rail being put into position and fastened

    - The rail in position connecting the two side plates

- Rear Mount Attachment of the Actuators

  - Move on to work with mounts for the actuators. [1-WIDE] Use ‘L’-shaped stock to machine a rear mount attachment and a bar to attach to the bottom of the mount. [2-CU-TXT]  Screw the bar into the bottom of the mount. [3-CU]

    - Talent at bench, getting rear mount attachments and bar

    - The rear mount and the bar [TEXT: See supplemental files for details]

    - The rear mount as the bar is being attached. Have several seconds of the completed assembly

  - At this point, get the selected actuator for the experiment. [1-MED] Attach the rear mount to the body of the actuator using the hole pattern in the actuator. [2-CU]

    - Talent getting actuator and putting it into position

    - The actuator as the rear mount is being attached to it

  - Slot the base of the mount to the frame to attach the rear actuator. Use two screws to secure it. The series of holes in the side plate allow the mount position to be adjusted. [1-CU] Duplicate the steps to mount an actuator on each of the side plates. [2-CU]

    - The rear mount being put into position on the frame. After a pause, screws being used to secure it

    - Demonstrate that an actuator has been mounted on each of the side plates

- Front Mount Attachment of the Actuators

  - Get the parts for supporting the front of the actuator. [1-WIDE] One part is an ‘L’-shaped piece with a hole for a tapered connector. On its side is a track to accommodate a plate. [2-CU] The plate for this track also has a track to accommodate fixtures. [3-CU]

    - Talent at bench with parts for the front mount

    - The ‘L’-shaped piece demonstrating the hole for the connector. If necessary, after a pause, the piece being rotated to show the track. The second sentence might also be a separate shot

    - The plate being put into the track on the ‘L’-shaped piece of the previous shot. Ideally the track on the plate would also be visible.

  - Machine an aluminum, cylindrical part to connect the load cell and actuator. [1-CU] Machine an aluminum, tapered, cylindrical part to connect the load cell to the fixture and carriage. [2-CU] Pass the cylinder into the hole of the front actuator mount and use a set screw to anchor the cylinder end. [3-CU]

    - The cylindrical, aluminum connector

    - The tapered, cylindrical connector. Possibly put this piece next to the piece of the previous shot to allow comparison

    - The tapered connector being put in the front actuator mount and screwed into place

  - Now, work with the fixtures for the front mounts. [1-MED] For vertical adjustments, machine a central vertical slot in the fixture holder. [2-CU] Attach the actuator front mount to the rectangular plate with the vertically aligned holes in its center. [3-CU]

    - Talent at bench getting fixture parts

    - Detail of the fixture holder

    - The front mount as it is attached to the place

  - Here is the completed assembly for the front mount on one side. [1-CU] Duplicate the steps to mount both the left and right actuators of the device. [2-CU]

    - The completed front mount assembly

    - The device to demonstrate that both sides have been fitted in the same way

- Results: Load Displacement Tests with the Uniaxial Loading Machine

  - (Voice talent: Pronounce “2-0” and “two-oh”) These are representative data sets from load testing knotted, 25.4 millimeter 2-0 sutures until failure. [1-LM] The grey dashed curve represents results from the device described in this protocol using a loading rate of 0.61 millimeters per second. [2-LM] The black curve represents results from a fixed-end device using a doubled loading rate or 1.22 millimeters per second. [3-LM] In all tests, failure occurred at the knot and measurements showed no statistical difference between the devices. [4-LM]

    - LAB MEDIA: Figure8.tiff

    - LAB MEDIA: Figure8.tiff (Video editor: Please call attention to the grey curve during this sentence)

    - LAB MEDIA: Figure8.tiff (Video editor: Please call attention to the black curve during this sentence)

    - LAB MEDIA: Figure8.tiff

- Conclusion (said by authors on camera)

  - Marnie Saunders: While attempting this procedure, it is important to always keep alignment in mind. During all aspects of the design process put measures into place to ensure that the device gives accurate and reproducible results.

  - Marnie Saunders: If you understand the basic principles of design for this device, they can be applied to develop any loading device and properly conduct mechanical tests.

  - Marnie Saunders: For safety reasons, only experienced machinists should machine the components of this or similar devices.

- Provided Media

Authors, Please list all images, movie files, or 3-D rendered animations that can be included in the video per editor’s request.  The step in the script/video where these images will be inserted should be specified.   For example:

- 6.2 –  0123_PIname_Figure1.tif -  dual color imaging of tumor angiogenesis at 40X

6.2 –  0123_PIname_Figure2.tif -  dual color imaging of tumor angiogenesis at 100X

Formats:  For static images we prefer .tiff, .eps, Illustrator, Powerpoint or Photoshop files at dimensions of at least 720X480 pixels and 300 dpi.  The higher resolution, the better.  Likewise any exported movie files should have at minimum these dimensions and be rendered to .mov, .mp4, or .avi files.

- Insert your media filenames here.

- General Preparation

- It’s critical for a smooth and organized shoot that all reagents are accounted for, in advance.

Any overnight or long incubation steps should be recognized and specimens/samples be prepared in advance so that prior steps can be recorded and shooting can continue with pre-prepared specimens/samples.

- All tubes/flasks should be pre-labeled neatly before we arrive.

Ex. Luciferase assay done in 96 well plates should be labeled with negative/positive control wells and experimental samples are labeled accordingly.

You will receive more detailed preparation instructions are included in the email accompanying the finalized script.