- Quick 19F{1H} spectrum to set up initial parameters

- Measure the 90 degree pulse width (P1, sample dependent)

- Set up parameter for T1 measurement experiment (SW, O1, P1, VDLIST[t1delay])

- Running T1 measurement experiment using an inversion-recovery sequence (t1ir)

- T1 experiment data processing

  - check TD(F1): number of points in VDLIST (in this case 16)

  - rser 16 - call in final spectrum with longest τ delay

  - ef – apply window function and FT

  - phase correction to the signal

  - save phase parameters to 2D dataset

  - xf2 – apply FT in F2 dimension only

  - go to Dynamics tab (within analysis) and select t1t2 mode

  - load full FID (#16 in this case) and it should be phased (from previous steps)

  - select Peaks/ranges and choose Manual integration. Integrate peak by holding left mouse button on one side of desired peak and dragging to the other before releasing left mouse button.

  - Export Regions to relaxation module and .ret.

  - Select Relaxation button. If options dialog appears, select default options with OK

  - To fit all relaxation data and calculate T1 value