# Submission ID #: 58986

# Scriptwriter Name: Bridget Colvin

Project Page Link: http://www.jove.com/files_upload.php?src=17990778

Title: High-Throughput and Comprehensive Drug Surveillance Using Multisegment Injection-Capillary Electrophoresis Mass Spectrometry

Authors and Affiliations: Meera Shanmuganathan1, Sabrina Macklai1, Clara Barrenas Cárdenas1, Zachary Kroezen1, Marcus Kim2, William Zizek3, Howard Lee3, and Philip Britz-McKibbin1

1Department of Chemistry and Chemical Biology, McMaster University

2Agilent Technologies Inc.

3Seroclinix Corporation

# Corresponding Author:

# Philip Britz-McKibbin

# britz@mcmaster.ca

Email addresses for Co-authors:

shanmm2@mcmaster.ca

macklais@mcmaster.ca

c.barrena1@usp.ceu.es

marcus.kim@agilent.com

kroezezp@mcmaster.ca

hlee@seroclinix.com

# bzizek@seroclinix.com

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy? N

2. Does your protocol include software usage? Y

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

Authors: please upload all screen captured files to your project page.

3. Which steps from the protocol section below are the most important for viewers to see?

3.1.-3.3., 3.6.-3.8.

4. What is the single most difficult aspect of this procedure and what do you do to ensure success?

3.2., 3.6.

5. Will the filming need to take place in multiple locations? N

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

- REQUIRED Interview Statements (Said by you on camera): All interview statements may be edited for length and clarity.

  - Philip Britz-McKibbin: v [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Philip Britz-McKibbin: The main advantage of our method referred to as multisegment injection-capillary electrophoresiss-mass spectrometry or MSI-CE-MS is that it allows for rapid screening of virtually an unlimited panel of drugs and their metabolites with high accuracy, minimal sample workup [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# OPTIONAL Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Meera Shanmuganathan: There is an urgent need for a high-throughput and affordable screening method to improve patient outcomes by confirming adherence and optimizing drug dosage regimes while revealing illicit drug usage [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Meera Shanmuganathan: These multiplexed separations platform can be applied to a variety of other research areas related to targeted or nontargeted metabolite profiling in complex biological samples [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Introduction of Demonstrator (Said by you on camera):

  - Zachary Kroezen: I am an undergraduate Coop student in our Chemistry program at McMaster University and I will be demonstrating the procedure [1][2].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

    - The named technician, post doc, student looks up from workbench or desk or microscope and acknowledges the camera

Ethics title card: (for human subjects or animal work, does not count toward word length total)

  - Procedures involving human subjects have been approved by the Hamilton Integrated Research Ethics Board.

# Section - Protocol

- Urine Sample Preparation

  - Begin by thawing deidentified morning urine samples from a representative cohort of ten clinically depressed patients with a known prescription drug history from -80-degree Celsius storage on ice [1-TXT] and vortexing the thawed samples for 30 seconds [2].

    - WIDE: Talent placing samples into rack onto ice TEXT: Avoid multiple free-thaw cycles/delayed RT storage

    - MED: Talent vortexing sample(s)

  - Next, sediment the samples by centrifugation [1-TXT] and combine three 10 microliter aliquots of each sample with 10 microliters of mixture of deuterated internal standards, 5 microliters of a 4-fluoro-L-phenylalanine and 3-chloro-L-tyrosine solution, and 25 microliters of deionized water [2].

    - MED: Talent adding tube(s) to centrifuge TEXT: 1 min, 14,000 x g

    - MED: Talent adding sample to one of three tubes, with d-IS and F-Phe/Cl-Tyr containers visible in frame

  - Then vortex the mixtures for 1 minute [1] and transfer 20 microliters of each sample into individual polypropylene vials [2].

    - CU: Mixture being vortexed

    - MED: Talent adding sample to vial

- Capillary Electrophoresis-Time-of-Flight-Mass Spectrometer (CE-TOF-MS) System Setup

  - To set the uncoated fused-silica capillary conditioning parameters, cut a 135-centimeter long 50-micrometer-inner diameter and 360-micrometer outer diameter capillary from a spool [1] and use a capillary window maker to remove about 7 millimeters of polyimide coating from both ends of each capillary [2].

    - WIDE: Talent cutting capillary from a large spool (Editor: The authors included 3 subshots below. However, with the way the VO is currently written, I’m unsure as to when these should be shown. The authors may need to provide additional clarity)

# 3.1.1A. [Added Shot]: Talent cutting about 2.5 cm of capillary from both ends using a diamond cutter

# 3.1.1B. [Added Shot]: Talent cleaning capillary with methanol soaked wipe and spraying CE-sprayer with compressed air

# 3.1.1C. [Added Shot]: Talent installing one end of capillary in CE sprayer

    - CU: Coating be removed from one end both ends of capillary

# 3.1.2A. [Added Shot]: Talent wiping both uncoated ends with methanol soaked wipe.

  - Next, install the capillary inlets into the capillary electrophoresis, or CE (C-E), cartridge by looping each capillary for two 360-degree turns [1], taking care to protrude about 2 millimeters of the capillary outlets out of the CE sprayer needle [2].

    - MED: Talent installing inlet/looping capillary

    - CU: Shot of outlet protruding 2 mm out of CE sprayer needle

  - When all of the capillaries have been installed, carefully load the CE cartridge into the CE system [1] and store the CE sprayer out of the ion source [2].

    - MED: Talent loading cartridge onto CE system

    - MED: Talent CE moving sprayer out of ion source Talent storing CE-sprayer with a wipe beside ion source

  - Place the liquid chromatography, or LC (L-C), sprayer into the ion source and select Flush in the system software to setup the capillary conditioning [1] [2].

    - MED: Talent moving LC sprayer into ion source (Editor: I’m unsure if the VO needs to be adjusted because the 3.4.2 was removed)

    - MED-over the shoulder: Talent selecting Flush function, with monitor visible in frame

  - Set the duration of methanol, 1-molar sodium hydroxide, deionized water, and background electrolyte flush to 30 minutes per flush [1] and place the isocratic pump on standby during the capillary conditioning period [2].

    - SCREEN: To be provided by Authors: Flush duration(s) being set

  - After the last flush, clean the capillary electrophoresis-mass spectrometer, or CE-MS (C-E-M-S), sprayer [1] and the CE electrode with methanol-soaked wipes [3.6.2A] and clean the CE-MS interface with a 1:1 isopropanol-water solution to remove any residual salt deposits [3].

    - MED: Talent wiping sprayer

    - MED: Talent wiping removing electrode (Editor: This shot may not need to be shown. If it is shown, it could be shown before 3.6.2A)

# 3.6.2A. [Added Shot]: Talent cleaning and wiping electrode (Editor: This shot should be shown instead of 3.6.2)

    - MED: Talent cleaning interface

  - Replace the LC sprayer with the cleaned CE-MS sprayer with a newly conditioned capillary into the coaxial sheath liquid interface [1].

    - MED: Talent placing CE-MS sprayer with capillary into interface

# 3.7.1A. [Added Shot]: Talent connecting the sheath-liquid line to CE-sprayer. (Editor: Authors may need to provide additional VO for this shot)

  - Then turn on the isocratic pump [1] and apply a voltage of 30 kilovolts for 15 minutes to ensure that a stable CE current profile is achieved prior to the urine analysis [2].

    - MED: Talent turning on pump

    - MED: Talent applying voltage

- CE Injection and Separation Condition Setup

  - In the system software, select Preconditioning [1], set the Flush to 600 seconds, and specify a background electrolyte vial position [2].

    - WIDE: Talent selecting Preconditioning, with monitor visible in frame

    - SCREEN: To be provided by Authors: Flush being set to 600 s and BGE vial position being set

  - Set the Injection to hydrodynamically inject the samples at 100 millibars for 5 seconds and to electrokinetically inject the background electrolyte spacers at 30 kilovolts for 75 seconds [1].

    - SCREEN: To be provided by Authors: Injection being set to hydrodynamically inject samples at 100 mbar for 5 s and to electrokinetically inject BGE spacers at 30 kV for 75 s

  - Set the Applied voltage to 30 kilovolts, the cartridge temperature to 25 degrees Celsius, and the total runtime to 40 minutes [1].

    - SCREEN: To be provided by Authors: Applied voltage being set, cartridge temperature being set, total runtime being set

  - Then, using the timetable, apply a pressure gradient of 2 millibars/minute from 0 to 40 minutes during the separation [1].

    - SCREEN: To be provided by Authors: Pressure gradient being set/applied

  - When the sample acquisition is complete, rinse the capillary at 50 millibars of pressure with background electrolytes overnight while the capillary is in the CE sprayer in the ion source [1-TXT].

    - MED: Talent rinsing capillary TEXT: See text for alternative capillary rinsing/storage details

- TOF-MS Setup

  - [1] [2].

    - WIDE: Talent approaching TOF-MS

    - MED: Talent checking CE unit

  - Set the TOF-MS positive ion detection to span a mass-to-charge ratio of 50-1,700 [1] with a data acquisition rate of 500 milliseconds/spectrum and confirm that both the profile and the centroid data are stored in a “.d” file format [2].

    - MED-over the shoulder: Talent setting m/z ratio, with monitor visible in frame

    - SCREEN: To be provided by Authors: Acquisition rate being set to 500 ms/spectrum, shot of profile and centroid data set to .d file formats

  - Set the electrospray ionization conditions to capillary and nozzle voltages of 2,000 volts each, the nebulizer gas to 10 pounds per square inch, and the drying gas delivery to 8 liters/minute at 300 degrees Celsius, with a sheath gas flow of 3.5 liters/minute at 195 degrees Celsius during acquisition [1-TXT].

    - SCREEN: To be provided by Authors: Vcap/nozzle voltage being set, nebulizer gas pressure being set, drying gas delivery rate being set, then sheath gas flow being set TEXT: Turn off capillary and nozzle voltages during nebulizer gas injection

  - Then set the MS voltage settings of the fragmentor, skimmer, and octpole one radio frequency to 120, 65, and 750 volts, respectively, and set the instrument control and data acquisition parameters according to the design of the experiment [1].

    - SCREEN: To be provided by Authors: Fragmentor, skimmer and Oc1 RF voltages being set, then some control and/or data acquisition parameters being set

- Data Analysis

  - For data analysis, open the multisegment CE-MS data in the system software [1] and, under Chromatograms, set the Extraction data and the Chromatogram and mass spectral data formats to Profile mode [2].

    - WIDE: Talent opening data in software, with monitor visible in frame

    - SCREEN: To be provided by Authors: Extraction data and Chromatogram and mass spectral data formats being set to Profile mode

  - Under Smoothing, select Quadratic-Cubic-Savitzky Golay and set the Function width to 15 points [1].

    - SCREEN: To be provided by Authors: Smoothing being selected, then Quadratic-Cubic-Savitzky Golay being set and the Function width being set

  - Then click Integrate MS and set the integrator to Agile [1].

    - SCREEN: To be provided by Authors: Integrate MS being clicked, then integrator being set to Agile

  - Under Peak filters, set the maximum number of peaks to the 11. Under View, click Integration peak list, and select Peak number, Retention time, Peak area, Peak height, and Signal-to-Noise ratio [1].

    - SCREEN: To be provided by Authors: Peak filters being selected, then maximum number of peaks being set to 11. Then View being opened and Integration peak list being clicked, and Peak number, Retention time, Peak area, Peak height, and Signal to Noise ratio being selected

  - Then save the parameters under a unique method name and apply the method to the process for the interpretation of each data set [1].

    - SCREEN: To be provided by Authors: Parameters being saved under unique method name, then method being applied to one data set

# Section – Results

- Results: Representative High-Throughput and Comprehensive Drug Surveillance

  - In this representative analysis of various isobaric-isomeric drugs of abuse and their metabolites [1], 30 resolved peaks [2] … from 10 independent samples of a drug mixture were detected without sample carry-over [3].

    - LAB MEDIA: Figure 1B

    - LAB MEDIA: Figure 1B: JoVE Video Editor: please emphasize peaks

    - LAB MEDIA: Figure 1B: JoVE Video Editor: please emphasize the 10 different 3-peaked clusters for each sample

  - Two other isobaric opioids could be fully resolved as 20 distinct peaks [1] with the full-scan data acquisition [2].

    - LAB MEDIA: Figure 1C: JoVE Video Editor: please emphasize tall peaks

    - LAB MEDIA: Figure 1C: JoVE Video Editor: please emphasize short peaks

  - Similarly, two amphetamine positional isomers were fully resolved in this multisegment-CE-MS analysis [1] to distinguish illicit methamphetamine abuse [2] from the potential misuse of phentermine, a prescribed stimulant that is used as an appetite suppressant for weight loss [3].

    - LAB MEDIA: Figure 1D

    - LAB MEDIA: Figure 1D: JoVE Video Editor: please emphasize tall peaks

    - LAB MEDIA: Figure 1D: JoVE Video Editor: please emphasize short peaks

  - In this representative drug screening analysis, a positive screening test result for methadone [1] was deduced by the detection of a large signal peak comigrating with a methadone-d3 peak with a low mass error [2].

    - LAB MEDIA: Figure 2A: JoVE Video Editor: please emphasize large blue peak

    - LAB MEDIA: Figure 2A: JoVE Video Editor: please emphasize small red peak under large blue peak

  - No other signals were detected in any other urine samples within the same run [1] and the detected methadone concentration exceeded 13-times the recommended cut-off limit compared to its measured ion response ratio in the sample and corrected by a fourfold urine dilution factor [2], confirming the patient’s adherence to methadone maintenance therapy [3].

    - LAB MEDIA: Figure 2A: JoVE Video Editor: please emphasize flat blue data line between cutoff and large peak and after large peak

    - LAB MEDIA: Figure 2A: JoVE Video Editor: please emphasize cutoff peak

    - LAB MEDIA: Figure 2A

  - Here a serial injection configuration in multisegment-CE-MS comprising five different drug calibrants was analyzed in duplicate [1] within a single run together with a synthetic urine blank [2]. The drug metabolites were detected as their protonated molecular ions above their screening cut-off limits [3].

    - LAB MEDIA: Figures 3A: JoVE Video Editor: please emphasize first 1-5 blue peaks in one block then second 1-5 blue peaks in one block

    - LAB MEDIA: Figures 3A: JoVE Video Editor: please emphasize blank peak(s)

    - LAB MEDIA: Figures 3B: JoVE Video Editor: please emphasize data line OR emphasize orange data points along data line

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Meera Shanmuganathan: Remember to protrude 2 millimeters of the capillary from the CE sprayer to achieve a stable spray formation (step 3.2) [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Philip Britz-McKibbin: Following this procedure, a large number of polar and ionic metabolite in urine can also be analyzed by MSI-CE-MS in support of discovery-based metabolomics research. Additionally, comprehensive surveillance of a wide array of exogenous metabolites and they metabolites can be achieved, including non-prescribed medications and illicit drugs prone to abuse [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Zachary Kroezen: This optimized method can be used to routinely screen other widely-used legal drugs of abuse, including cannabis and alcohol metabolites, as well as to identify synthetic or “fake” urine samples [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Zachary Kroezen: Remember to always use caution when handling biological fluids [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.