## Sheet1
Maintenance and calibrations
Real-time PCR systems | Monthly | Perform background calibrations every month
| 18 months | Perform background, spatial and dye calibrations every 18th months
Centrifuge | 1 year | Calibrate for revolutions per minute and temperature through external or internal calibration services
Glove Box | Daily | Visually inspect elements, particularly for damage to the exposed surfaces of the HEPA filters, gloves, o-rings and hoses. Make sure duct clamps are tight and in place. Perform leak pressure test. Test the pressure alarm.
| 6 months | Change the HEPA filter
| 1 year | Calibrate the system
Autoclave | Weekly | Clean the water tank and racks using a mild non-abrasive detergent
| 3 months | Calibrate timer and gauges
| 1 year or every 50 cycles | Inspect, clean thoroughly, test and calibrate
Refrigerator and Freeezer | 6 months | Check fan motor, evaporator coils, vacuuming condensing coils and condensor filters and replace batteries as needed
| 1 year | Calibrate freezer through internal or external calibration services