Rebuttal
Editorial comments:
General:
1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

The manuscript was proofread by 3 other people.

2. Please ensure that the manuscript is formatted according to JoVE guidelines–letter (8.5” x 11”) page size, 1-inch margins, 12 pt Calibri font throughout, all text aligned to the left margin, single spacing within paragraphs, and spaces between all paragraphs and protocol steps/substeps.

Manuscript was formatted according to the guidelines given above.

3. Please revise lines 93-97, 160-162, and 216-219 to avoid overlap with previous work.

Lines 93-97 has been rephrased as:  V0 is the calibration constant of the sun photometer, R is the earth-sun distance in astronomical units, Vdark  is the dark voltage recorded when light is blocked from passing through the hole on the top bracket of the sun photometer, V is the sunlight voltage recorded from the sun photometer when light passes through the hole on the top bracket, aR represents the attenuation of light due to Rayleigh scattering, P and P0   are the measured and standard atmospheric pressure respectively and m is the relative air mass.

160-162 has been rephrased as: τ1 and λ1 can be replaced with τ2 and λ2 in equation 3 to get the same value for τ3. This calculation is used to compare τ values obtained by two instruments that uses different wavelengths. Ideally the two instruments must be used at the same locality. In our case it must be noted that the two instruments were ~ 96 km apart.

216-219 has been rephrased as: The agreement on the red channel is lower because of the effects of heating on the red LED. The red LED is more sensitive to temperature than the green LED. Agreement for both channels is improved when data collectors control the exposure of the instrument to direct sunlight between measurements.

Summary:
1. Please revise the summary to be 10-50 words long.
Summary is reduced to 47 words.

Protocol:
1. There should be at least 1 page of filmable steps in the protocol section; please expand this (see below).
2. Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed, including specific software/hardware steps. Alternatively, add references to published material specifying how to perform the protocol action. If revisions cause a step to have more than 2-3 actions and 4 sentences per step, please split into separate steps or substeps.

More “how to” details were added to the protocol section.

Figures, Tables, and Figure Legends:
1. Please upload each Figure individually to your Editorial Manager account as a .png or a .tiff file (i.e., 6 files in total). Please combine all panels of one figure into a single image file.

All the images have been saved as TIFF with a resolution 1920 x 1080 pixels.

2. Please remove the embedded figures and table from the manuscript. All figures should be uploaded separately to your Editorial Manager account. Each figure must be accompanied by a title and a description after the Representative Results of the manuscript text.

All embedded figures and tables have been removed from the manuscript and will be uploaded separately. The title and description are given after representative results.

3. Please obtain explicit copyright permission to reuse any figures from a previous publication (if applicable). Explicit permission can be expressed in the form of a letter from the editor or a link to the editorial policy that allows re-prints. Please upload this information as a .doc or .docx file to your Editorial Manager account. The Figure must be cited appropriately in the Figure Legend, i.e. “This figure has been modified from [citation].”

We have no figures from previous publications.

4. Please remove ‘Figure 1’ etc. from the Figures themselves.

Figure labels have been removed from the TIFF figure files.

5. Please define error bars and arrows in the Figure legends.

Error bars and the arrows have been defined in the  figure legends.

6. Figure 1: It’s unclear what this demonstrates separately from what would be presented in the video-either remove or clarify.

Figure 1 was removed from the manuscript

7. Figure 3: Please explicitly explain panel c in the legend.

Panel c was removed from figure 3 since the same data is shown in figure 4.

8. Figure 3b: Please include a label for this panel. Also, please capitalize ‘summer’.

Figure 3b is now clearly labelled and summer is capitalized.

Discussion:
1. As we are a methods journal, please revise the Discussion to explicitly cover the following in detail in 3–6 paragraphs with citations:
a) Critical steps within the protocol
b) Any limitations of the technique

Discussion is rewritten.

References:
1. Please do not abbreviate journal titles.

Journal titles are written in full in the reference section.

Table of Materials:
1. Please ensure the Table of Materials has information on all materials and equipment used, especially those mentioned in the Protocol.

All materials and equipment used are included in the Table of materials.


Reviewers' comments:

Reviewer #1: 
Manuscript Summary:
The manuscript provides nice and simple instructions how to measure the AOT without high end equipment. The titel appropriately describes the method presented. The abstract has a nice length and provides a nice overview of the paper. The introduction describes the presented method and explains why the presented method can be useful. The steps in the protocol are clearly explained and the results are reasonable. The results show the magnitude of the deviations to expect from a nearby AERONET station.

There are some suggestions how to further improve this paper, below:

Introduction:
Provide motivation why to measure AOT? (Health, meteorology), Why do you focus on these four wavelengths (667 nm, 551 nm, 532 nm, 490 nm)?

The following statements were added into the introduction:

The impact of aerosols on the climate and on public health are among the dominant topics in current environmental research24. Aerosols affect the weather by scattering or absorbing light and thermal radiation from the sun and by acting as condensation nuclei in the formation of clouds. Aerosols also play a role in the dispersal of pathogens in the air and they can cause or enhance respiratory and cardiovascular diseases.

For our comparison we focused on these 4 wavelengths, 667 nm, 551 nm, 532 nm and 490 nm. We chose these because they are the 4 AERONET wavelengths closest to the GLOBE sun photometer wavelengths. To make the comparison, we extrapolated AOT values at these wavelengths for XULA site.

What make the wavelengths 505 nm and 625 nm suitable for this?

Generally, sun photometry is based on single wavelengths.  Filter-based photometers have very narrow response bandwidth (~few nanometers). The GLOBE sun photometer is an LED based sun photometer. LED response bandwidths are typically several tens of nanometers. 505nm and 625nm were chosen because the LEDs for these wavelengths have the smallest bandwidths. Other than that there is no particular reason to select these particular colors.

Protocol:
Please clarify whether the voltmeter is a part of the photometer or it is a separate part which has the be manually connected to the photometer (the second sentence in the summary might be read this way)

The following statement was added to the to the protocols section:

The voltmeter on the Globe sun photometer is in-built on the sun photometer.

12 A: please describe the way clouds near the sun are recorded in more detail.

Results:
Please add a y axis label to all the plots

Figure 2: please make the description of this plot more clear. What does "calculated by Globe" mean? This can easily be mixed up with the measured data., because it is measured with the "GLOBE" photometer.
In Figure 3, the "b" is cropped

Protocol step 12A was rephrased to:

## Use the Globe cloud chart to observe and record the clouds near the sun. This is done by looking into the sky and checking off observed features from the GLOBE cloud chart.  Visible cirrus clouds are easy to observe because of their characteristic thin wispy strands. Invisible cirrus clouds are inferred from unexpected voltage readings on the sun photometer.

Figure 2 has been modified and the description clarified, it now shows the daily variation of AOT for the red channel and the green channel for the month of October. The cropped part of Figure 3 has been corrected

Table of materials:
Please include a GPS
Parts of the description of the "Watch" are missing.

Labquest2 GPS has been included in the Table of materials and   the description on the watch has been corrected 


Reviewer #2:
Manuscript Summary:
The manuscript describes a method to monitor aerosol extinction (in terms of AOT) in the visible spectral range using a low-cost hand-held instrument (a sun photometer). The protocol is in principle useful, as it is part of large-scale student project.

Major Concerns:
The manuscript is poorly written, with a systematic incorrect terminology. This must be addressed during the revision stage of the manuscript. A sample dataset obtained with this method is compared with more established automated AOT measurements from a global network; by the way, this comparison is not acceptable, in my opinion, because the validation observatory is too far from the sun photometer site - and in fact the validation results are not good. Mathematical formalism and a few aspects in AOT calculations must be clarified (see Minor Concerns).

The mathematical formalism for comparing AOT values obtained from two different instruments is given in detail below. We added a few of these details into the manuscript.

Wavelength, optical thickness, and atmospheric turbidity are related through Angstrom’s turbidity formula:

τ = β•λ-α			(2)

where τ is the optical thickness, β is Angstrom’s turbidity coefficient, λ is wavelength and α is the Angstrom exponent. α and β are independent of wavelength and they are properties of the atmosphere being measured. For two different wavelengths,

τ1 = β•λ1-α
         	           τ2 = β•λ2-α

from which we get

τ1/(λ1-α) = τ2/(λ2-α)

Solving for α:

α = ln(τ1/τ2)/ln(λ2/λ1)			(3)

From the two τ equations above we can solve for β

β = τ1•λ1α = τ2•λ2α

Given τ at two different wavelengths, the τ at a third wavelength can be deduced for the same atmosphere.

ln(λ3/λ1) α = ln(τ1/τ3) = ln(τ1) - ln(τ3)
             ln(τ3) = ln(τ1) - ln(λ3/λ1)α

τ3 = exp[ln(τ1) - ln(λ3/λ1) α]		(4)

This analysis is used for a rigorous comparison of AOT values obtained by two instruments that uses different wavelengths. Ideally the instruments should be at the same locality. In our case the two instruments are ~90km apart. Thus, at this stage our data can only be used for basic analysis and trend comparison as we did.

Minor Concerns:
1) Title: Please provide a more specific title (e.g., "...using a GLOBE handheld sun photometer"), as this protocol is specific to this sun photometer. Same in the Abstract: please mention the photometer type.

The title was changed to “Measurement of Aerosols Optical Thickness of the Atmosphere using the GLOBE Handheld Sun Photometer” and the GLOBE sun photometer is specifically mentioned in the abstract.

2) L41: "Microscopic": Aerosols size span submicron to millimeters

L41 was changed to “Atmospheric aerosols are minute solid and liquid particles (ranging from submicron to millimeters) suspended in the air”.

3) L44: "Light...": and thermal radiation

Line 44 was changed to “Aerosols in the atmosphere reduces the amount of solar energy reaching the earth’s surface by scattering or absorbing light and thermal radiation from the sun”.

4) L45: "Arbitrary": that's not "arbitrary" of course, maybe "episodic"

Line 45 was changed to “There are seasonal and annual changes as well as episodic changes due to events such as large dust storms, wild fires or volcanic eruptions”5-8.

5) L47: "Sunlight...": it depends on the spectral region. AOT in the UV/VIS mesures "sunlight extinction" while AOT in the TIR measures "thermal radiation extinction"

The GLOBE sun photometer is an LED based sun photometer. The green LED spectral response ranges from 400nm-575nm and the red LED spectral response ranges from 575nm-675nm. Thus they measure sunlight extinction in the Visible range.

6) L52: "Miles...": please use SI units

The unit miles were changed to km throughout the manuscript.

7) L53-54: Being "handheld", so requiring a continuous human non-automatic operation, this is a strong statement that must be mitigated

AOT measurements are done 6 times a day. Unlike an automatic operation like AERONET, students observe the sky and make a decision on whether to accept the measurement or to disregard it.

8) L64: I recommend to mention that other, more precise handheld sunphotometers have been used in the past to carry out research studies, e.g. in areas of difficult access, for exemple:
Sellitto, P., G. Salerno, A. La Spina, T. Caltabiano, L. Terray, P.‐J. Gauthier, and P. Briole (2017), A novel methodology to determine volcanic aerosols optical properties in the UV and NIR and Ångström parameters using Sun photometry, J. Geophys. Res. Atmos., 122, 9803-9815, doi:10.1002/2017JD026723.

This line was added at L64: AOT measurements with small and transportable handheld sun photometers have been successfully used in the past to carry out research studies in remote and hard to access areas17

9) L79-80: How cirrus clouds are detected?

Visible cirrus clouds are easy to observe because of their characteristic thin wispy strands. Invisible cirrus clouds are inferred from unexpected voltage readings on the sun photometer. As part of our Advanced Earth Science Class, students observe clouds on a daily basis using the GLOBE cloud protocol. All our daily cloud observations are compared with NASA’s cloud observation satellites. Observations that coincide with satellite observations receive a ‘match’ email which summarizes both ground and satellite observations. Overtime students became very experienced in cloud observations and in many cases their observations and the NASA satellite observation matched.

10) Eq1: How Rayleigh scattering-related optical depth is calculated?

The contribution to AOT due to Rayleigh scattering is derived from the fundamental physics of the atmosphere. Bucholtz [21] produced tabulated values of  based on standard atmospheres at sea level and has given analytical representations for those tabulations in the form

Where  is specified in microns and A, B, C and D are coefficients obtained in the standard atmosphere table. For the red channel   0.05793 and for the green channel it is  0.13813.

11) How ozone-related optical depth is calculated?

The contribution to AOT due to ozone is calculated using climatological and latitude-dependent average ozone values. Satellite-based instruments such as Total Ozone Mapping Spectrometer (TOMS) are sources of such data. For the GLOBE handheld sun photometer, a typical ozone contribution to non- Rayleigh optical thickness is ~0.01 for green channel and ~0.03 for the red channel. (Ref 22).

12) L100-101: I don't understand this sentence, please explain

Line 100-101 was removed from the manuscript.

13) Eq2: Where is tau2 and lambda2 in this equation? How do you calculate alpha?

For clarity, that part of the manuscript was expanded as follows:

To make a comparison between the two stations, we extrapolated AOT at 667nm, 551 nm, 532 nm and 490 nm for the XULA site. This is done using the XULA site’s atmospheric parameters.  For any given site and instrument, the optical thickness τ, the wavelength λ and the atmospheric turbidity coefficient β are connected through the Angstrom’s turbidity formula

τ=βλ-α		(2)

Where α is the Angstrom’s exponent. α and β are independent of the wavelength at which the optical thickness is measured. They are parameters that describe the atmosphere being measured. Given AOT at two different wavelengths (λ1=505 nm and λ2=625 nm, for our sun photometer), and the measured AOT (τ1 and τ2), the Angstrom exponent α for the XULA site is calculated from the equation,

α=lnτ1τ2/lnλ2λ1 	     (3)

The AOT (τ3) at a third wavelength, λ3 (λ3 being any of the 4 wavelengths of the AERONET site) can be extrapolated for the same XULA atmospheric conditions using the equation:

(4)

τ1 and λ1 can be replaced with τ2 and λ2 in equation 3 to get the same value for τ3. This calculation is used to compare τ values obtained by two instruments that uses different wavelengths.

14) L164-165: I don't understand here: XULA is a site and Globe is an instrument? What are the different series in the comparisons of Fig2?

L164-165 was replaced by:

Figure 1 shows a sample of the typical daily average AOT values calculated using equation 1. This figure shows the AOT data for both the green and the red channels of the GLOBE sun photometer for the month of October.

15) L200-201: One might argue that there is no reason to compare the two datasets due to their distance and typical aerosol horizontal variability...

I agree that we cannot do a rigorous comparison of the two datasets. Because of the distance between the sites, the extrapolated AOT values are at best approximations and cannot be considered for rigorous scientific examination. They were used here for basic analysis and trend comparison. In future we plan to acquire the portable Calitoo sun photometer with blue (465nm), green (540nm) and red (615nm) channels. With these we can validate the GLOBE handheld sun photometer at the XULA site and use both instruments at the AERONET site.

16) L328-330: Is it not trivial and to avoid to be mentioned in the paper?

L328-330 is removed from the manuscript.

17) L330-332: Cirrus just add extinction to the observation, so an additional apparent AOT, please use correct terminology.

L330-332 was changed to: Cirrus clouds adds an additional extinction to the sunlight from the sun.

Reviewer #3: 
Manuscript Summary:
This manuscript used two handheld Sun Photometers to measure the aerosol optical depth (AOD) at Xavier University of Louisiana campus site at two wavelengths. The results were compared with the measurement obtained from the AERONET site. The procedures and the mathematical calculation to obtain the AOD were discussed in the manuscript, as well as some critical steps.

Major Concerns:
The method is lack of a few key considerations, which makes the values obtained from the handheld Sun Photometers doubtful. Equations (1) and (2) are two major equations that the authors used to obtain the AOD. However, equation 1 is different from the original citation (citation 18), because the citation 18 mentions it is important to consider the absorbance of water vapor and other gas species that can absorb in the visible range. For instance, there are water absorption peaks around ~505nm and ~600 nm, so has the author considered subtracting these absorptance caused by the water molecules in the atmosphere? How about other gases?

The method also lacks direct comparison with other instrument to validate the results. Although the authors compared the results with the data obtained from the AERONET site, however there was significant difference between these two sources during 278-288th day of the year (Figure

The correct citation for equation 1 should be [16], David R. Brooks, Development of an inexpensive handheld LED-based Sun photometer for the GLOBE program, Journal of Geophysical Research. 106 (5), 4733-4740 (2001).

The GLOBE sun photometer has two channels at 505nm and 625nm. The absorption of most gases in the atmosphere is outside this band and their contribution to AOT is usually ignored. Only water vapor and ozone are the gases likely to affect AOT especially on the red channel. The contributions to AOT due to ozone is calculated using climatological and latitude-dependent average ozone values. Satellite-based instruments such as Total Ozone Mapping Spectrometer (TOMS) are sources of such data. For the GLOBE handheld sun photometer, a typical ozone contribution to optical thickness is ~0.01 for green channel and ~0.03 for the red channel. This is cited from ref 22. These values are subtracted from the value obtained in equation 1.

The absorption by water vapor occurs in three regions of the spectrum. Rotational transitions absorb in the microwave and far infrared, vibrational transitions absorb in the mid infrared and near infrared. Because of the wide spectral response of our LEDs, it is possible that the red LED’s responds curve is coincident with water vapor absorption region in the near infrared. Ref 24 showed that the estimated error introduced by absorption due to water vapor in the near infrared region is between 0.0015 and 0.0012 AOT units. Measurements made carefully with our GLOBE sun photometer are only accurate to about 0.02 AOT units. Because of this limitation we can also ignore the effects of water vapor.

In future we plan to acquire the portable Calitoo sun photometer with blue (465nm), green (540nm) and red (615nm) channels. With these we can validate the GLOBE handheld sun photometer at the XULA site and also use both instruments at the AERONET site.

2). The difference was near 100%, instead of 5% claimed by the author in the main text. In addition, the AERONET site was too far (60 miles away) to be used as a good validation for the method. Figure 3 also shows a significant difference of AOD between the AERONET site and the campus site, however the author could not give a valid reason to convince the reader that the method proposed was valid. For instance, a recent paper by Toledo et al. (JGR, 2018) shows the effects of water and other trace gases could heavily impact the AOT results. Unless the author can provide direct evidence to validate the method described in the manuscript, the quality of the data is doubtful.

Because of the distance between the sites, the extrapolated AOT values are at best approximations and cannot be considered for rigorous scientific examination. They were used here for basic analysis and trend comparison.

Because of the limitations imposed by the design of the GLOBE sun photometer, we can only measure AOT to 2 decimal places with a minimum uncertainty of about ± 0.02 AOT units. It is true that this presents a substantial percentage error particularly for clear skies where the AOT can be as low as 0.05 AOT units. However, this level of accuracy is still extremely useful to the atmospheric science community.

Thirdly, the graphs in the manuscript could not meet the standard for atmospheric sciences academic publication. The author should use Matlab or Igor to make serious scientific plots instead of Excel. The x and y axes did not have labels and units either.

We currently do not have Matlab or Igor. However, our plots have been converted to TIFF images and the x and y axis are labelled

The discussion part explaining why AERONET site and the campus site had significant difference was not clear. The author did not provide any concrete reasons for the difference and attribute the AOD peaks in February and May to random factors. Did the authors do research to justify whether there were any significant regional pollution? Did the author correlate the AOD data with any major pollution events or forestry fires?

One very importance parameter for a new method is the standard deviation or the error caused by this method compared with the true value. What is the error bar range of this method? The authors need to perform such calculations to show the error bar of the proposed method.

The data for the XULA site and the AERONET site both shows a peak in February and May. We cannot corelate this data with any major pollution events or forest fires. To understand this trend, we need to take measurements for many seasons to add to this preliminary data and see if a definite trend emerges.

During each measurement time, at least 5 values of the voltage V and the dark voltage Vdark are taken and AOT values are calculated using equation 1. Error analysis considerations shows that the retrieval of AOT using the GLOBE sun photometer is most sensitive to errors in calibration constant V0 and the measured voltages V-Vd. The refrence voltage V0 is a result of calibration measurements performed by IESRE at Mauna Loa Observatory. The uncertainty in the two sun photometers we have is between 3% and 5%. The uncertainty in the value of V-Vd recorded by students was calculated based on the standard deviation of the of the 5 values of V and Vd recorded during each measurement time. This ranged between 2% and 6%.

Minor Concerns:
The author used the abbreviation AOT however the full spelling was "aerosol optical depth". The abbreviation of "aerosol optical depth" is "AOD" and the abbreviation of "aerosol optical thickness" is "AOT". Please use the correct abbreviation in the manuscript.

The author stated that their calculation and the value from the GLOBE was only less than 5% difference, but from Figure 2 it shows the difference could be nearly 100% at certain days. The author should make the proper correction in the manuscript.

Aerosol optical depth was replaced with aerosol optical thickness. Figure 2 was replaced and the proper correction was made in the manuscript.