# Submission ID #: 59302

# Scriptwriter Name: Qingyun Ping

# Project Page Link: http://www.jove.com/files_upload.php?src=18087158

Title: Measuring Enzymatic Stability by Isothermal Titration Calorimetry

Authors and Affiliations:  W. K. Dindi Chan1, Marco Mason1, Lee D. Hansen2, and Jason Donald Kenealey1

1Department of Nutrition, Dietetics and Food Science, Brigham Young University, Provo, UT, USA

2Department of Chemistry and Biochemistry, Brigham Young University, Provo, UT, USA

# Corresponding Author:

# Jason D. Kenealey

jason_kenealey@byu.edu

# Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? N

2. Does your protocol include software usage? Y

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot. (You do not need to include steps that will be screen captured. Please do not list entire sections.)

3.5 The loading of the enzyme into the sample cell

3.7 Preparing the titration syringe with substrate

3.8 Removing the buret handle of the ITC and placing the titration in it etc.

4.1 Setting up ITCrun and adjusting the injection interval, volume and number of injections

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1-2 individual steps using the step numbers listed in this document. (Please do not list entire sections.)

3.7 Preparing the titration syringe with substrate - make sure that this step is done correctly and with care to prevent air bubbles in the syringe and bending the syringe tip. This will affect the accuracy of the data.

5. Will the filming need to take place in multiple locations? N

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

- REQUIRED Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

  - Dindi Chan: Using enzymes as catalysts in reactions can reduce the amount of energy needed for the reaction and reduce toxic byproducts.  However, the use of enzymes in industry is limited because enzymes are not stable and can be expensive. This protocol could reduce the time that it takes to determine if conditions are deleterious to enzyme activity or if modifications to the enzyme have increased its stability [1].

    - INTERVIEW

  - Dindi Chan: The main advantage of this technique is that it requires fewer man hours than traditional enzyme stability assays.  Additionally, this protocol can be used with a wide range of enzymes because it can be used with any reaction that produces or absorbs heat [1].

    - INTERVIEW

# Section - Protocol

- Preparing Samples

  - To begin, prepare 0.1 molar sodium acetate buffer [1]. Measure 800 milliliters of distilled water in a 1000-milliliter graduated beaker [2-TXT]. Then, weigh 8.2 grams of anhydrous sodium acetate and add it to the beaker [3].

    - WIDE: Talent prepares for the experiment.

    - MED: Talent measures water in beaker. TEXT: 0.1 M sodium acetate buffer Video editor: Keep the text until 2.3.3

    - MED: Talent weighs chemical.

    - [Split Shot] Adds to beaker

  - Place the beaker on a stir plate [1]. Place a stir rod into the beaker, and turn on the stir plate to stir until completely dissolved [2]. When the anhydrous sodium acetate is completely dissolved [3], use a calibrated pH meter to measure the pH of the solution [4].

    - MED: Talent places beaker on stir plate.

    - MED: Talent places stir rod into beaker and turns on the plate.

    - CU: Shot of the beaker with chemicals completely dissolved.

    - MED: Talent measures pH.

  - Use a pipette to add 1 molar hydrochloric acid or sodium hydroxide accordingly [1] to obtain the desired pH at 4.6 [2]. Add distilled water into the beaker until the total volume reaches 1000 milliliters [3].

    - MED: Talent adds acid or alkaline to solution.

    - CU: Shot of the pH meter showing 4.6.

    - MED: Talent adds distilled water into beaker.

  - To prepare the enzyme solution, in a 15-milliliter graduated cylinder, measure 8 milliliters of the 0.1 molar sodium acetate buffer [1-TXT]. Then, add the buffer solution into a 15-milliliter conical tube with enzyme [2], and shake vigorously until the enzyme has dissolved [3].

    - MED: Talent measures buffer in a cylinder. TEXT: 10-30 mg/mL range enzyme solution Video editor: Keep the text until 2.5.2

    - MED: Talent adds enzyme into cylinder with solution.

    - MED: Talent shakes the cylinder.

  - Add more buffer solution until the total volume reaches 10 milliliter [1]. Store the enzyme solution at 4 °C until use [2].

    - MED: Talent adds more solution to the tube.

2.5.1.B [Split Shot] MED: Talent transfers solution into a 15mL conical tube.

    - WIDE: Talent stores the solution in a refrigerator.

  - To prepare the substrate solution, weigh out the substrate and place it into a 100-milliliter glass beaker [1]. Use a graduated cylinder to measure 20 milliliters of the buffer solution and then add it to the glass beaker [2].

    - MED: Talent weighs substrate and puts into beaker.

    - MED: Talent measures buffer solution and add to beaker.

  - Place the beaker on a stir plate and place a magnetic stir rod into the beaker [1]. Turn on the heat and adjust the stirring speed accordingly [2]. Allow stirring to continue until the substrate has dissolved [3].

    - MED: Talent puts the beaker on stir plate and places in a stir rod.

    - CU: Talent turns on the heat and adjusts speed. Close up of the hand and settings.

    - MED: Shot of the beaker on stir plate stirring.

  - Pour the substrate solution into a 50-milliliter conical tube and add the sodium acetate buffer prepared previously [1] until the total volume is 45 milliliters [2]. Shake the tube to mix [3]. Store the substrate solution at room temperature until use [4].

    - MED: Talent pours solution into tube and adds buffer.

    - CU: Shot of the tube with solution volume at 45 ml.

    - MED: Talent shakes the tube.

    - MED: Talent places the solution on a platform.

- Preparing the ITC (Isothermal Titration Calorimetry) Instrument

  - On the computer [1], open ITCrun (pronounce as I-T-C-run) and click Set up [2]. Click Stirring rate and set to 350 RPM (pronounce as R-P-M) [3].

    - MED: Talent operates on the computer.

    - SCREEN: Talent opens ITCrun and clicks Set up.

    - SCREEN: Talent sets up stirring rate.

  - Check the syringe size is at 50 microliters [1]. Set the temperature to 55 degrees Celsius and press Update [2]. Wait for at least 1 hour before proceeding to allow enough time for the ITC (pronounce as I-T-C) instrument to heat up or cool down as needed [3].

    - SCREEN: Talent checks the syringe size.

    - SCREEN: Talent sets temperature.

    - MED: Shot of the ITC instrument.

  - Before loading the enzyme, ensure that the reference cell is loaded with 350 microliters of distilled water [1]. To clean the sample cell, fill the loading syringe with 500 microliters of 2% cleaning solution [2-TXT]. Carefully insert the needle into the sample cell [3], fill the cell, and slowly remove the liquid using the same syringe [4].

    - MED: Talent checks the reference cell of the ITC instrument.

3.3.1.B [Split Shot] CU: Talent removes the insert from the reference cell

3.3.1.C [Split Shot] MED: Talent check reference cell by drawing water out of the cell

    - MED: Talent fills the syringe. TEXT: See Table of Materials for cleaning solution preparation. Videographer: Take multiple shots, as this will be used later.

    - ECU: Talent inserts the needle into cell. Close up of the needle. Videographer: Take multiple shots, as this will be used later.

    - CU: Talent fills and removes liquid using the syringe. Close up of the syringe and cell. Videographer: Take multiple shots, as this will be used later.

  - Dispose the liquid into a beaker [1]. Repeat this step twice with 2% cleaning solution, three times with 70% ethanol and then wash ten times with distilled water [2].

    - MED: Talent disposes liquid into beaker.

    - Use 3.3.2 & 3.3.4.

  - After the sample cell has been cleaned, fill the loading syringe with 450 microliters of enzyme solution [1], carefully insert the needle all the way to the bottom of the sample cell [2] and slowly press the plunger down to the 100-microliter line to prevent formation of air bubbles [3].

    - CU: Talent fills the syringe with enzyme solution. Close up of the gradient on the syringe.

    - Use 3.3.3.

    - CU: Talent presses the plunger. Close up of the plunger and the 100-microliter line.

  - Then, wash the 50-microliter titration syringe with distilled water three times by placing the needle tip into water to slowly take up the water into the syringe [1], then dispensing the water into a waste container [2]. Rinse with substrate solution three times to remove residual water [3].

    - MED: Talent takes up water into syringe.

    - MED: Talent dispenses water into waste container.

    - MED: Talent takes up and dispenses substrate solution.

  - After that, fill the titration syringe with substrate solution by drawing the solution up until the syringe is full without any air bubbles [1]. With the syringe still in the substrate solution, remove the plunger [2] and allow approximately 2 microliters of air to enter the top of the syringe and reinsert the plunger [3].

    - CU: Talent draws the solution. Close up of the syringe showing no air bubbles.

    - CU: Talent removes the plunger

    - CU:: Talent adjusts the syringe to allow air enter the top, and reinserts the plunger. Close up of the syringe and the air on top.

  - Remove the buret handle of the ITC [1], place the syringe inside the buret handle and screw until tight [2]. Wipe the tip of the titration syringe with a lint-free tissue, then carefully place the buret handle into the ITC instrument and lock it in place [3].

    - MED: Talent removes handle of ITC

    - MED: Talent places syringe into handle and screws.

    - MED: Talent places the handle into the instrument.

- Setting up the ITCrun and Analyzing Data

  - For the experiment setup, select incremental titration. Click Insert to setup the injections [1]. Adjust the injection interval to 5400 seconds, injection volume to 4 microliters and number of injections to 4 [2]. Press OK to confirm settings [3].

    - SCREEN: Talent selects incremental titration.

    - SCREEN: Talent setup injection parameters.

    - SCREEN: Talent presses OK.

  - In the Equilibration box, select Auto-equilibrate and Large expected heats [1-TXT]. Set the initial baseline to 300 seconds [2]. To start the run, click the Start symbol next to Stirring rate and then click Start which is located next to the wrench symbol [3]. Save the file and allow the instrument to run [4].

    - SCREEN: Talent selects in the equilibration box. TEXT: Selecting Small expected heats will increase the equilibration time.

    - SCREEN: Talent sets initial baseline.

    - SCREEN: Talent clicks two start symbols

    - MED: Shot of the instrument running.

  - To analyze data, open the file in NanoAnalyze [1]. Click Data and select Data columns. Select all data, copy and then paste the data into Microsoft Excel [2].

    - SCREEN: Talent opens file.

    - SCREEN: Talent transfers data into excel.

  - Adjust zero baseline by adding the value required at 300 seconds to make it zero [1]. Apply this correction to the entire column of heat rate values [2]. Then, plot graphs of the minimum or maximum values against the time at which the value occurred during the titration [3].

    - SCREEN: Talent adjusts baseline.

    - SCREEN: Talent applies correction to entire column.

    - SCREEN: Talent plots graphs.

# Section – Results

- Results: Enzymatic Stability

  - In this protocol, ITC data traces of lactase activity is shown with four sequential injections of lactose into a lactase solution in pH 4.6 buffer [1] at 55 degrees Celsius [2], 45 degrees Celsius [3], 35 degrees Celsius [4], and 25 degrees Celsius [5]. A no enzyme control was performed at 55 degrees Celsius for the heat of mixing [6].

    - Figure 1

    - Figure 1- Video editor: emphasize the part for 55 °C.

    - Figure 1- Video editor: emphasize the part for 45 °C.

    - Figure 1- Video editor: emphasize the part for 35 °C.

    - Figure 1- Video editor: emphasize the part for 25 °C.

    - Figure 1- Video editor: emphasize the part for control.

  - For each injection of enzyme, there is an initial exothermic heat of mixing [1], and subsequently the lactase catalyzed endothermic hydrolysis of lactose occurs until the reaction is completed and the heat rate returns to the baseline [2].

    - Figure 1- Video editor: emphasize the positive peaks for the four temperature levels and the control.

    - Figure 1- Video editor: emphasize the negative peaks as well as the returning to baseline for the four temperature levels.

  - The endothermic peak minimum following each injection is a little less than the preceding injection because of the decrease of enzymatic activity [1]. As expected, the enzyme stability for each injection decreases with increasing temperature [2].

    - Figure 1- Video editor: emphasize the straight lines for the four temperature levels.

    - Figure 2E

  - The control experiment shows lactase activity at 25 degrees Celsius [1] had an 8% decrease in enzyme activity due to dilution after four injections [2]. Considering the fourth injection shows a 73% decrease in the assay [3], the actual loss of enzyme activity was therefore 65% [4-TXT].

    - Figure 3- Video editor: show figure 3 without the inset.

    - Figure 3- Video editor: zoom in the negative peak and make the inset appear. Emphasize the peak of the grey line and black line, and put an arrow pointing up from the black peak to grey peak and TEXT: 8%.

    - Figure 1- Video editor: show figure 1 with temperature at 25 °C only, and emphasize the fourth injection.

    - Figure 1- Video editor: show figure 1 with temperature at 25 °C only, and emphasize the fourth injection. TEXT: 73% - 8% = 65%

  - Thus, the dilution of the lactase resulted from each injection had a relatively small effect of 11% on enzyme activity [1-TXT].

    - Figure 1. TEXT: 1 - 65% / 73% = 11%

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Dindi Chan: It is important that the buffer used to make the enzyme and substrate solutions are matched as closely as possible. Mismatch of pH or concentration can result in a larger heat of mixing, which can mask the reaction heat [1] [2].

    - Use 2.4.1 & 2.6.1

    - INTERVIEW

  - Dindi Chan: You can apply this protocol to study other enzymes and determine their stability by measuring the rate of heat released or absorbed during the course of a reaction [1].

    - INTERVIEW