# Submission ID #: 59420

# Scriptwriter Name: Laura Rigolo

# Project Page Link: http://www.jove.com/files_upload.php?src=18123148

# Title:

Functional MRI in Conjunction with a Novel MRI-Compatible Hand-Induced Robotic Device to Evaluate Rehabilitation of Individuals Recovering from Hand Grip Deficits

# Authors and Affiliations:

Mark P. Ottensmeyer1,2, Shasha Li2,3,4, Gianluca De Novi1,2, A. Aria Tzika2,3,4

1Medical Device & Simulation Laboratory, Department of Radiology, Massachusetts General Hospital, Cambridge, MA, USA

2Harvard Medical School, Boston, MA, USA

3NMR Surgical Laboratory, Center for Surgery, Innovation and Bioengineering, Department of Surgery, Massachusetts General Hospital Boston, MA, USA

4Athinoula A. Martinos Center of Biomedical Imaging, Department of Radiology, Massachusetts General Hospital, Boston, MA, USA

# Corresponding Author:

A. Aria Tzika			(atzika@hms.harvard.edu)

# Email addresses for Co-authors:

Mark P. Ottensmeyer	(mpo@alum.mit.edu)

Shasha Li			(SHASHA.LI@mgh.harvard.edu)

# Gianluca De Novi		(gdenovi@seas.harvard.edu

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? (Y/N)   N

Can you record movies/images using your own microscope camera? (Y/N) N/A

If no, JoVE will need to record the microscope images using our scope kit (through a camera port or one of the oculars). Please list the make and model of your microscope.

2. Does your protocol include software usage? (Y/N) Y

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot. (You do not need to include steps that will be screen captured. Please do not list entire sections.)

2.2/2.3 (connections through penetration panel)

2.6 (mounting of device to MR table, protecting from being knocked over by patient)

2.8 (secure device at comfortable position for patient)

2.10/2.11 (dynamometer function to set patient maximum force level)

3.5 (visual metronome display and subject gripping/releasing in synchrony)

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1-2 individual steps using the step numbers listed in this document. (Please do not list entire sections.)

3.1 Ensure proper securing/fixation of the subject’s head to avoid head motion imaging artifacts

5. Will the filming need to take place in multiple locations? (Y/N) Y– MR room and conference room at MGH Charlestown Navy Yard.

If yes, how far apart are the locations? – same building, different floors

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

- REQUIRED Interview Statements: (Said by you on camera)

  - Dr. Shasha Li: This protocol is used to monitor the rehabilitation of patients with hand grip deficits. Chronic stroke patients and those with other neurological diseases involving motor deficits may benefit from this technique [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Dr. A. Aria Tzika: MR imaging allows monitoring of the adaptability, in other words, neuroplasticity of the brain, resulting from rehabilitation, rather than simply monitoring grip performance [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

OPTIONAL Interview Statements: (Said by you on camera)

  - Dr. Mark P. Ottensmeyer: With suitable modification of the force stimulus device, this method can be applied for rehabilitation of disabilities affecting other regions of the body.  This device and any modifications must retain MR-compatibility [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

  - Dr. Mark P. Ottensmeyer:

# It’s important to ensure that the MR-CHIROD has been properly connected and

# its functions verified before bringing a subject into the MR room [1].

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

# Introduction of Demonstrator: (Said by you on camera)

  - Dr. A. Aria Tzika: Dr. Li will be demonstrating operation of the MR scan sequence and Dr. Ottensmeyer will operate the controls of the Magnetic Resonance Compatible Hand Induced Robotic Device, that we call MR-CHIROD [1].

    - The named technician, post doc, student looks up from workbench or desk or microscope and acknowledges the camera. (Author Comment: introduction shot done in the library with Li and Ottensmeyer standing to either side of Tzika, not done in the scanner control room)

Ethics title card: (for human subjects or animal work, does not count toward word length total)

  - Procedures involving human subjects have been approved by the Partners Healthcare Institutional Review Board (IRB) at Massachusetts General Hospital.

# Section - Protocol

# Note to videographer:  The MRI scanner is a very strong magnet that cannot be turned off.  To avoid serious damage and injury: Do not bring a camera or any other equipment into the MRI room. Follow directions of MRI staff at all times to determine a safe filming distance.  (Disregard if the Mock scanner room is used.)

- Scanner Room Setup

  - Prior to starting this experiment, first obtain informed consent from the subject and thoroughly screen them for MRI safety [1]. Do not proceed with the scan if the participant has any potential MRI contra-indications [2].

    - WIDE: Talent and the subject are sitting in a waiting area (or similar area outside the scanner). Talent goes over the consent form/experiment with the subject.

    - MED: Talent goes over the MRI safety checklist with the subject, confirming answers on the form.

  - To begin set up, first bring the Magnetic Resonance-Compatible Hand Induced Robotic Device into the MRI room and place it near the penetration panel [1]. Then, insert the three eighths inch pneumatic tube into the pass-through tube in the panel into the adjacent MRI support room [2]. (Videographer note, important step)

    - Talent brings the device through the doorway to the MRI room and places it near the penetration panel.  (film through doorway – do not enter the MR room)

    - Talent inserts the 3/8-inch pneumatic tube into the pass-through tube in the panel into the adjacent MRI support room.  (Videographer note – if this cannot be filmed from the MR door, film the tube being pushed through the pass through into the adjacent MRI support room. You may be able to get a shot through the pass through (these are often a tube through the wall). (Author Comment: FINAL NOTE – could not bring camera into MR room, filmed passage of 3/8” tube through large, metal “pass-through” tube from external side in the MRI support room)

  - Next, connect the force sensing and encoder cables to the 9-pin D-shaped connector on the MRI scanner room side of the panel [1]. Then connect the support room force sensing and encoder cables to the DSUB (pronounce: dee-sub) connector on the external side of the penetration panel as shown here [2]. (Videographer note, important step)

    - Talent connects the MRI scanner room force sensing and encoder cables to the 9-pin D-shaped connector on the scanner room side (can only be filmed in Mock scanner). If this cannot be filmed in the room: show the force sensing and encoder cables that will be plugged in, perhaps talent holds them in their hand and/or points to the pins/connector.  (Author Comment: FINAL NOTE – could not bring camera into MR room, so 2.3.1 was skipped.)

    - Talent connects the support room force sensing and encoder cables to the DSUB connector on the external side of the penetration panel (outside scanner room).

  - Connect the 3/8-inch pneumatic tube fitting, emerging from the penetration panel, to the outlet of the interface power unit pressure regulator outlet [1]. Then connect the 4-millimeter pneumatic tube to the outlet of the compressor and the inlet of the air filter on the interface power regulator [2].

    - Talent connects the 3/8-inch pneumatic tube fitting, emerging from the penetration panel pass-through to the outlet of the interface power unit.

    - Talent connects the 4 mm pneumatic tube to the outlet of the compressor and the inlet of the air filter.

  - After extending and lowering the scanner bed, attach the bottom half of the head coil [1] and have the volunteer lay down, making sure that they are resting comfortably with their arms extended [2].

    - After extending and lowering the bed, talent places the lower half of the head coil onto the bed and assures it is attached.

    - Talent helps the subject lay down on the scanner bed, they place their arms down by their side, and talent places a cushion under the knees. (Author Comment: FINAL NOTE – first version of shot includes 2.5.1, 2.5.2, 3.1.1 and 3.1.2.  Additional shots of 3.1.1 and 3.1.2 were also made.  3.1.1 and 3.1.2 are moved here and relabeled as 2.5.3 and 2.5.4)

# 3.1. 	Attach the top of the head coil and use small foam pads to immobilize the participant’s head [1]. Also attach the communication ball on the volunteer’s chest, and provide instructions on its use [2].

    - Talent attaches the top piece of the head coil, then places foam pads around the subject’s head. (Editor: Originally numbered 3.1.1)

    - Talent tapes the communication ball to the subject’s head, then speaks to them providing instructions. (Editor: Originally numbered 3.1.2)

  - Now loosely install the robotic device on the side of the patient opposite to their brain lesion using the corresponding bed-slot [1]. Then, with the volunteer’s elbow resting on the table to support their arm, move the device handle to the webbing between thumb and forefinger and help them grab the handles [2]. (Videographer note, important step)

    - Talent loosely install the Robotic Device opposite to the brain lesion, using the bed slot.  (Suggest setting up on the opposite side of the table from the penetration panel, so that we can film the next step.)

    - The volunteer’s elbow is resting on the table as talent helps them grab the handles.

  - If the device is on the opposite side of the table from the penetration panel, position the cables and the pneumatic tube so that they pass under the table rather than over the patient [1].

    - Talent positions the cables and the pneumatic tube so that they pass under the table.  (Note, may cut this shot if the device is not on this side. Info will remain in the manuscript) (Author Comment: FINAL NOTE – This shot was not filmed because the device was positioned on the near side of the table.  Voice-over can still be used immediately following note “[1]” in item 2.6.) (Editor: I’ve left this VO in just in case it can be used after 2.6.1 as the authors suggest. However, if it is too long or cannot easily be used, omit it entirely)

  - Next, instruct the volunteer to squeeze and push or pull the device until they have the most comfortable position for squeezing [1]. Then, secure the device firmly in place by tightening the plastic nuts using an MR-compatible wrench [2]. (Videographer note, important step)

    - Talent speaks to the subject, then they squeeze and push or pull the device until they have the most comfortable position.

    - Talent secures the device firmly in place by tightening the plastic nuts using an MR-compatible wrench.

  - Now, run the custom user interface stimulus program [1]. Set the pressure to the minimal setup level to automatically push the handle to the end-stop [2], verifying display of motion and force waveforms [3].

    - MED Over shoulder: Talent starts the user interface. Videographer: Please obtain some extra footage of talent at the computer using the UI, in case this is needed elsewhere.

    - SCREEN: Provided by authors: 59420_screenshot_1.mp4, time index 0:00.00.  (“MR-CHRODv3 Control Interface. Setting up comms with device.” Then “connected to electronics module…”)

    - MED Over shoulder: Talent verifies the display of motion and force waveforms.

  - Next set the force level and instruct the volunteer to completely squeeze 2 to 3 times for approximately 2 seconds [1]. Observe whether the volunteer can complete a squeeze at that force level [2]. (Videographer note, important step)

    - SCREEN: Provided by authors 59420_screenshot_1.mp4, time index 0:41.84. (“instruct the subject that grip force testing will begin now” appears, then “OK” is clicked on.)

    - Film through the MRI window (or door) as the volunteer squeezes the handle. If this is difficult to film, show talent looking through the window, as if they can observe the subject doing this.

  - Gradually increase the force level and repeat squeeze attempts until they cannot complete a squeeze [1]. This measurement serves as the volunteer’s maximum grip strength [2]. The program will automatically calculate 60%, 40%, and 20% of the maximum force level for use during testing [3]. (Videographer note, important step)

    - MED Over shoulder: Talent increases the force level on the UI. (or general shot of UI use)

    - Film through the MRI window (or door) as the volunteer squeezes the handle but cannot complete a squeeze.

    - SCREEN : Provided by authors: 59420_screenshot_1.mp4, time index 1:13, before the last increase of force level from 200 to 250 and demonstration by the subject that they can grasp/release at the new level.  End at 1:29 (before the section setting up the graphical display.)

- Scan Session

  - [1].[2]

# (Move all of 3.1 to after 2.5)

    - Talent attaches the top piece of the head coil, then places foam pads around the subject’s head.

    - Talent tapes the communication ball to the subject’s head, then speaks to them providing instructions.

  - Next, confirm that the communication ball works [1], then position using the laser alignment before moving the table and the participant into the scanner’s isocenter [2].

    - The subject squeezes the ball, then talent presses the button on the side of the scanner to silence it.

    - Talent aligns the head with the laser, then pushes the button to send the participant into the scanner.

  - Now, using the user interface, generate the instruction and stimulus images, and set the system to apply the first force level and wait for a trigger signal from the MRI scanner [1]. The program will display a set of instructions to remind the volunteer how to respond to the visual stimulus [2].

    - SCREEN : Provided by authors: 59420_screenshot_1.mp4, start at 1:31 (instruction to let go), continue through about 1:53 at which time the force level will be at the first set point, or “Prepare MR sequence”.

    - SCREEN : Provided by authors: Instructions are shown in: 59420_screenshot_2.mp4, time index 0:00.00 to 0:05.50.

  - The program will wait for the scanner to provide a trigger signal, then remove the instructions [1].  and show a “fixation cross” that the volunteer should focus on [2].

    - SCREEN : Provided by authors: 59420_screenshot1.mp4, time index 1:49.90, signaled by text display in the Matlab command window.(trigger# 1.0000  at 3.1736 appears.)

    - SCREEN : Provided by authors: crosshair displayed. 59420_screenshot_2.mp4, time index 0:06.00 to 0:26.50.

    - Note: critical step statement with this step.

  - When the fMRI scan acquisition begins, a visual metronome will be displayed, in the form of a growing and shrinking circle [1]. The volunteer should completely squeeze and release the handle synchronously with this stimulus [2].  (Videographer note, important step)

    - SCREEN : Provided by authors: 59420_screenshot_2.mp4, time index 0:26.60 to 0:47.56 for metronome (growing and shrinking circle)

    - The volunteer completely squeezes and releases the handle synchronously. (film from observation window or through door).  If possible, we see the visual metronome is only displayed on the projector.  If this cannot be filmed, pan down to show images being captured in real time on the MR console.  Videographer: Please obtain some extra footage of scanner running while the subject performing the task, in case this is needed elsewhere.

  - Rest periods will separate stimulus periods, during which time the fixation cross will again be displayed [1]. Observe the live plots of force and displacement to monitor the force output and participant task performance [2].

    - SCREEN : Provided by authors: 59420_screenshot_2.mp4, start at 0:46 to show metronome briefly, then changes to crosshair (rest) at 0:47.56.

    - SCREEN : Provided by authors: 59420_screenshot_1: 2:53, (rest, flat line) followed by cycles on the graph.

  - Once the experiment is complete, let the participant know they can relax and to let go of the handle [1]. Finally, collect a series of anatomical scans. [2].

    - The participant relaxes and releases the handle. (film through window).  If this is hard to film, perhaps show talent speaking through the intercom to the subject.

    - MED over shoulder: Talent scrolls through an anatomical MRI on the scanner console computer, while another scan is being acquired. (Or mock acquisition.)

OPTIONAL – Critical Step Statement:

# Dr. Mark P. Ottensmeyer, Step   3.4:   The scanner sends trigger signals at the start of each at each phase of the scan, which are used to synchronize the metronome display with the acquisition repetition times. (Editor: This statement is very close to the first statement in the conclusion. If they’re both delivered as written, then one can be omitted if they feel similar. However, both can be kept if needed)

# Section – Results

- Results: MRI-Compatible Hand-Induced Robotic Device Motor Task

  - This figure shows typical motor task results [1]. Here we see fMRI activations, superimposed on a brain outline [2], and as pseudo-color on a three-dimensional cross-sectional view of the volunteer’s anatomical image. M1 indicates the Primary motor cortex, and SMA indicates the Supplementary motor cortex. [3]. This image shows pseudo-color activations rendered on a brain template.

    - LM: Figure 5: Show whole figure.

    - LM:  Note, all text in blue to be done by a JoVE video editor: Show Figure 5A alone with second sentence.

    - LM: Figure 5B appears next to 5A.  Indicate M1 when spoken, then SMA when spoken.

    - LM: Show Figure 5 C

  - This graph shows actual force output, measured in units of force as a function of time. The output was recorded in real time. The Black bar corresponds with the 60 second stimulus and rest period.

    - Figure 5D. JoVE video editor: indicate the black bar when spoken.

  - Here a single-voxel time course of activation is shown, chosen from a voxel at the somatosensory area at the location of the cross-hairs in this image.

    - LM: Figure 5 E and B. JoVE video editor: 5E appears, then 5B appears alongside it when “location of the cross-hairs” is spoken.

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Dr. Shasha Li: All subjects must be properly trained in performing the metronome-tracking grip motions in advance. In addition, synchronization between the visual stimulus and the MRI sequence is crucial.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera B-roll suggestion: step 3.4

  - Dr. Shasha Li: Additional imaging modalities can be used in this protocol, including Diffusion Tensor Imaging to detect white matter fiber tract orientation and growth which is also expected to change with rehabilitation.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Dr. A. Aria Tzika: Development of this process will allow the demonstration that stroke recovery continues beyond six months post-injury, suggesting that therapy and associated insurance reimbursement should continue beyond this time.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Dr. Mark P. Ottensmeyer: Performing any MR experiment is hazardous due to the powerful magnetic field generated by the scanner.  Therefore, all subjects and experimenters must follow MR-safety guidelines, including screening for MR contraindications.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.