# Submission ID #: 59513

# Scriptwriter Name: Anthony Iannazzi

# Project Page Link: https://www.jove.com/account/file-uploader?src=18151513

# Title: Preparing a Celadonite Electron Source and Estimating its Brightness

# Authors and Affiliations:

# Evelyne Salançon, Alain Degiovanni, Laurent Lapena, Roger Morin

# CINaM, Aix-Marseille University, CNRS, UMR 7325, Marseille, France

# Corresponding Author:

# Evelyne Salançon			salancon@cinam.univ-mrs.fr

# Email Addresses for Co-authors:

# degiovanni@cinam.univ-mrs.fr

# lapena@cinam.univ-mrs.fr

# morin@cinam.univ-mrs.fr

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? Yes

Can you record movies/images using your own microscope camera? Yes

2. Does your protocol include software usage? No

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot.

2.2.1, 2.3.2, 2.4.2, 2.9.1, 3.3.2, 4.4.2

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1-2 individual steps using the step numbers listed in this document.

2.9.1

The most difficult aspect was to find a way to deposit only one micro-crystal at the apex of the carbon fiber. This step is essential.

5. Will the filming need to take place in multiple locations? Yes

If yes, how far apart are the locations? 2 corridors.

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

- REQUIRED Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Salançon Evelyne: This protocol aims explaining how to make a celadonite electron source. These sources have shown a long lifetime and its brightness is equivalent to that of field emission metal tips [1].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.

  - Salançon Evelyne: Using this electron source in a projection microscope associated with an electrostatic lens makes it possible to work with a large source-object distance. This prevents the source crash-risk and image distortion by decreasing local field effect on the object [1].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.

# Section - Protocol

- Preparation of the Source

# 2.0.      To begin this procedure, insert the source into stainless-steel tube with an inner diameter of 90 micrometers. Insert a tungsten wire – with a diameter of 50 micrometers – into the tube [1], and roll the tube under a knife blade to cut it to the required length [2].

# 2.0.1. Talent shows the tube and the tungsten wire.

# 2.0.2. Talent cuts the tube.

  - Fix the source support underneath an optical microscope [1]. Insert a 10 micrometer carbon fiber into the stainless-steel tube [2] and glue the carbon fiber to the tube with silver lacquer [3].

    - Establishing shot of the talent approaching the optical microscope, and securing the source support underneath the microscope. Videographer: This is one of the most important steps for viewers to see.

    - Talent inserts a carbon fiber into the stainless-steel tube. (Author Comment: We also filmed a shot called 2.1.2scope) (Editor: I’m not sure where or how the authors want this shown. If it looks good enough, perhaps use it here instead of 2.1.2)

    - Talent glues the carbon fiber to the tube with silver lacquer. (Author Comment: We also filmed a shot called 2.1.3scope) (Editor: I’m not sure where or how the authors want this shown. If it looks good enough, perhaps use it here instead of 2.1.3)

  - Under a binocular microscope, use cutting tweezers to cut the fiber so that between 100 micrometers and 3 millimeters are left outside of the stainless-steel tube [1].

    - SCOPE: *To be provided by authors. Microscope shot of the fiber being cut with cutting tweezers. Authors, please upload all microscope videos to your project page.

  - Next, grind the celadonite with a mortar and pestle [1]. Weigh out 0.2 milligrams of celadonite powder [2] and dilute it in 10 milliliter of deionized water [3].

    - Talent uses a mortar and pestle to grind the celadonite.

    - Talent weighs the celadonite powder. Videographer: This is one of the most important steps for viewers to see.

    - Talent dilutes the celadonite powder in deionized water.

  - Place an ultrasound tip directly into the celadonite-containing water [1] and use an ultrasonic frequency of 30 kilohertz and a power of 50 watts for 30 seconds to break up the aggregates [2].

    - Talent places an ultrasound tip directly into the celadonite-containing water.

    - Talent turns on the sonicator and sonicates the water. Videographer: This is one of the most important steps for viewers to see.

  - To prepare the deposition environment, connect a capillary holder to a pressure controller [1]. Maintain the capillary holder under an optical microscope with a multidirectional micro-manipulator [2].

    - Talent connects a capillary holder to a pressure controller.

    - Talent places the capillary holder into a micro-manipulator under an optical microscope.

  - Place the support under the microscope with the carbon fiber facing the capillary holder [1].

    - Talent places the support under the microscope.

  - Next, fix a glass capillary in a puller jaw [1]. Using Table 1 of the text protocol, ensure that the puller parameters are set properly according to the patch pipette size, and pull a micropipette with an internal-end diameter between 2 – 10 micrometers to allow the dispersed celadonite to flow without obstruction [2].

    - Talent fixes a glass capillary in a puller jaw.

    - Talent checks the pullers parameters, and pulls a micropipette as described.

  - Then, fill the micropipette with the celadonite-containing water [1]. Under the microscope, mount the micropipette on the capillary holder [2] and align the micropipette and the carbon fiber [3].

    - Talent fills the micropipette with celadonite-containing water.

    - SCOPE: *To be provided by authors. Microscope shot of the micropipette being mounted on the capillary holder. Authors, please upload all microscope videos to your project page. (Author Comment: The beginning of the film named shot2.8.3_2.9.1_2.9.3 corresponds to the alignment of the carbon fiber and the micropipette, and to the pressure apply. The meniscus is growing.) (Editor: Looks like 2.8.2 – 2.9.2 were combined. The file name mentions 2.9.3, so this might be in the combined files as well, but the authors didn’t mention it in their note)

    - SCOPE: *To be provided by authors. Microscope shot of the micropipette being aligned with the carbon fiber. Authors, please upload all microscope videos to your project page.

  - Move the micropipette to a distance approximately between 2 and 10 micrometers from the apex of the carbon fiber [1]. Apply the progressive pressure on the wide end of the micropipette such that a drop forms at the tip but does not fall, which will wet the apex of the carbon fiber [2-TXT]. After this, retract the micropipette [3].

    - SCOPE: Microscope shot of the micropipette being moved close to the apex of the carbon fiber. Videographer: This is one of the most important steps for viewers to see, and one of the most difficult to perform.

    - SCOPE: Microscope shot of a drop forming at the tip of the micropipette. TEXT: Typically, apply 100 mbar. Videographer: Please also capture a shot of the talent applying pressure on the wide end of the micropipette.

    - SCOPE: Microscope shot of the micropipette being retracted.

- Kicking-off the Source

# 3.0.      Under a microscope [1], insert the source into the source-support [2].

# 3.0.1. Talent places the source-support under the microscope.

# 3.0.2. SCOPE: The stainless-steel tube is inserted into the source-support.

  - Install the source holder under the vacuum [1]. Connect the carbon fiber and the object to two high-voltage electrical feedthroughs [2]. Check the electrical continuity of each contact [3-TXT] and install the flange on the experimental setup [4].

    - Talent installs the source holder under the vacuum.

    - Talent connects the carbon fiber and the object to two high-voltage electrical feedthroughs.

    - Talent checks the electrical continuity of each contact. TEXT: Contacts: Anode-object, lens and screen.

    - [Added Shot]: Talent installed the vacuum flange with the source on the experiment.

  - After this, turn on the vacuum pumping [1]. Connect a nano-ammeter of a caliber in the microampere range between the object and the electrical ground [2].

    - Talent turns on the vacuum pump.

    - Talent connects a nano-ammeter between the object and the electrical ground.

  - Increase the negative bias voltage applied to the source slowly, at approximately 1 volt per second [1]. If the anode is 1 millimeter away from the source, the kick-off takes place at around 2 kilovolts when the intensity suddenly increases [2].

    - Talent increases the negative bias voltage applies to the source.

    - Shot of the talent as they monitor the source during the kick-off. Alternatively, film a shot of the readout at the intensity suddenly increases. Videographer: This is one of the most important steps for viewers to see.

  - Then, immediately decrease the voltage to stabilize the intensity by some hundred nano-amperes [1]. At the beginning, the intensity can fluctuate over several orders of magnitude [2].

    - Talent decreases the voltage.

    - Shot of the readout showing that the intensity is fluctuating over several orders of magnitude.

  - Let the system fluctuate for several hours [1], until the fluctuations decrease [2]. Cut off the voltage when the fluctuations are lower than 10 percent [3].

    - Talent walks away from the system as it continues to fluctuate. Alternatively, film a shot of the readout, showing the fluctuations decrease.

    - Shot showing the fluctuation decreasing.

    - Talent cuts off the voltage.

- Source Characterization

  - To begin, use the rotating flange to turn the source towards the simple projection setup to observe the electron beam [1]. Use the micro-manipulator to decrease the source-to-screen distance and obtain the entire spot on the screen [2]. Measure the source-to-screen distance [3-TXT].

    - Talent uses the rotating flange to turn the source towards the simple projection setup.

    - Talent uses the micro-manipulator to decrease the source-to-screen distance.

    - Talent, looking at the screen, measures the source-to-screen distance. TEXT: D = source-to-screen distance.

  - Take pictures of the screen by using the rotating flange to change the angle between the electron beam and the normal to the screen [1]. Plot the gray-level intensity profile along one axis, and determine the emission radius at a given source-to-screen distance [2-TXT]. Calculate the cone angle as outlined in the text protocol [3].

    - Talent uses the rotating flange to change the angle between the electron beam and the normal to the screen, and takes pictures.

    - LAB MEDIA: Figure 3. TEXT: R = emission radius. Video Editor: Show only Figure 3B.

    - Talent calculates the cone angle. Any action taken in this process can be filmed for this shot.

  - After this, measure the emission intensity versus the voltage applied to the source, with the intensity measured at the anode and the voltage applied at the carbon fiber [1]. Create a Fowler-Nordheim plot of the celadonite source as outlined in the text protocol [2]. The curve will show a decreasing straight line with saturation for highest voltage. The longest straight line is the signature of the field emission process [3].

    - Talent measures the emission intensity versus the voltage applied to the source.

    - Talent creates a Fowler-Nordheim plot of the celadonite source. Alternatively, the talent can be filmed at a workstation computer as they review a previously generated plot.

    - LAB MEDIA: Figure 4.

  - To measure the source size, use the rotating flange to turn the source towards the electrostatic lens [1]. Produce a projection image containing a huge Fresnel diffraction patter along the edge of an object [2-TXT].

    - Talent uses the rotating flange to turn the source towards the electrostatic lens.

    - SCOPE: *To be provided by authors. Talent produces a projection image containing a huge Fresnel diffraction patter along the edge of an object. TEXT: Magnification of ~20,000x is required. Authors, please upload all microscope videos to your project page.

# (Author Comment: From 00:00 to 00:49 : intensity adjustment with lens 1 activated. At 00:49 : G=718. Then, the object is approached towards the source, the magnification increases. At 01:04 : G=1015. At 01:14 : G=1359. At 01:17 : G=1890. At 01:24 :G=2797. At 1:40 : G=5355. At 01:46 : G=6442. We stop the approach, the lens 1 is cut and L2 is activated. At 02:00 : highest growing for L2, and reactivation of L1 to obtain a magnification of 23010 at 02:04.) (Editor: I’m not sure how much of this, or which parts, the author wants shown. I’d either show the last actions, or attempt to speed up the clip to show most of it)

  - Measure the sharpest visible detail on the image on-screen, and calculate the source size as outlined in the text protocol [1].

    - LAB MEDIA: Figure 5.

# Section – Results

- Results: Analyzing the Celadonite Electron Source and Estimating its Brightness

  - Several scanning electron micrograph images of celadonite deposited on carbon fibers were obtained at 15 kilovolts. Sources exhibit one, sometimes two, crystals at their apex [1].

    - LAB MEDIA: Figure 1.

  - However, the use of the SEM involves another support for the carbon fiber, which is hard to mount and demount without breaking. It is safer to attempt direct electron emission. Tests in a projection microscope show that every source prepared this way emits. The kick-off is required only once [1].

    - LAB MEDIA: Figure 2.

  - Most of these sources show one single point source. The emission profile indicates only one continuing image without any other spot [1].

    - LAB MEDIA: Figure 3. Video Editor: Show only Figure 3C.

  - The Fowler-Nordheim plot exhibits 10 orders of magnitude straight and saturation at higher voltage. The saturation regime obtained for a given voltage depends on the structure, but the slope decreases systematically for higher current intensities from about 10 microamperes [1].

    - LAB MEDIA: Figure 4.

  - The source size is then estimated by measuring the smallest detail on the image produced [1]. This image is the Fresnel diffraction pattern of the object [2]. Here, loss of interference fringes is attributed to the size of the source [3].

    - LAB MEDIA: Figure 5. Video Editor: Emphasize Figure 5A during “…the smallest detail on the image produced”.

    - LAB MEDIA: Figure 5. Video Editor: Hold the emphasis from 5.5.1.

    - LAB MEDIA: Figure 5.

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Salançon Evelyne: In this protocol, the most important is to obtain a single celadonite deposited on the apex of a tip-like shape conductor to be able to approach an object towards the celadonite to image this object [1]. Likely, the crucial step described in 2.9 where a small drop of the well dosed celadonite containing water is deposited at the apex of the fiber [2].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.

    - B-roll suggestion: Shots from 2.9.

  - Salançon Evelyne: Using the celadonite electron source in a projection microscope equipped with an electrostatic lens allows working with a large source-object distance. This makes it possible to develop off-axis holography techniques to explore magnetic and electric fields around nanometric objects, for example [1].

    - INTERVIEW: Named author says the statement above in an interview-style shot while looking slightly off-camera.