# Submission ID #: 59553

# Scriptwriter Name: Melissa Ceo

# Project Page Link: https://www.jove.com/account/file-uploader?src=18163713

Title: Electrochemical Roughening of Thin-film Platinum Macro and Microelectrodes

Authors and Affiliations: Anna N. Ivanovskaya1*, Anna M. Belle1*, Allison Yorita1, Fang Qian2, Supin Chen1, Angela Tooker1, Rose García Lozada1, Dylan Dahlquist1, Vanessa Tolosa1

1Engineering Directorate, Lawrence Livermore National Laboratory

2Physical and Life Science Directorate, Lawrence Livermore National Laboratory

*These authors share equal contribution

# Corresponding Author:

# Anna M. Belle

# belle1@llnl.gov

# Email addresses for Co-authors:

# ivanovskaya1@llnl.gov

# yorita1@llnl.gov

# qian3@llnl.gov

# supin@neuralink.com

# tooker1@llnl.gov

# garcialozada1@llnl.gov

# dahlquist2@llnl.gov

# vanessa@neuralink.com

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy? N

2. Does your protocol include software usage? Y

3. Which steps from the protocol section below are the most important for viewers to see? 2.1.1, 2.2.1, 2.2.2, 4.1.1, 4.1.2, 4.7.1

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? 4.1

5. Will the filming need to take place in multiple locations? N

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

Authors, these headshots will be used for the JoVE Dedicated Author Webpage. Here is one example if you wish to take a look.

- REQUIRED Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Anna Ivanovskaya: Our protocol further develops an existing approach used over the last four decades, extends the applicability of electrochemical roughening to thin films and opens the door for miniaturization [1].

    - INTERVIEW: Anna Ivanovskaya says the statement above in an interview-style shot, looking slightly off-camera.

  - Anna Belle: By increasing the surface area of platinum electrodes without an additional coating, the devices are more robust and can last longer than a plated device for electrical stimulation [1].

    - INTERVIEW: Anna Belle says the statement above in an interview-style shot, looking slightly off-camera.

OPTIONAL Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Allison Yorita: We have worked with many different electrode geometries and found that parameters like electrode size, shape and layout influence roughening. We encourage researchers to vary pulsing parameters for their specific electrodes [1].

    - INTERVIEW: Allison Yorita says the statement above in an interview-style shot, looking slightly off-camera.

# Section - Protocol

- Cleaning the Pt Electrode(s) Before Initial Characterization and Surface Roughening

  - First, submerge the electrode tip of a device in a 500 millimolar perchloric acid solution that also contains a platinum wire counter electrode and mercury sulfate reference electrode [1].

    - Talent submerges the electrode tip of a device in a perchloric acid solution containing a platinum wire counter electrode and mercury sulfate reference electrode. Videographer: The authors have indicated that this step is important (questionnaire question 3).

  - Connect several short electrodes of a multielectrode device together as the working electrode [1]. Then, connect the working, counter and reference electrodes to a potentiostat [2].

    - Talent connects several short electrodes of a multielectrode device together. Videographer: The authors have indicated that this step is important (questionnaire question 3).

    - Talent connects the working, counter and reference electrodes to a potentiostat. Videographer: The authors have indicated that this step is important (questionnaire question 3).

  - To electrochemically clean the surface of the electrodes by repetitive potential cycling, first open the EC-Lab software of the potentiostat. To apply cyclic voltammograms, or CVs to the electrodes, press the plus sign to add the electrochemical technique under the Experiment tab. In the pop-up window, Insert Techniques will appear [1].

    - SCREEN: 2.2.4.1.mp4. Show up to 0:19.

  - Click on Electrochemical Techniques. When that expands, click on Voltamperometric Techniques. When that expands, click on Cyclic Voltammetry [1].

    - SCREEN: 2.2.4.1.mp4. Show from 0:19 to 0:27.

  - In the Experiment window, fill in the appropriate parameters. Under Safety/Advanced Settings, select the Electrode Connections as CE to ground [1].

    - SCREEN: 2.2.4.1.mp4. Show from 0:33 to 1:26.

  - Connect the working, counter and reference electrodes to the instrument leads as shown on the Electrode connection diagram [1].

    - Talent connects the working, counter and reference electrodes to the instrument leads.

  - Press the Run button and select the filename to begin the experiment [1].

    - SCREEN: 2.2.4.1.mp4. Show from 1:26 to 1:37.

  - Perform repetitive potential cycles until the voltammograms visually appear to overlap from one cycle to the next, which typically occurs after 50 to 200 CVs [1].

    - SCREEN: 2.2.4.2.mp4

- Electrochemical Characterization of the Electrode Surface Before Roughening

  - To perform the electrochemical characterization, submerge the electrode tip of the device in a beaker of deoxygenated 500 millimolar perchloric acid that also contains a platinum wire counter electrode and mercury sulfate reference electrode [1]. Video Editor: The authors have indicated that this step is important (conclusion statement 6.1).

    - Talent submerges the electrode tip of the device in a beaker of perchloric acid containing a platinum wire counter electrode and mercury sulfate reference electrode.

  - Under the Experiment tab in the EC-Lab software, press the plus sign to add the electrochemical technique. In the pop-up window, Insert Techniques will appear. Click on Electrochemical Techniques. When that expands, click on Voltamperometric techniques. When that expands, click on Cyclic Voltammetry [1].

    - SCREEN: 3.2.1.mp4. Show from 0:17 to 0:26.

  - In the Experiment window, fill in the appropriate parameters. Under Safety/Advanced Settings, select the Electrode Connections as CE to the ground [1].

    - SCREEN: 3.2.1.mp4. Show from 0:30 to 1:15.

  - Connect the working, counter and reference electrodes to the instrument leads as shown on the Electrode connection diagram [1].

    - Talent connects the working, counter and reference electrodes to the instrument leads.

  - Press the Run button and select the filename to begin the experiment [1].

    - SCREEN: 3.2.1.mp4. Show from 1:16 to 1:25.

  - Perform repetitive potential cycles until the voltammograms visually appear to overlap from one cycle to the next [1]. If the two cathodic peaks of a platinum CV are poorly resolved, estimate the electrode surface area from the double layer capacitance at the electrode-solution interface [2].

    - SCREEN: 3.3.1.mp4

    - LAB MEDIA: Poorly Resolved Peaks.png, Poorly resolved peaks.emf

  - To measure the impedance spectra of a single electrode under open circuit conditions, first submerge the electrode tip of the device in PBS that also contains a platinum wire counter electrode and mercury sulfate reference electrode [1]. Connect one electrode at a time as the working electrode [2].

    - Talent submerges the electrode tip of the device in a beaker of PBS containing a platinum wire counter electrode and mercury sulfate reference electrode.

    - Talent connects one electrode at a time to the instrument leads.

  - Under the Experiment tab in the EC-Lab software, press the plus sign to add the electrochemical technique. In the pop-up window, Insert Techniques will appear. Click on Electrochemical Techniques. When that expands, click on Impedance Spectroscopy. When that expands, click on Potentio Electrochemical Impedance Spectroscopy [1].

    - SCREEN: 3.6.1.mp4. Show from 0:13 to 0:22.

  - In the Experiment window, fill in the appropriate parameters. Under Advanced/Safety Settings, select the Electrode Connections as CE to the ground [1].

    - SCREEN: 3.6.1.mp4. Show from 0:29 to 1:27.

  - Connect the working, counter and reference electrodes to the instrument leads as shown on the Electrode connection diagram [1].

    - Talent connects the working, counter and reference electrode to the instrument leads.

  - Press the Run button and select the filename to begin the experiment [1].

    - SCREEN: 3.6.1.mp4. Show from 1:27 to 1:35.

- Electrochemical Roughening of a Macroelectrode and Characterization of Electrode Surface After Roughening

  - Submerge the electrode tip of the device in a beaker of deoxygenated 500 millimolar perchloric acid that also contains a platinum wire counter electrode and mercury sulfate reference electrode [1]. Then, connect an individual electrode as the working electrode and apply the pulsing paradigm to roughen the electrode [2]. Video Editor: The authors have indicated that this step is important (questionnaire question 4 and conclusion statement 6.1).

    - Talent submerges the electrode tip of the device in a beaker of perchloric acid containing a platinum wire counter electrode and mercury sulfate reference electrode. Videographer: The authors have indicated that this step is important (questionnaire question 3).

    - Talent connects an individual electrode and applies the pulsing paradigm. Videographer: The authors have indicated that this step is important (questionnaire question 3).

  - Begin the roughening protocol with a series of oxidation/reduction pulses between minus 0.15 volts and 1.9 to 2.1 volts at 250 Hertz with a duty cycle of 1 to 1 for 10 to 300 seconds [1-TXT].

    - LAB MEDIA: Figure 1. TEXT: -0.15 V, 1.9 – 2.1 V, 250 Hz, 10 – 300 s.

  - Open the VersaStudio program for the PAR potentiostat. Expand the Experiment menu and select New. In the Select an Action pop-up window, choose Fast potential pulses and enter the desired file name when prompted. The Fast potential pulses line will then appear under the Actions to be performed tab [1].

    - SCREEN: 4.3.1.mp4. Show from 0:11 to 0:23.

  - Under the Properties for Fast Potential Pulses, enter the Number of pulses as 2, Potential 1 as minus 0.59 volts versus Ref for 0.002 seconds, and Potential 2 as 1.56 volts versus Ref for 0.002 seconds [1]. Under Scan properties, enter the Time per point as 1 second, the number of cycles as 50,000 for a 200 second duration [1].

    - SCREEN: 4.3.1.mp4. Show from 0:23 to 1:06.

  - Under Instrument properties, enter the  Current range as Auto. Program the potentiostat with a prolonged application of a constant reduction potential by first pressing the plus button to insert a new step. Click on Chronoamperometry. Enter the Potential as minus 0.59 volts, the Time per point as 1 second and the Duration as 180 seconds. Press the Run button to start the roughening [1].

    - SCREEN: 4.3.1.mp4. Show from 1:06 to 1:38.

  - The program will stop automatically when the roughening procedure is complete [1].

    - SCREEN: 4.4.1.mp4

  - After the roughening is finished, determine the increase in effective surface area of the macroelectrodes as previously described [1].

    - Talent places the macroelectrode under the microscope and looks through the microscope to confirm roughening has occurred. Videographer: The authors have indicated that this step is important (questionnaire question 3).

# Section – Results

- Results: Characterization of Roughened Thin-Film Platinum Electrodes

  - A schematic showing the voltage application for roughening both macroelectrodes and microelectrodes is shown here [1].

    - LAB MEDIA: Figure 2

  - Optical microscopy can be used to visualize the difference in the appearance of a roughened macroelectrode or microelectrode [1].

    - LAB MEDIA: Figures 3 and 4 – Video Editor: Highlight figure 3 at mention of “roughened macroelectrode” and figure 4 at mention of “microelectrode”.

# In addition, electrochemical characterization of the platinum surface using impedance spectroscopy and cyclic voltammetry can clearly show the increased active surface area of a roughened macroelectrode and microelectrode [1].

    - LAB MEDIA: Figures 1 and 5 – Video Editor: Highlight figure 1 at mention of “a roughened macroelectrode” and figure 5 at mention of “microelectrode”.

  - The relationship between surface roughness and the number of roughening pulses applied for macroelectrodes is shown here [1].

    - LAB MEDIA: Figure 4

  - An example of different roughening parameters to maximally increase electrode active surface area for different electrode geometries is shown here [1].

    - LAB MEDIA: Table 1

# Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

  - Anna Ivanovskaya: Use high purity electrolytes for roughening (steps 3.1, 4.1). Dilute high purity perchloric acid with deionized water and use only dedicated glassware [1].

    - INTERVIEW: Anna Ivanovskaya says the statement above in an interview-style shot, looking slightly off-camera.

  - Anna Ivanovskaya: This procedure will improve techniques that benefit from high surface area, such as enhancing optical spectroscopy signals of surface adsorbed species, increasing electrochemical reactor efficiency and improving biosensor characteristics [1].

    - INTERVIEW: Anna Ivanovskaya says the statement above in an interview-style shot, looking slightly off-camera.

  - Anna Belle: This approach to electrode roughening will allow researchers to roughen the surfaces of thin-film electrodes for many different applications without compromising the integrity or lifetime of the electrode [1].

    - INTERVIEW: Anna Belle says the statement above in an interview-style shot, looking slightly off-camera.

  - Allison Yorita:  Perchloric acid is hazardous. When working with this reagent, use all appropriate personal protective equipment and only handle in a fume hood [1].

    - INTERVIEW: Allison Yorita says the statement above in an interview-style shot, looking slightly off-camera.