TITLE:

Implementation of a nonlinear microscope based on stimulated Raman scattering

AUTHORS & AFFLILIATIONS:

Rajeev Ranjan1, Maria A. Ferrara1, Maurizio Indolfi1, Luigi Sirleto1

1National Research Council (CNR), Institute for Microelectronics and Microsystems, Unit of Naples , 80131, Italy

Corresponding Author:

Luigi Sirleto 		(luigi.sirleto@cnr.it)

Email Addresses of Co-authors:

Rajeev Ranjan 		(rajeev.ranjan@na.imm.cnr.it)

Maurizio Indolfi 	(maurizio.indolfi@na.imm.cnr.it)

Maria A. Ferrara 	(antonella.ferrara@cnr.it)

KEYWORDS:

microscopy, imaging systems, nonlinear optics, stimulated Raman scattering, ultrafast optics, label free imaging, bio-imaging.

SUMMARY:

In this manuscript, the implementation of a stimulated Raman scattering (SRS) microscope, obtained by the integration of an SRS experimental set-up with a laser scanning microscope is described. Our SRS microscope is based on two femtosecond (fs) laser sources: a Ti-Sapphire (Ti:Sa) and a synchronized optical parametric oscillator (OPO).

ABSTRACT:

Stimulated Raman scattering (SRS) microscopy  uses near-infrared excitation light, therefore it shares many of the multiphoton microscopic imaging properties. SRS imaging modality can be obtained using commercial laser-scanning microscopes by equipping them with a non-descanned forward detector with proper bandpass filters and lock in amplifier (LIA) detection scheme.

A schematic layout of a typical SRS microscope is the following: two pulsed laser beams, i.e., the pump and probe directed in a scanning microscope, must be overlapped in both space and time at the image plane and focused by a microscope objective into the sample through two scanning mirrors (SMs) which raster the focal spot across an X–Y plane. After interaction with the sample, transmitted output pulses are collected by upper objective and measured by a forward detection system inserted in an inverted microscope. Pump pulses are removed by a stack of optical filters, whereas the probe pulses which are the result of the SRS process occurring in the focal volume of the specimen, are measured by a photodiode (PD). The readout of the PD is demodulated by the LIA to extract the modulation depth. A two dimensional (2D) image is obtained by synchronizing the forward detection unit with the microscope scanning unit.

In this paper, we describe and successfully demonstrate the implementation of SRS microscope, reporting label free image of  polystyrene beads with the diameter of 3 µm. It is worth noting that SRS microscopes are not commercially available, so in order to take advantage of their characteristics, the home-made realization is the only possible option. Since SRS microscopy is becoming popular in many fields, we believe that a careful description of SRS microscope implementation could be very useful for the scientific community.

INTRODUCTION:

In life science applications, SRS microscopy has emerged as powerful tool for label-free imaging The basic idea of SRS microscopy is to combine the strength of vibrational contrast and its ability to acquire images in a few seconds.

In SRS, the sample is excited by two collinear laser beams (pump signal and Stokes signal) at different frequencies. When the difference in frequencies is equal to a molecular vibration of the sample, a stimulated and coherent excitation of molecular bond vibration modes (third order non-linear process) occurs, which causes a significant increase of Stokes signal. Unlike linear Raman spectroscopy, SRS exhibits a nonlinear dependence on the incoming light fields and produces coherent radiation. SRS has two fundamental advantage: speed, which makes images less sensitive to artefacts arising from sample movement or degradation, and an excellent signal-to-noise ratio (SNR). In addition, SRS exhibits an identical spectrum as the spontaneous Raman, and SRS signal is linearly proportional to the concentration of the chemical bond excited1-5.

In our microscope, a femtosecond (fs) SRS experimental set up is integrated with an inverted optical microscope equipped with a fast mirror scanning unit (see Figure 1)6-8. Two pulsed laser sources are used to implement this microscope. The first is a fs-Ti:Sa with a pulse duration of approximately 140 fs, with a repetition rate of 80 MHz, and emission wavelengths in the range of 680 nm – 1080 nm. The second, used as probe beam and pumped by Ti:Sa, is a femtosecond synchronized optical parametric oscillator (SOPO), with a pulse duration of approximately 200 fs, having a repetition rate of 80 MHz, emitting in the range of wavelengths 1000 nm – 1600 nm. We note that the minimum photon energy difference between Ti:Sa and SOPO beam is 2500 cm-1. Therefore, using this combination of laser systems, only the high frequency C–H region (2800–3200 cm-1) of Raman spectra can be explored6-8.

In order to set up an SRS microscope, we have to deal with ‘three crucial issues’, which are described in the successive paragraphs. The first is the implementation of a high-frequency modulation transfer method (see Figure 2 and step 2.1 of protocol for description). In an SRS experimental investigation, a crucial parameter is the sensitivity of the system. SRS signal is detected as a small change of the intensity of excitation beam, and hence it is easily deteriorated by shot noise and laser intensity noise. We overcome this noise issue by implementing high frequency modulation. In this method, the pump beam is modulated by an electro-optic modulator (EOM) and the modulation transferred to the probe beam is measured with a PD, after blocking the pump beam with an optical filter (stimulated Raman gain (SRG) detection mode). The PD output is connected by a low pass filter to a lock-in amplifier (LIA) which demodulates the measured signal. By increasing modulation frequency of the beam to frequencies above 1 MHz, the intrinsic limit of PDs can be obtained.

The second, is the installation of a mechanical mount  which permits to carry out forward detection and at the same time to preserve microscope observation in brightfield. In addition it has to reduce the noise due to mechanical vibration during the generation of images and to allow the precise repositioning of detection system (see Figure 3 and step 2.2 of protocol).

The third is the synchronization of the signal acquired by the phase-sensitive detection scheme with the beam position onto the sample monitored by the scan head of the microscope. In order to realize images, the SMs require three TTL signals made available by the microscope controller connected to the scan head unit : “pixel clock”, “line sync”, and “frame sync”. The synchronization is achieved by controlling using a PCI card the three TTL signals and the acquisition of a voltage signal at the output channel of LIA6-8. A home made software has been developed and described in references6-8 while the hardware of the synchronization system is reported in Figure 4.

A fundamental procedure in carrying out SRS imaging is the microscope alignment. It is realized in 4 steps, which are described in the successive paragraphs. The first, is the spatial overlap of two beams (see step 3.1 of protocol for description). In our experimental set up, the two beams were spatially collinearly combined by a dichroic mirror. The preliminary step is the alignment of OPO and Ti:Sa, so that each of  them reaches the microscope. Then, considering OPO as a reference beam and taking advantage of a position sensitive detector, the Ti:Sa is spatially overlapped to OPO.

The second, is the temporal overlap of two beams, (see step 3.2 of protocol for description). Even if pump and OPO beams are perfectly synchronization9, since they follow slightly different beam paths inside the OPO housing,  at the OPO exit, they  have a time delay of about 5 ns and spatial difference of 5 cm. Therefore, Ti:Sa and OPO require being retimed optically to ensure temporal overlap at the sample. This is typically accomplished with a finely tunable optical delay line, which in our case is inserted between the Ti:Sa and the microscope (see figure 1). In order to obtain the temporal overlap of two beams, two techniques are used. The first one is carried out using a fast PD and an oscilloscope, while the second one is based on auto and cross optical correlation. Using the first technique,  a rough overlap of two beams is obtained (uncertainty of 10 ps), while an accurate temporal overlap of two beams is obtained by using a crosscorrelator (resolution of 1 fs).

The third, is the alignment of the two beams inside the microscope, (see step 3.3 of protocol for description)., A preliminary white light observation of sample allows to individuate the desired field of view (FOV). Afterwards, laser beams, entering the microscope by a side port of microscope, are aligned in order to reach the PD mounted on the upper part (figure 3) . However, for a correct image acquisition, the setting of a number of parameters is required, for example pixel dimension and pixel dwell time. The sampling frequency must respect the constraint imposed by Nyquist’s theorem in order to preserve all information in an image, while for a correct correspondence between the spatial coordinates of pixels and the SRS value measured in each pixel, the integration time of LIA should be equal or comparable to the pixel dwell time.

In the final step of microscope alignment, a number of tests are carried out in order to optimize the spatial and temporal alignment (see step 3.4 of protocol). A number of transmission images (TI), for both Ti:Sa and OPO, are acquired in order to optimize spatial overlap. In a TI, a single beam is used and the transmitted beam intensity from the sample is measured by a PD. In the case of TI realized by OPO, the PD output signal is directly connected to PCI card, while in case of TI realized by Ti:Sa, the PD output signal is connected to LIA and the analog output of LIA is connected to PCI card. The transmission images are very useful to optimize the FOV, the illumination, the focal position of microscope objectives and to check if the two beams are spatially overlapped6-8.

The optimization of pump and probe beams temporal overlap is obtained by scanning the delay line with steps of 0.001 mm corresponding to 3.3 fs time-shift and carrying out a SRS measurement in a single point of a polystyrene beads sample of 3 µm diameter. The amplitude of SRS signal measured values from LIA,  as a function of the probe–pump delay provides  a maximum, in correspondence of exact temporal overlap of two beams6-8. Before concluding, we point out that all the discussed steps are mandatory in order to obtain a good quality image.

PROTOCOL:

- Start up the Laser system

- Check if the temperature of chillers is maintained at or below 20°C.

- Check if the humidity control unit  is working properly and humidity is maintained at a value around 40%.

- Turn on the Ti:Sa laser, strictly following the instructions in the manual.

- Set the wavelength to 810 nm.

- Turn on the OPO and the connected mini-computer. Run the application which controls the OPO laser.

- Select bypass if 100% of Ti:Sa laser output  is required at the exit of OPO box.

- Deselect bypass if 20% of Ti:Sa laser output and 80% of OPO laser output are required at the exit of OPO box.

- Open the shutter of Ti:Sa and release the Ti:Sa beam to OPO input.

- Release the two laser beams at OPO exit by clicking on the signal-out and pump-out.

- Wait till both lasers Ti:Sa and OPO are stabilized (it takes around 45-60 minutes).

- Verify the beams spot at the exit of OPO box for both Ti:Sa and OPO using paper detector card and check the power using a power meter.

1.12 Tune the OPO laser wavelength to 1076 nm.

1.13 Reduce power to about 10 mW for each laser beam in order to perform alignment.

- Setting up the microscope

2.1. Implementation of high-frequency modulation transfer method

- Carry out the optical alignment procedure of the modulator, such that the Ti:Sa beam enters and exits without any distortion.

- Turn on the function generator and generate a TTL signal (a square wave with amplitude 5 V, offset 2.5V, frequency 5 MHz).

- Divide the TTL signal in 2 parts using a T junction; one for EOM while the other one for the Lock-in amplifier (LIA) (see Figure 2).

- Check all the signal levels with an oscilloscope.

- Turn on the LIA and connect the generator output channel to the reference channel of LIA.

- Connect the generator output channel to the high voltage power amplifier of EOM.

- Turn on the amplifier and set the voltage almost to the maximum level. Monitor the beam power at the exit of EOM.

2.2. Integration of mechanical mounting to fix the PD and assign X, Y relative motion

Note: Here with microscope we introduce an external mount equipped with micrometer which have motion control in X and Y direction.

- Mount two travel translation Stages in order to allow movements along the X and Y directions (see figure 3).

- Fix the stage on a Ø1.5" Post of appropriate height.

- Mount the PD to an external mount.

- Maximize the beam intensity at PD adjusting the PD positions (x and y coordinate) using the micrometers attached to the mount (see figure 3).

- Alignment of Microscope

3.1. Spatial overlap of the beams

Note: Considering the OPO beam as a reference, and taking advantage of a position sensitive detector, the Ti:Sa is overlapped to OPO according to the following procedure:

- Carry out the alignment for OPO laser beam in order to reach the microscope. Carry out the alignment for Ti:Sa laser beam in order to reach the microscope.

- Place the laser beam position sensors detectors in two positions in between DM1 and M4, the first close to DM1 and the second close to M4, and detect, for each position, the both X and Y coordinates of the OPO beam (follow Figure 1).

- Verify, that the X and Y coordinates of Ti:Sa laser beam are the same of OPO in both the positions of position sensors detectors. If in some positions the coordinates of Ti:SA and OPO are not coincident, tune the tilt of the adjacent mirror in order to compensate the difference (follow figure 1).

- Follow the same procedure to align the Ti:Sa beam positions with respect to OPO for the path in between M4-M5 (follow figure 1).

3.2. Temporal synchronization of the beams

3.2.1. Procedure using fast photodiode plus an oscilloscope

- Stop propagation of Ti:Sa and OPO beams and put a fast detector in front of OPO beam (in between M6 and M7).

- Connect the trigger signal provided by Ti:Sa laser box with oscilloscope in channel 2.

- Connect the detector cable with oscilloscope in channel 1 and visualize OPO temporal profile.

Record the time (abscissa) measured by the oscilloscope corresponding to its maximum value, namely t1

- Stop OPO beam and release Ti:Sa beam.

- Visualize Ti:Sa temporal profile and record the time (abscissa) corresponding to its  maximum value , namely t2.

3.2.1.7 Minimize the difference t1-t2 using the delay line, in order to overlap the two beams. In our case the minimum measurable difference is 10 ps.3.2.1.8 Remove the fast detector in between M4 and M5.

3.2.2. Procedure using an autocorrelator

Note: In the schematic diagram figure 1, we install autocorrelator without interfering the optical paths of the beams. In addition, we introduce an additional mirror mounted on a flip-flop mount (referred to as FFM/AM) in between M6 and M7 in order to divert the beam into the autocorrelator.

- Flip the AM in order to direct the beam into autocorrelator.

- Stop the Ti:Sa and release OPO.

- Set the beam distance adjustment screw micrometer of autocorrelator in Normal position (8.35 mm).

- Power ON the autocorrelator controller and start the software application on the personal computer controlling it .

- Project the OPO beam from FFM/AM to the input mirror into the autocorrelator.

- Control the reflection spot (using paper detector card) of the beam on the alignment window of autocorrelator.

- In case of no-beam intensity or low-beam intensity adjust the position and orientation of FFM/AM to optimum extent, also try to adjust the input mirror (mounted on the autocorrelator) to maximize the laser pulse signal. The autocorrelator signal is obtained as shown in the figure 5 (a).

- Stop OPO and projecting the Ti:Sa beam from FFM/AM to input mirror into the autocorrelator. Repeat the steps 3.2.2.6 and 3.2.2.7. The autocorrelator signal is obtained  as shown in the figure 5 (b).

- Set the beam distance adjustment screw micrometer to Cross position (7.30 mm).

- Release both beams.

- Scan the delay line in order to obtain the two beams OPO and Ti:Sa overlapped. The cross-correlator signal is obtained   as shown in the figure 6

- Again flip the mirror FFM/AM, so the beams can  reach M7 and the scan head of microscope.

3.3. Microscope alignment

3.3.1. White light Microscopic observation

Note: Prior to microscopic observation be sure that the microscope is properly aligned

3.3.1.1. Prepare the test sample; it consists in a phosphate buffer solution in which polystyrene beads, with a diameter of 3 μm, are dispersed. The solution is placed inside a sandwich of two glass slides.

3.3.1.2. Turn on the microscope and the power supply of white light. Follow the manual for the observation in white light.

3.3.1.3. Use condenser to illuminate the sample. Use objective of 60X to collect light. Place the sample onto the stage. Optimize the focal position of 60X microscope objective.

3.3.1.4. Select the FOV of interest. Take a CCD image of the sample (Figure 7).

3.3.1.5. Turn off the power supply of white light.

3.3.2. Microscope alignment with femtosecond laser beams: OPO and Ti:Sa

- Remove condenser using escape button to retract temporarily the 60X microscope objective lens. Move the 60X microscope objective lens off the optical path, rotating the nosepiece.

- Mount the detector to the upper part of microscope using the external mechanical mount. Connect the detector output by a low pass filter of 50Ω to oscilloscope to monitor the OPO signal.

- Turn on the processor which control the scanner head. Project the OPO beam into the scanner head of the microscope.

- Check the position of the beam inside the microscope, make sure that the location of the beam is in center or near to the center.

- Check the position of the beam inside the head of PD and check that the location of the beam is in center or near center microscope.

- Maximize the power measured by detector using x-y translator.

- Switch the beam from OPO to Ti:Sa and verify they both are on almost the same position. It indicates that both beams are well aligned.

3.3.2.8 Finalize the beam alignment, introducing again the 60X microscope objective lens rotating back the nosepiece.

3.3.2.9 Use refocus button in microscope to regain the finalized focus to the 60X microscope objective lens.

3.3.2.10 Place the objective with magnification of 40X in place of condenser without touching or disturbing the sample.

3.4	Optimizations of spatial and temporal synchronization of beams

3.4.1.	Temporal synchronization

- Set the power of Ti:Sa and OPO measured before the microscope to 30 mW, for both beams. Set the wavelength of OPO at different value with respect to the previous one, (so that pump and probe are not in resonance with vibrational frequency of beads).

- Release both beams, i.e. Ti:Sa and OPO, so that they enter the microscope.

- Run the scanning delay line computerized translator and record the measured intensity by LIA for each position of delay line. Wait till the delay line scanning is complete. The obtained temporal profile is visualized in figure 8 (a).

- Set the wavelength of OPO to 1076 again ( so that pump and probe are in resonance with vibrational frequency of beads). Repeat points 3.4.1.3, .. the obtained temporal profile is visualized in figure 8 (b).

- Set the obtained overlap beam position in the delay line, in order to acquire SRS images.

3.4.2. Spatial synchronization of beams

Note: The transmission images are very useful to optimize the FOV, the illumination, the focal position of microscope objectives and to check if the two beams are spatially overlapped.

3.4.2.1. Transmission image acquisition of OPO

- Stop the Ti:Sa beam and reduce the OPO power to 8 mW.

- Connect detector readout to the data acquisition card.

3.4.2.1.3. Run the data acquisition program along with microscope scanning console.

3.4.2.1.4. Save the file and process the data to get the image. The raw image appears like in figure 9(a).

3.4.2.2. Transmission image acquisition of Ti:Sa

- Stop the OPO beam and reduce the Ti:Sa power to 2.5 - 4.5 mW.

- Connect the detector with LIA and LIA readout with data acquisition card.

3.4.2.2.3. Repeat the steps 3.4.2.1.3 and 3.4.2.1.4. The raw image appears like in Figure 9 (b).

- SRS image acquisition

Note: A dedicated algorithm has been realized in order to store data. It supports the following image formats: 512px × 512px and 256px × 256px, with acquisition times 16 s, 8 s, 4 s and 2 s.

- Introduce a stack of filters in between the objective 40X and the PD, in order to remove the pump pulses (Ti:Sa) and acquire only the Stokes signal (OPO).

- Set the pump signal to 810 nm with a focused power of 8 mW, and the probe signal to 1076 nm with a focused power of 8 mW, in order to investigate a typical C-H bond of polystyrene (Raman shift of 3054 cm−1).

- Connect the detector with LIA, and LIA readout to data acquisition card.

4.4. Set the image acquisition pixel format as per requirement (we have supported format of 512px × 512px and 256px × 256px,) and set the acquisition time.

NOTE: We have options of 16 s, 8 s 4 s and 2 s.

4.5. Run the program, which controls the microscope controller.

4.6. Run the dedicated algorithm program which acts as synchronization between the microscope controller, detection system and DAQ (see figure 4).

4.7. Save the matrix file once the acquisition is completed.

4.8. Import the raw data file and save the image in required format (typically, we save in .tif format) using ImageJ software. The image is shown in the figure 10.

REPRESENTATIVE RESULTS:

An example of SRS measurement, i.e. SRS measurement in a single point of the sample, is reported in figure 7. When the beam is not overlapped in time or space, the obtained result is reported in figure 7 (a). In off-resonance the amplitude of signal measured by LIA is zero, while the phase of signal measured by LIA jumps among negative and positive value. In resonance, i.e. when the beams are overlapped in space, moving the delay line in an appropriate range, the obtained results are reported in figure 7(b). The signal measured by LIA increases and reaches its maximum when the beams are perfectly overlapped, while the phase start to get a fixed value during the time that the beams are overlapped.

The absorption image(s) obtained using a single beam (Ti:Sa or OPO) of the same polystyrene beads are represented by the figure 9 (a) and 9 (b), with the scale bar of 6 µm. In order to acquire the SRS images, the delay line is set to the position achieved in the figure 7(b), a typical SRG image is shown in the figure 10 with the scale bar of 6 µm.

FIGURE AND TABLE LEGENDS:

Figure 1: Schematic layout of the f-SRS microscope system.

OPO = Optical Parametric Oscillator; Ti:Sa = Ti:Sapphire laser; M1-M4 = Femtosecond Mirror; M5, M6, M7, FFM/AM = Broadband Mirror DM1, DM2 = Dichroic Mirror; FFM/AM = Flip-flop Mirror/ Autocorrelator Mirror; ACF =  Autocorrelation Function, EOM = Electro-Optic Modulator; FG = Function Generator; GM = Galvo Mirror; Obj1, Obj2 = Objective 1 and 2; PD = Photodiode; DAQ = Data acquisition system; PC = Personal Computer.

Figure 2: Schematic of electronic scheme for phase detection system.

Inset figures: Before sample represents the two lasers beams before interaction inside the sample and After Sample represents the modified probe due to interaction of probe and pump inside the sample. Time represented in ns whereas laser output is in mW.

Figure 3: Pictorial representation of photodiode mount with mechanical mounting system.

Figure 4: Schematic of Data acquisition system; PD = photodiode, LIA =  Lock-in amplifier, DS = Detection system, MC = Microscope control, DAQ= Data acquisition system, PC = Personal computer

Figure 5: Autocorrelator signal with OPO and Ti:Sa.

(a) OPO signal on the Autocorrelator

(b) Ti:Sa signal on the Autocorrelator

Figure 6: Raman peak of the Autocorrelation Function (ACF).

Figure 7: Peak with lock in amplifier (LIA).

Figure 8: CCD image of the sample polystyrene beads. Scale bar: 6 µm

Figure 9: Absorption images of OPO (a) and Ti:Sa (b) of polystyrene beads. Scale bar: 6 µm

Figure 10: SRS image of polystyrene beads. Scale bar: 6 µm.

DISCUSSION:

The recent development of SRS microscopy has boosted the interest in the study of biological structures such as lipids, which are vital to cells and fundamental for the cellular architecture. They are implicated in multiple physiological pathways: making up the biological membranes, serving as fuels, they are biosynthetic precursors and signal transducers10. The lipids are packed into specialized intracellular organelles, termed lipid droplets (LDs), which occur abundantly in adipose and steroid-producing cells, but they are also present in other cellular lines. Their sizes vary from some tens of nanometers to tens of micrometers in diameter11-12. LDs participate in a broad variety of physiological processes, such as lipid storage, and they figure prominently in common pathologies, for example, the altered cholesterol metabolism has a major health impact from cardiovascular diseases to neurological disorders13-14.

In order to visualize lipids inside the cell, traditionally, fluorescence microscopy techniques, using labeled fixed cells with neutral lipid-specific dyes, are used.10 However, the lipids are smaller than proteins and DNA; therefore, their structures and functions are more susceptible to changes and unwanted artifacts, occurring adding fluorophores15-16. SRS has proven to be particularly powerful for studying lipid-rich structures. Lipids are abundant in C-H2 groups. Thus in their Raman spectrum, the relatively isolated peaks associated with C-H bond vibrational states at 2845 cm-1 provide a unique signature for lipids inside a cell, distinguishing them from other chemical species. With this technique it is possible to image, for example, the hydrocarbon bonds (C-H) in the chemical composition. We note that it is rather difficult to distinguish a target biomolecule from the sea of the other related species inside cells since the differentiable vibrational signatures are finite and many bio-molecular species share similar chemical bonds. However, it is possible to add tiny Raman-active vibrational probes (e.g., alkynes and stable isotopes), which exhibits sensitivity, speciﬁcity, and biocompatibility for imaging of small biomolecules17.

In many in vivo biological and biomedical applications, simultaneous mapping of different chemical species in the same sample is extremely important for the investigation of the co-distribution or dynamic correlation between pairs of biomolecules18-19, as well as multicolor imaging with multiple chemical contrasts is considered necessary. By tuning the frequency of the pump beam or the Stokes beam in sequential scans, it is possible to image different Raman modes of the sample18. However, this wavelength-tuning approach may lose the co-localization information of different Raman modes, especially when the sample is in a dynamic environment18.

As a consequence of nonlinear excitation, SRS offers intrinsic 3D resolving capabilities of the selected chemical bond within biological samples20. Volume reconstruction of the selected chemical bond and its spatial distributions can be simply achieved by collecting SRS images at different focal plane along the z axis. Since the images are acquired with high spatial and temporal resolutions, we are also able to obtain several key information (such as 3D structure, chemical composition et cetera) about the biological sample.

ACKNOWLEDGMENTS:

We appreciate V. Tufano from IMM_CNR for his valuable technical assistance and Giacomo Cozzi, product specialist from Nikon Instruments, for useful discussions and continuous support. This work was partially supported by Italian National Operative Programs PONa3 00025 (BIOforIU) and by Euro-Bioimaging large-scale panEuropean research infrastructure project.

DISCLOSURES:

The authors declare that there is no conflict of interest.

REFERENCES:

- Saar, B. G., Freudiger, C. W., Reichman, J., Stanley, C. M., Holtom, G. R. & Xie, X. S. Video-Rate Molecular Imaging in Vivo with Stimulated Raman Scattering. Science 330 (6009), 1368–1370 (2010).

- D. Zhang, P. Wang, M. N. Slipchenko, and J.-X. Cheng, “Fast Vibrational Imaging of Single Cells and Tissues by Stimulated Raman Scattering Microscopy,” Acc. Chem. Res. 47(8), 2282–2290 (2014).

- A. Alfonso-García, R. Mittal, E. S. Lee, and E. O. Potma, “Biological imaging with coherent Raman scattering microscopy: a tutorial,” J. Biomed. Opt. 19(7), 071407 (2014).

- J.-X. Cheng and X. S. Xie, “Vibrational spectroscopic imaging of living systems: An emerging platform for biology and medicine,” Science 350(6264), aaa8870 (2015).

- C. H. Camp, Jr. and M. T. Cicerone, “Chemically sensitive bioimaging with coherent Raman scattering,” Nat. Photonics 9(5), 295–305 (2015).

- D’Arco, A., Brancati, N., Ferrara, M. A., Indolfi, M., Frucci, M. & Sirleto, L. Subcellular chemical and morphological analysis by stimulated Raman scattering microscopy and image analysis techniques. Biomedical Optics Express 7(5), 1853 (2016).

- Annalisa D'Arco, Maria Antonietta Ferrara, Maurizio Indolfi, Vitaliano Tufano and Luigi Sirleto. Label-free imaging of small lipid droplets by femtosecond-stimulated Raman scattering microscopy. Journal of Nonlinear Optical Physics & Materials 26(4) (2017)

- Ranjan, R., D’arco, A., Ferrara, M. A., Indolfi, M., Larobina, M. & Sirleto, L. Integration of stimulated Raman gain and stimulated Raman losses detection modes in a single nonlinear microscope. Optics Express 26(20), 26317 (2018).

- Reid, D. T., Sun, J., Lamour, T. P. & Ferreiro, T. I. Advances in ultrafast optical parametric oscillators. Laser Physics Letters 8, (1) 8–15 (2011).

- A. Zumbusch, W. Langbein, and P. Borri, “Nonlinear vibrational microscopy applied to lipid biology,” Prog. Lipid Res. 52(4), 615–632 (2013).

- M. Suzuki, Y. Shinohara, Y. Ohsaki and T. Fujimoto, Lipid droplets: Size matters, J. Electron Microsc 60(1), S101–S116 (2011).

- Rizzatti, V., Boschi, F., Pedrotti, M., Zoico, E., Sbarbati, A., & Zamboni, M. Lipid droplets characterization in adipocyte differentiated 3T3-L1 cells: size and optical density distribution, Eur. J. Histochem. 57(3), 159–162 (2013)

- A. Alfonso Garcia, S. G. P sterer, H. Riezman, E. Ikonen and E. O. Potma, D38-cholesterol as a Raman active probe for imaging intracellular cholesterol storage, J. Biomed. Opt. 21(6), 061003-1:8 (2016).

- Mukherjee, S., Zha, X., Tabas, I., & Maxfield, F. R., Cholesterol distribution in living cells: fluorescence imaging using dehydroergosterol as a fluorescent cholesterol analog. Biophysical journal, 75(4), 1915-25 (1998).

- S. Fukumoto and T. Fujimoto, Deformation of lipid droplets in fixed samples, Histochem. Cell Biol. 118(5) 423–428, (2002).

- A. Kinkel et al., Oil red-O stains non-adipogenic cells: A precautionary note, Cytotechnology 46(1) 49–56, (2004).

- L. Wei, F. Hu, Z. Chen, Y. Shen, L. Zhang, and W. Min, Live-Cell Bioorthogonal Chemical Imaging: Stimulated Raman Scattering Microscopy of Vibrational Probes, Acc. Chem. Res. 49(8), 1494–1502 (2016).

- Yasuyuki Ozeki , Takuya Asai, Jingwen Shou, and Hironobu Yoshimi. Multicolor Stimulated Raman Scattering Microscopy With Fast Wavelength-Tunable Yb Fiber Laser. IEEE Journal of Selected Topics in Quantum Electronics, 25(1), 1-11, (2019).

- Saar, B. G., Contreras-Rojas, L. R., Xie, X. S. & Guy, R. H. Imaging Drug Delivery to Skin with Stimulated Raman Scattering Microscopy. Molecular Pharmaceutics 8(3), (969–975 (2011).

- Chen, X., Zhang, C., et al. Volumetric chemical imaging by stimulated Raman projection microscopy and tomography. Nature Communications 8, 15117 (2017)