# Submission ID #: 59614

# Scriptwriter Name: Michael Linnes

# Project Page Link: http://www.jove.com/files_upload.php?src=18182318

# Title: Implementation of a Nonlinear Microscope Based on Stimulated Raman Scattering

# Authors and Affiliations: Rajeev Ranjan, Maria A. Ferrara, Maurizio Indolfi, and Luigi Sirleto

National Research Council, Institute for Microelectronics and Microsystems, Naples, Italy

# Corresponding Author:

Luigi Sirleto 		(luigi.sirleto@cnr.it)

# Email addresses for Co-authors:

Rajeev Ranjan 	(rajeev.ranjan@na.imm.cnr.it)

Maurizio Indolfi 	(maurizio.indolfi@na.imm.cnr.it)

Maria A. Ferrara 	(antonella.ferrara@na.imm.cnr.it)

Author Questionnaire:

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? (N)

2. Does your protocol include software usage? (Y)

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot.

2.1, 2.3, 3.5, 4.2, 4.3, 4.6

4. What is the single most difficult aspect of this procedure and what do you do to ensure success?

2.3 and 3.5

5. Will the filming need to take place in multiple locations? (N)

# Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

- REQUIRED Interview Statements:

# Why is your protocol significant?

  - Rajeev Ranjan: This protocol can help scientists interested in nonlinear microscopy to understand the key components, the setting up and the alignment procedure of a microscope based on Stimulated Raman Scattering.

# What is the main advantage of this technique?

  - Maurizio Indolfi: The main advantages of SRS microscopy are its abilities to perform label free imaging based on vibrational contrast and to acquire an image in a few seconds.

OPTIONAL Interview Statements:

# Are there any specific areas of research that this method could provide insight into?

  - Maria A. Ferrara: SRS microscopy has taken label-free imaging to new heights, especially in studies of complex biological structures such as lipids, which are fundamental to cells and cellular architecture.

# Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Luigi Sirleto: SRS signal is detected as a small change in the intensity of the probe’s beam and it can be corrupted by noise. Therefore, in order to be successful using this technique, an accurate design is crucial.

# Section - Protocol

- Alignment of the Microscope – Spatial Beam Overlap

  - Rajeev Ranjan:“The lasers used in this video have very high power, hence it is important to wear goggles for eye protection and to remove any metallic ornaments from your hands.”[1]

    - INTERVIEW: Author says above statement interview style.

  - To begin, align the OPO and the titanium-sapphire laser beams so that they both reach the microscope. [1-TXT]

    - WIDE: Talent works at the table with the laser setup, performing a laser adjustment. TEXT: OPO: Optical parametric oscillator laser (1076 nm); Ti:Sa (810 nm); Power: 10 mW.

  - Next, place the laser beam position sensors detectors in two positions in between dichroic mirror 1 and mirror 6, [1-TXT] the first position is located close to dichroic mirror 1 and the second one is close to mirror 6. [2] For each position, use the sensors to detect the x and y coordinates of the OPO beam [3].

    - Talent holds a sensor detector and points out DM1 and M6 TEXT: See Figure 1 for array schematic.

    - Talent points out the first and 2nd sensor position in the order described.

    - Videographer films the sensors from the left side of the optical bench.

  - Importantly, verify that the x and y coordinates of the titanium-sapphire laser beam are of the same OPO in both positions of the sensors detectors. [1] If in some positions the coordinates do not coincide, tune the tilt of the adjacent mirror to compensate. [2]

    - Talent uses a detector card and shows the beams both simultaneously and alternately.

    - Talent adjusts one of the mirrors.

  - Follow this same procedure to align the titanium-sapphire beam positions with respect to OPO for the path in between mirrors 6 and 7. [1-TXT]

    - Talent points out M6 and M7 TEXT: *Follow along in the text protocol to perform temporal synchronization of the beams.

- Alignment of the Microscope – Temporal synchronization of the beams - Use of the Autocorrelator

  - Install an additional mirror on a flip-flop mount in-between mirrors 6 and 7 and flip the mount of the mirror to direct the beam into the autocorrelator. [1]

    - Author points to the described mirror and flips it.

  - Power on the autocorrelator controller, start the software application on the computer controlling it, [1] and set the beam distance adjustment screw of the autocorrelator to its normal position at 8.35 millimeters. [2]

    - Talent turns on the autocorrelator controller, sits at computer and begins application.

    - Talent sets the beam distance adjustment screw to 8.35 mm.

  - Then, stop the titanium-sapphire beam and release and project the OPO beam from the additional mirror to the input mirror of the autocorrelator. [1] Try to adjust the input mirror to maximize the laser pulse signal as shown here. [2]

    - Talent adjusts the flip-flop mount and then the input mirror - Videographer: Film from the right side of the optical table.

    - LABMEDIA: Figure 5 Video Editor: Zoom into the left graph

  - Next, stop the OPO beam and release and project the titanium-sapphire beam from the flip-flop mounted mirror to the input mirror and into the autocorrelator. [1] Repeat the optimal beam adjustment until the autocorrelator signal shown here is obtained. [2]

    - Talent adjusts the flip-flop mount and then the input mirror - Videographer: Film from the right side of the optical table.

    - LABMEDIA: Figure 5 Video Editor: Zoom into the right graph.

  - Now, set the beam distance adjustment screw to the Cross position at 7.30 millimeters. [1] Release both beams and scan the delay line to overlap the two beams. [2] Obtain a resulting cross-correlator signal as shown here. [3]

    - Talent sets the beam distance adjustment screw to 7.3 mm.

    - Talent releases both beams and sits at computer and begins application.

    - LABMEDIA: Figure 6.

  - Then, flip the flip-flop mounted mirror so that the beams can reach mirror 7 and the scan head of the microscope. [1].

    - Talent flips the flip-flop mirror and shows the path of beams.

- Alignment of the Microscope – Aligning the Femtosecond Laser Beams: OPO and Ti:Sa.

  - Remove the condenser and use the escape button to temporarily retract the 60x objective lens. [1] Then, rotate the nosepiece to move the 60x objective lens off the optical path. [2]

    - Talent removes the condenser.

    - Talent presses escape and rotates the nosepiece.

  - Next, mount the detector to the upper part of microscope using the external mechanical mount. [1] Connect the detector’s output through a 50 Ohm low-pass filter to an oscilloscope. [2]

    - Talent mounts the detector as described.

    - Talent connects the detector through a low-pass filter and into an oscilloscope.

  - Now, turn on the processor that controls the scanner head and project the OPO beam into the scanner head of the microscope. [1]

    - Talent turn on the processor and project the OPO beam into the scanner head.

  - Monitor the OPO signal and Maximize the power measured by the detector using an x-y translator. [1]

    - Talent monitors the OPO signal and moves the x-y translator and shows the change of the signal to the oscilloscope.

  - Then, switch the beam from the OPO laser to the titanium-sapphire laser and verify that a signal is also obtained for the titanium-sapphire laser. This indicates that both beams are well-aligned. [1]

    - Talent switches the beams and measures Ti:Sa power.

  - Finalize the beam alignment by rotating back the nosepiece to introduce the 60x objective. Then, use the refocus button on the microscope to regain the finalized focus to the 60x microscope objective lens. [1]

    - Talent rotates the 60x objective back into place and shows the refocus button on the microscope.

  - Finally, place the objective with a magnification of 40x in place of the condenser without touching or disturbing the sample. [1]

    - Talent mounts the 40x objectives.

- Optimization of Temporal Synchronization of Beams

  - Set the power of the titanium-sapphire and OPO lasers measured before the microscope to 30 milliwatts for both beams. [1]

    - Talent measures the power of OPO.

  - Then, set the wavelength of OPO laser to a different value with respect to the previous one so that the pump and probe are not in resonance with the vibrational frequency of the beads. [1]

    - Talent sets wavelength of OPO

  - Next, release both beams so that they enter the microscope. [1] Run the scanning delay line computerized translator and record the measured intensity using the lock-in amplifier for each position of the delay line. [2] Wait until the delay line scanning is complete. [3]

    - Talent releases both beams and sits at computers and begins application.

    - LABMEDIA: Figure 7 (left).

  - Now, set the wavelength of the OPO back to 1076 nanometers so that pump and probe are in resonance with the vibrational frequency of the beads. [1]. Run the scanning delay line computerized translator and wait till the delay line scanning is complete. [2]

    - Talent sets wavelength of OPO.

    - LABMEDIA: Figure 7 (right).

  - Finally, set the obtained overlap beam position in the delay line for next acquisition of stimulated Raman scattering Images. [1]

    - Talent sets the obtained beam position.

- Optimization of Spatial Overlap of Beams -Transmission Image Acquisition of OPO and Ti:Sa

  - To optimize the spatial synchronization of the beams, begin by stopping the titanium-sapphire beam and reducing the OPO power to 8 milliwatts. [1] Next, connect the detector readout to the data acquisition card. [2]

    - Talent reduce powers of OPO beams.

    - Talent connects the detector to the DAQ.

  - Run the data acquisition program along with the microscope scanning console. When finished, save the file and process the data [1] to get an image like the one shown here. [2]

    - Talent seats at computers and begins application.

    - LABMEDIA: Figure 9a

  - Next, stop the OPO beam and reduce the titanium-sapphire power to 2.5 to 4.5 milliwatts. [1] Connect the detector with the lock-in amplifier and its readouts with the data acquisition card. [2]

    - Talent stop the beam and reduce power of Ti:Sa.

    - Talent connects the detector with the lock-in amplifier and its readouts with the data acquisition card.

  - Then, again run the data acquisition program along with the microscope scanning console. [1] When finished, save the file and process the data to get an image like the one shown here. [2]

    - MED Over the Shoulder: Talent begins application at computer.

    - LABMEDIA: Figure 9b

- SRS Image Acquisition

  - Introduce a stack of filters in-between the 40x objective and the photodiode to remove the pump pulses and acquire only the Stokes signal. [1]

    - Talent introduces filters.

  - Then, set the pump signal to 810 nanometers with a focused power of 8 milli-Watts and set the probe signal to 1076 nanometers with the same focused power of 8 milli-Watts to investigate a typical carbon-hydrogen bond of polystyrene with a Raman shift of 3054 inverse centimeters. [1]

    - Talent sets power for both beams in the order listed.

  - Connect the detector with the lock-in-amplifier and the lock-in-amplifier’s readout to the data acquisition card. [1]

    - Talent connects the detector to LIA and LIA readout to DAQ.

  - Finally, set the pixel format and acquisition time in microscope program and run it and the data acquisition system, saving the matrix file once it is complete. [1]

    - SCREEN: Screen capture_(figure_10)_SRS_Image_Acquisition.mp4.

# Section – Results

- Results: Nonlinear Microscope Based on Stimulated Raman Scattering

  - An example measurement from a single point of the sample is displayed here. [1]

    - LABMEDIA: Figure 7

  - When the beam is not overlapped in time or space, the obtained result is off-resonance where the amplitude of signal as measured by a lock-in-amplifier is zero. [1]

    - LABMEDIA: Figure 7 - Video Editor: Zoom into The top left graph.

  - The phase of this signal, however, jumps between negative and positive values. [1]

    - LABMEDIA: Figure 7 - Video Editor: Zoom out and back into the bottom left graph.

  - If the beams are overlapped in space, the signal increases and reaches its maximum when the beams are perfectly overlapped in time, [1] while the phase starts to achieve a fixed value during the time at which the beams are overlapped. [2]

    - LABMEDIA: Figure 7 - Video Editor: Zoom out and back into the top right graph.

    - LABMEDIA: Figure 7 - Video Editor: Zoom out and back into the bottom right graph.

  - In a transmission image, a single beam is used, and the transmitted beam intensity from the sample is measured by a photodiode. On the left, the transmission image was obtained using OPO while on the right the transmission image was obtained using titanium-sapphire. [1]

    - LABMEDIA: Figure 9 - Video Editor: Label the left image “TI - OPO Laser” and the right image “TI - Ti:Sa laser”

  - A typical example of an SRS image is shown here, in which label-free images of polystyrene beads with diameters of 3 µm are reported. [1]

    - LABMEDIA: Figure 10 - Video Editor: Zoom into figure 10.

# Section - Conclusion

- Conclusion Interview Statements:

# What is most important thing to remember when attempting this procedure?

  - Luigi Sirleto: In order to obtain a high quality image based on SRS microscopy, the alignment of microscope is critical. Therefore, all indicated steps in the protocol have to be carried out carefully.