Editorial Comments:

• Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammatical errors.

• Textual Overlap: Significant portions show significant overlap with previously published work. Please re-write lines 20-26, 99-102, 137-139, 414-419, 425-431, 440-445 to avoid this overlap.

Following the suggestion we amended the same in revised manuscript.

•Protocol Language:Please ensure that all text in the protocol section is written in the imperative voice/tense as if you are telling someone how to do the technique (i.e. “Do this”, “Measure that” etc.) Any text that cannot be written in the imperative tense may be added as a “Note”, however, notes should be used sparingly and actions should be described in the imperative tense wherever possible.

1) Examples NOT in imperative voice : 2.2.1, 2.2.2, 2.2.4, 2.3.1-2.3.8, 4.1,4.2, etc.

Following the suggestion we amended the same in revised manuscript

•Protocol Detail: Please note that your protocol will be used to generate the script for the video, and must contain everything that you would like shown in the video. Please add more specific details (e.g. button clicks for software actions, numerical values for settings, etc) to your protocol steps. There should be enough detail in each step to supplement the actions seen in the video so that viewers can easily replicate the protocol. Some examples:

1) 1.2: what is the humidity limit?                   See line 172

2) 1.11: Mention wavelengths tuned to.                See line 175 and 185

3) 2.1: Please describe in detail how this is done.

4) 2.2: Mention wavefrom shape, frequency, amplitude, pulse width etc. See line 191 and 192

5) 2.7: What is the proper voltage level?                     See line 199

Following the suggestion we added in the revised version of  paper all required details.

•Protocol Numbering: Please adjust the numbering of your protocol section to follow JoVE’s instructions for authors, 1. should be followed by 1.1. and then 1.1.1. if necessary and all steps should be lined up at the left margin with no indentations.

There must also be a one-line space between each protocol step.

Following the suggestion we amended the same in revised manuscript

•Protocol Highlight: Please highlight ~2.5 pages or less of text (which includes headings and spaces) in yellow, to identify which steps should be visualized to tell the most cohesive story of your protocol steps. Please see JoVE’s instructions for authors for more clarification. Remember that the non-highlighted protocol steps will remain in the manuscript and therefore will still be available to the reader.

1) The highlighting must include all relevant details that are required to perform the step. For example, if step 2.5 is highlighted for filming and the details of how to perform the step are given in steps 2.5.1 and 2.5.2, then the sub-steps where the details are provided must be included in the highlighting.

2) Some of your shorter protocol steps can be combined so that individual steps contain 2-3 actions and maximum of 4 sentences per step.

3) The highlighted steps should form a cohesive narrative, that is, there must be a logical flow from one highlighted step to the next.

4) Please highlight complete sentences (not parts of sentences). Include sub-headings and spaces when calculating the final highlighted length.

5) Notes cannot be filmed and should be excluded from highlighting.

6) Please bear in mind that software steps without a graphical user interface/calculations/ command line scripting cannot be filmed.

• Discussion: JoVE articles are focused on the methods and the protocol, thus the discussion should be similarly focused. Please ensure that the discussion covers the following in detail and in paragraph form (3-6 paragraphs): 1) modifications and troubleshooting, 2) limitations of the technique, 3) significance with respect to existing methods, 4) future applications and 5) critical steps within the protocol.

In order to comply with these requirements in the discussion of revised version we added two new paragraphs (line 442-459).

• Figures: Please remove the embedded figures from the manuscript. Figure legends, however, should remain within the manuscript text, directly below the Representative Results text.

• Figure/Table Legends: Please expand the legends to adequately describe the figures/tables. Each figure or table must have an accompanying legend including a short title, followed by a short description of each panel and/or a general description.

• Commercial Language:JoVE is unable to publish manuscripts containing commercial sounding language, including trademark or registered trademark symbols (TM/R) and the mention of company brand names before an instrument or reagent. Examples of commercial sounding language in your manuscript are Eclipse TE-2000-E, Nikon, C2, Nikon, EOM CONOPTICS 3350-160 KD*P, Chameleon Ultra II, (SOPO-Chameleon Compact OPO, Semrock FF875-Di01-2536, Newport MOD MILS200CC, (Coherent Chameleon Ultra II, Coherent Compact OPO, Thorlabs VRC5, PowerMax PM10, Coherent, Electro-Optic, TG5011 AIM – TTi, (SR844, Thorlabs PT1/M,

- Please use MS Word’s find function (Ctrl+F), to locate and replace all commercial sounding language in your manuscript with generic names that are not company-specific. All commercial products should be sufficiently referenced in the table of materials/reagents. You may use the generic term followed by “(see table of materials)” to draw the readers’ attention to specific commercial names.

• Table of Materials:Please revise the table of the essential supplies, reagents, and equipment. The table should include the name, company, and catalog number of all relevant materials/software in separate columns in an xls/xlsx file. Please include all instruments used etc.

• If your figures and tables are original and not published previously or you have already obtained figure permissions, please ignore this comment. If you are re-using figures from a previous publication, you must obtain explicit permission to re-use the figure from the previous publisher (this can be in the form of a letter from an editor or a link to the editorial policies that allows you to re-publish the figure). Please upload the text of the re-print permission (may be copied and pasted from an email/website) as a Word document to the Editorial Manager site in the "Supplemental files (as requested by JoVE)" section. Please also cite the figure appropriately in the figure legend, i.e. "This figure has been modified from [citation]."

Reply to Reviewers:

Reviewers' comments:

Reviewer #1:

Manuscript Summary:

The manuscript reports the implementation of a stimulated Raman scattering (SRS) microscope obtained by the integration of an SRS spectroscopy set-up with a laser scanning microscope. This SRS microscope is based on two femtosecond (fs) laser sources: a Ti-Sapphire (Ti:Sa) and a synchronized optical parametric oscillator (OPO).

The developed system appears to be very important for bio chemical application and of course for the research in Material science.

It is worth noting that SRS microscopes are not commercially available, so in order to take advantage of their characteristics, the home-made realization is the only possible option.

Paper sounds well, the experimental part is clearly reported and discussed allowing to develop the system in other labs.

I have highly appreciated this work.

We sincerely thank the reviewer and appreciate for time you spent reviewing our manuscript. We also like to thanks for the positive comments and appreciation of our work.

Reviewer #2:

Manuscript Summary: The manuscript submitted by Ranjan et al. presents the implementation of a stimulated Raman scattering (SRS) microscopy by combining SRS spectroscopy with a laser scanning microscope system (LSM). Since no commercial SRS microscope system is available at the moment, if efficiently presented, this manuscript could bring a detailed guide of how to set up such a system on an existing LSM, and become useful to the imaging community.

We thank the esteem reviewer for his appreciation of our work.

However, I found many major weaknesses of this manuscript.

We would like to thanks reviewer for his constructive suggestions. We highly acknowledge the reviewer for highlighting the major and minor concerns about the manuscript. We are replying the raised issue point by point.

Major Concerns:

Overall, I found the writing requires much rework, and ask JOVE to require a major revision from the authors, before considering it for publication.

In details, I find many sentences and paragraphs need to be better arranged and organized. Also the English should be used with more care and in the professional style common in science community. References are not used effectively. Reference 1-9 are all together used in the “Discussion” while no reference was used in the “Introduction”.

Following the comments, we have thoroughly revised the English of our manuscript. Additionally, we have included more reference in the introduction section of the manuscript.

- Line 19: What is the “problem of other coherent Raman techniques,”? Name a few.

Coherent Antistokes Raman scattering suffers from nonresonant background. Due to a nonresonant background, CARS spectrum is different from the corresponding spontaneous Raman spectrum. This background complicates spectral assignment; it causes difficulties in image interpretation and limits detection sensitivity[7]

In the revised version we substitute ‘other coherent Raman techniques’ with ‘coherent antistokes Raman scattering’ (see line 18).

- Line 22 and Line 33 stated about the availability of a commercial system. Since there isn’t any yet available, state that and indicate the significance of this work.

We rewrote the abstract, according to the suggestion of reviewer (see line 34-42).

- From Line 40-141 (Introduction): The “Introduction” is the part I find to need most revision. In the introduction, the readers would expect a certain level of introduction of the field, the history of the technical developments, the current status and limitations of the systems available, what the authors achieved (improvements/breakthroughs) or how the authors overcame the limitations, and the significance of the work (academically, financially or etc.). However, the “Introduction” of this manuscript barely state any of these aspects. I indeed found the earlier paragraphs in the “Discussion” (Line 409-424) are more appropriate to be placed in the “Introduction”. Without such content, the current “Introduction” looks more likely a method section of a paper. I understand that these descriptions are necessary, and due to the format of JOVE, they are too long to be put in “Protocol section.” Thus if it is inevitable to put the technical details in the “Introduction,” they have to be better arranged in a more logical flow, with a smooth transition from the regular introduction content (Line 409-424).

We move the line 408-424 to the beginning of introduction according to the suggestion of reviewer ( see line 44-57).

- Line 44-46: The first laser stated is the Ti:Sa (pump) laser, but it is stated as “Beam2” in Fig. 1. Try to state the two lasers in the same order as in the Figure.

We have revised and reproduced the figure accordingly.

- Line 53-59 summarizes three issues of the SRS microscopy, and the following three paragraphs give detailed descriptions of those issues (Line 60-88). This arrangement is not indicated or stated anywhere, so the readers would get confused initially, also because Line 53-59 state elsewhere of the manuscript (2.1~2.3). For a better flow, indicate this way of flow by inserting “(and will be described in the following paragraphs)” at the end of Line 52.

We inserted the suggest sentence suggested by reviewer at the end of line 68 of revised version.

- Line 52: “some important issues” -> avoid using words like “some.” State exactly what these issues are in a proper category.

In the revised paper we substitute ‘some important issue’ with ‘three crucial issues’ (see line 68).

- On the other hands, Line 53-59 do not deliver the essential idea of each issue. For example, Line 55-56 stated about the mechanical arrangement, but skips why this became an issue (which was stated later in Line 70-75, the sensitivity and noise problems). The summary of each issue should be concise and compressive at the same time. Put lines between paragraphs. Currently it is hard to identify different paragraphs.

We have inserted a part of line 70-75 in the summary of the issue according the suggestion of reviewer.

- Start the three paragraphs (Line 60-88) by restating the issues stated in Line 53-59, so that it would read clearly that these paragraphs are describing the issues in details. For example, Line 76 should start with a sentence including the word, “synchronization” Cite the proper figures in any paragraph describing the system details.

We have rearranged the revised version of paper according to the suggestions of reviewer (see line 91-93)

- Line 62: the “Therefore” seems not right. To deliver the message logically, the sentence should be like “We overcame this noise issue by implementing a high frequency modulation…”

Following the suggestion, we amended the same in revised manuscript (see line 74)

- Line 63: “According to” seems not right. “this process” should mean the previously stated “high-frequency modulation transfer method, but this method was never explained in detail. Therefore, instead of "According to this process,” “In this method” should be used.

Following the suggestion, we amended the same in revised manuscript (see line 75)

- Line 72: “white light” -> does this mean the brightfield imaging?

Yes. We substitute white field for brightfield into the text (see line 88).

- Line 82: What is “Nikon controller”? According to Fig. 3 it should be what is called “Microscope controller” but no detail of what it is described.

We substitute “Nikon controller” with ‘microscope controller connected to the scan head unit’ (see line 101).

- Line 95-125: These are the details of the four steps stated in Line 91-94. Again, the readers should be guided properly for those paragraphs. Put “(and will be described in the following paragraphs)” at the end of Line 90. Also start each paragraph of Line 95-125 by restating the according step stated in Line 91-94.

We inserted ‘’will be described in the following paragraphs” (see line 108 of revised version).

- Line 99: “Recently, OPO technology has allowed perfect synchronization…” -> Can a reference be placed here?

Yes, we added the reference 9 in revised paper.

D.T. Reid, J. Sun, T.P. Lamour, and T.I. Ferreiro. Advances in ultrafast optical parametric oscillatorsLaser Phys. Lett., 1–8 (2010) / DOI 10.1002/lapl.201010085

- Line 101: “5 cm” -> is this correct? This looks too big spatial difference.

Yes

- Line 117-118: “To realize the image… software is used” -> This is not relevant to the “steps” topic (Line 93). This sentence can be placed in a separate paragraph in the end of “Introduction”.

We revised it and address the issue (see line 164-165).

- Line 126: Avoid a value word like “some”. In fact, these tests are stated in Line 128-140, but this flow is not well indicated. Thus, state it by "we carry out the tests described below, …"

Following the suggestion, we amended the same in revised manuscript (see line 148).

- Line 133: "The only difference" -> What difference to what? The sentence including this phrase is describing the difference between the procedures of obtaining transmission images between OPO and Ti:Sa, but this sentence can be very confusing. Restate this in a clear way.

We revised it and address the issue (see line 151-155).

- Line 185 as well as Figure 3: Figure 3 is too simple, and the visualization of section 2.3 is very weak by Figure 3. Either by putting more words or graphical descriptions in Figure 3, support the idea of 2.3. Then describe each component of Figure 3 (especially the "Detection System" and "Microscope controller" well in the description (from Line 188-206). Currently it is not very clear which parts are describing what. "Note" can be used to divide parts in Line 188-206, so that readers would know what is described in Fig. 3.

We note that Figure 3 describes the hardware of synchronization system, while paragraph 2.3 describe the basic idea of our ‘in house software’ realized to obtain the synchronization. In the revised version of figure 3.a, we add graphical descriptions and new words and also change the name of paragraph 2.3. The new one is Synchronization software.

- Line 190, 192 and 194: Put quotation marks around the TTL signal, so “frame sync”, “pixel clock” and “line sync”

Following the suggestion, we amended the same in revised manuscript (see line 100-101).

- Line 192-193: This sentence is very confusing. Make it simple and clear.

We substitute the sentence with a new one (The basic idea of Syncronization software is the following: the frame sync controls the start acquisition in post trigger synchronism mode, line sync manages the memory area and the image realization, while the pixel clock is used as SRS signal acquisition sampling rate. A dedicate algorithm in LabVIEW12 has been realized in order to store data. Here, we describe the main steps), (see line 208-212).

- Line 208, 220 and many other places: State “Figure 1” in the note or steps. The readers should know what Figures to look at when reading the steps.

Thanks for suggestion. We have amended the manuscript accordingly.

- Line 221: The title is strange. It should be like “Procedure using fast photodiode plus an oscilloscope”

We have implemented this change in the subtopic (see line 248).

- Line 238: Put the “AM” and “autocorrelator” in the Fig. 1

We have reproduced the figure 1 following suggestion.

- Line 242-263: No visualization is given. Video can include this part, or photos can be placed in the manuscript.

We have already Figure 4a (Line 279), 4b (Line 283) and figure 5 (287) which are dedicated to this section.

- Line 242: What is “ACF”? It was never introduced.

We remove it from the text

- Line 244: What is “APE application”? APE is a company, so either state the software name, or just state it like software #1 and put the details in the materials list. Same for "pulseLink."

APE application is the support and control software for the Autocorrelator.

- Line 246, 257 and many other: Avoid using the word, “now”

We revised the manuscript following the suggestion.

- Line 262 and several other places in steps: this should be written in the imperative tense.

We revised the manuscript following the suggestion

- Line 271: What manual? The microscope manual?

Yes.

- Line 276: Avoid the word “you” but instead state like “FOV of interest”

Thanks for suggestion we amended it (see line 302).

- Line 282-303: Visualization is very weak. Adding some photos will help.

We additionally added the picture of the upper section of microscope where the photodiode is mounted with mechanical mount (see figure 3a).

- Line 311: What to do with Figure 7(a) and (b)? Repeat the work until (b) is obtained?

We rewrote the paragraph 3.4.1. and we repet the work until (b) is obtained (see line 331-345).

- Line 349-360: SRS images of beads are obtained, but what is the useful and additional information we can extract out of SRS image of beads compared to the conventional brightfield or epi images? Also can there be a biological sample images obtained by this SRS microscope?

Being  the beads, a  homogeneous sample, we don’t expect additional information we can extract out of SRS image compared to the conventional brightfield or epi images. Concerning biological sample, in order to elucidate this point we added two new paragraphs in discussion (line 442-459).

- Figure 1 & Figure 2: I suggest combining Figure 1 and 2 since they are very much related and can help readers to understand if combined. Color code should be consistent. Figure 1 put green color for the Beam 2 (pulse laser) but Figure 2 uses blue color. Also indicate each objective differently (like Obj1 and Obj2), since Figure 1 and 2 have objective in the opposite orientation.

In the revised version of figures 1 and 2, color code are consistent and focused and collected objective are indicated in the same way (Obj1 and Obj2). However, we don’t prefer to combine them.

- Figure 2: indicate where the “T-junction” is.

We have indicated in figure 2 where the T-jucntion is.

- Figure 3: put more details as stated before. Also the arrows and figures should be better aligned (also for Figure 2).

Following the suggestion we amended the same in revised manuscript.

- Line 409-424: as stated before. Those should be moved into the “Introduction”.

In revised our manuscript we move line 409-424 at the beginning of ‘Introduction’.

- Line 409: “is a great opportunity” is a too vague statement. State the advantages specifically.

We revised our manuscript and we introdude the sentence , SRS microscopy of biological samples is one of the most important player for future challenges of label free imaging.

- Line 410-413: The sentence is too long.

Following the suggestion, we have shortened it (see line 44-47).

- Line 425-438: Each paragraph should deliver a single idea, but this one has multiple stories. For example, Line 430-432 describes a limitation of SRS microscopy, but it doesn’t serve any role. Try to make each paragraph to deliver a single idea, and if otherwise, separate it into several.

We have amended it.

- Line 450: “very expensive” give a rough estimate of the cost.

Roughly estimate figure is about € 400 KEURO

- Line 446-447: This is the main selling point of this manuscript, so it should be put into the “Introduction” as well as “Abstract”

We rewrote the abstract section of the manuscript according to the suggestion of reviewer (see line 34-42).

- Line 451-452: This sentence should be written in a more professional way, like "becoming popular in varying fields."

We rewritten the sentence (sse line 40).

- Line 455: delete "really"

Deleted.

Reviewer #3:

Manuscript Summary:

The manuscript provides a good amount of detail about the construction and implementation of a stimulated Raman scattering microscope system based on a Nikon microscope frame. Methods in beam alignment and signal acquisition conditioning are discussed wherein an SRS image of polystyrene is lastly demonstrated.

We thank reviewer for the constructive inputs about our manuscript.

Major Concerns:

1. The article's abstract describe the construction of an SRS spectroscopy setup however it was not clear that spectrometers were used or any kind of spectroscopy was performed. The pump and Stokes beams were described to have pulse durations of 140 fs and 200 fs, respectively. I presume that the bandwidths of the laser beams span several nm or wavenumbers and tuning the wavelengths of the lasers at 809 nm and 1073 nm not only excites the 3054 cm−1 but more likely a large span of the CH vibrational region. Without any sort of dispersion mechanism for the two beams, this setup prevents vibrational spectroscopy with a reasonable value for spectral resolution. Is there by any chance a dispersion mechanism is inserted somewhere in the setup, if so, what would be the spectral resolution of the system? Otherwise, the use of the term “"spectroscopy”, should be refrained from the manuscript as it may confuse readers.

According to the suggestion of  reviewer, we refrained from the manuscript the term ‘spectroscopy’

Minor Concerns:

The article is written with great detail and may be very helpful for a lab with the aim of building an SRS microscope or even something similar. However, there are numerable inconsistencies and vagueness in the article draft which are listed below:

1. line 65 mentions PD which I assume stands for photodiode. Please include manufacturer information and model.

Yes, PD stands for Photodiode. Acronyms PD is also defined in line 31 of revised version. The manufacture is Thorlabs and the model is DET10N/M. Following the guidelines and policies of JOVE we can’t add the commercial details of the products, but we have described all the details in material list.

2. line 74: in order to avoid small oscillation of detection system --> does this mean mechanical vibration?

Yes. This mechanical vibration may generate from the mechanical mounting. In the revised paper we substitute small oscillation with mechanical vibration (sww line 82).

3. Please correct typographical errors in optic elements:

3.1 Line 96: Typo in Semrock FF875-Di01-2536, should be FF875-Di01-25x36.

3.2 Line 103: Typo in Newport MOD MILS200CC, should be Newport M-ILS200CC

We have corrected and added the commercial information(s) of all the products from the manuscript following the guidelines and policy of JOVE in the material list.

4. Ti:Sa and OPO are often mentioned in the article and they clearly mean pump and Stokes. Please revise as suited.

We have revised the same as suited.

5. Line 192: The pixel clock adjusts the reading by the buffer of the voltage values to be referred as pixel, acting as a sample clock ---> Confusing. Please rephrase.

We have rearranged it (see line 217-218 of revised version).

6. “the peak overlap of two beams is obtained with an approximation of 10 ps” ---> Seems vague, what does approximation of 10 ps mean? Is this device resolution?

It is the uncertainty of our measurements (we specify inside the text see line 126)

7. Line 213: I could not find a Newport SI-PSD9, perhaps this is the Newport CONEX-PSD9? If so, please revise.

Yes, it is CONEX-PSD9. We have added the commercial information(s) of all the products from the manuscript following the guidelines and policy of JOVE in the material list.

8. Line 213 to Line 217: The pump beam (Ti:Sa) does not seem to pass through M1 so the placement of laser beam position detector at M1 would only be useful in detecting the Stokes beam. How would the placement of the sensor at M1 allow the detection of the pump beam?

Following the suggestion, we amended the same in revised manuscript. We rewrote the 3.1 paragraph (see lines 239-249).

9. Line 242: The abbreviation ACF was not previously mentioned, please clarify what ACF means.

ACF indicates the autocorrelator signal. We remove the abbreviation from the text.

10. Line 283: Typo in "Photodiode DET10A". I believe this should be Thorlabs DET10A.

Thanks for the figuring out the typo. We have corrected it. We have added the commercial information(s) of all the products from the manuscript following the guidelines and policy of JOVE in the material list

11. In sections 3.2b and 3.3b, laser beams are incident on detectors (autocorrelator and photodiode) and the use of large powers may saturate or damage these detectors. A short mention of the laser powers may provide additional guidance to the readers.

We have mentioned the laser power in revised manuscript (see line 335 and 377)

12. Line 349: An example of SRS spectroscopy measurement --> remove "spectroscopy" if no spectroscopy was actually demonstrated.

We remove ‘spectroscopy’ in revised version.

13. Figure 1: Since there are already a lot of details included, a placement of optical filters may somehow add usefulness to the figure.

We added optical filter into the figure1.

14. Figure 9 shows a streaky and noisy image. From your research group's other publications, e.g. D'Arco et al. Biomed. Opt. Express 7, 1853-1864 (2016), Ranjan et al. Opt. Express 26, 26317-26326 (2018), shown are higher quality images of polystyrene beads. Perhaps using a similar image with better quality may improve the impact of this paper to future readers. If possible, try to explain why streaks appeared on the first image vs. the new. Explain steps to minimize image streaks/noise.

In submitted version of paper, we inserted a rough image (without using image processing technique), while in reference cited by the reviewer an image processing was used. In revised version we improve the quality of image, using a simple filtering process.

15. Line 440 - 441: distinguish a target biomolecule from the sea of the other related species... I assume "sea" is a typo, please revise.

No sea is not a typo. It is right. This paragraph describes the difficulty in identifying a particular type of biological species from the large amount of species which share same chemical bonds.

16. The manuscript seems to mention some jargon on the image acquisition details (sec. 2.3). I would suggest, though not heavily required, about adding a figure related to the image acquisition system in order to clarify concepts and terms.

We have revised figure 3 in order to clarify concept a terms

17. Reference #3 is about CARS and not SRS. Adding this to the list of references in relation to its citing paragraph (line 424) does not seem to be appropriate. Please revise accordingly.

We have removed reference 3 and updated our list of references.