# Submission ID #: 59689

# Scriptwriter Name: Melissa Ceo

# Project Page Link: http://www.jove.com/files_upload.php?src=18205193

Title: A Two-Step Pyrolysis-Gas Chromatography Method with Mass Spectrometric Detection for Identification of Tattoo Ink Ingredients and Counterfeit Products

# Authors and Affiliations: Ines Schreiver, Christoph Hutzler, Andreas Luch

Department of Chemical and Product Safety, German Federal Institute for Risk Assessment (BfR)

# Corresponding Author:

# Ines Schreiver

# Email: Ines.Schreiver@bfr.bund.de

# Email addresses for Co-authors:

# Christoph.Hutzler@bfr.bund.de

# Andreas.Luch@bfr.bund.de

# PLEASE READ THE INSTRUCTIONS IN THE GRAY BOXES CAREFULLY AND USE TRACK CHANGES WHILE MAKING ANY EDITS TO THE DOCUMENT.

# This document has several sections on separate pages, so take care to view each page.

Author Questionnaire:

Authors, please fill out the unanswered questions below.

1. Microscopy: Does your protocol involve video microscopy? N

2. Does your protocol include software usage? Y

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot. (You do not need to include steps that will be screen captured. Please do not list entire sections.)

2.1, 2.2, 2.5, 2.6

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1-2 individual steps using the step numbers listed in this document. (Please do not list entire sections.)

Authors, please answer this question with the steps listed here in the Protocol section below for use by the videographer.

5. Will the filming need to take place in multiple locations? N

Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

Authors, these headshots will be used for the JoVE Dedicated Author Webpage. Here is one example if you wish to take a look.

- REQUIRED Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

- The total introduction length (i.e., Required and Optional Interview Statements) cannot exceed 150 words.

- Restrict the length of each statement to no more than 30 words.

- Please answer the questions below in full sentences to highlight the significance of your protocol. You will be expected to memorize and deliver these sentences as spoken interview statements during filming.

- Indicate the full name of each author who will give each statement. If only one author is giving the REQUIRED statements, the same author may speak both statements.

# Why is your protocol significant? OR What key questions can this method help answer?

  - Andreas Luch: Tattoo as well as other inks are complex mixtures of ingredients. Our method provides a quick and easy way to analyze multiple components at the same time.

# What is the main advantage of this technique?

  - Christoph Hutzler: Pyrolysis online coupled to gas chromatography with mass spectrometric detection allows the analysis of volatile and non-volatile compounds in two consecutive runs.

OPTIONAL Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

- The following OPTIONAL questions may be answered to provide additional introductory information about your protocol.

- These OPTIONAL statements must be spoken by different authors than those who gave the Required Interview Statements.

- The length of each OPTIONAL statement is restricted to no more than 30 words and contributes to the total introduction length, which cannot exceed 150 words.

- Indicate the full name of each author who will give each OPTIONAL statement.

# Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Ines Schreiver: In addition to pure analysis of ink ingredients, the method can also be used to differentiate original and counterfeit products.

# Are there any specific areas of research that this method could provide insight into?

# Can this method be applied to any other systems?

  - Christoph Hutzler: In general, the method can be used for the analysis of all kinds of liquid samples. The data evaluation approach may also be applied to any other pyrolysis data.

# How would you expect an individual who has never performed this technique to struggle?

# Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Ines Schreiver: As with most analytical techniques, data analysis of pyrograms is the Achilles heel of the method. Therefore, we provided libraries with pyrograms and single analyte spectra to ease interpretation of the data.

# Why is visual demonstration of this method critical?

  - Author Name: ___________(Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

# Introduction of Demonstrator: (Said by you on camera)

- Please use this statement ONLY if any of the individuals who will be demonstrating the procedure on camera have not given a required or optional Introduction interview statement already.

- Include the full name(s) of the person(s) demonstrating the experiment followed by their title (e.g., technician, post doc, grad student, clinician, etc.)

- Also indicate the full name of the author who will introduce the demonstrator(s).

  - Author Name: Demonstrating the procedure will be _________ (name of the person or persons), a _________ (technician, post doc, grad student) from my laboratory. (Add additional mention of demonstrators as necessary).

    - Interview style: Author saying the above

    - The named technician, post doc, student looks up from workbench or desk or microscope and acknowledges the camera.

Section - Protocol

- Read through the entire protocol carefully to understand what you will need on the filming day and prepare accordingly.

- The two-digit numbers (e.g. 2.1., 2.2.) represent the “steps” of you protocol and will be read by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the “shots” that our videographer will capture at your lab.

- To ensure that your protocol can be filmed in a single work day, the protocol is restricted to 30 steps and/or 60 shots.

- It is critical for a smooth and organized shoot that all materials and work spaces are prepared and labeled (if applicable) in advance.

- Any specimens/samples that require long or overnight incubation steps should be prepared in advance. (i.e. day 0 sample preparation will be filmed on the day of the shoot; day 1 samples should be prepared the day before the shoot so their processing can be filmed on the day of the shoot/after their overnight culture/treatment/etc.)

- Tattoo Ink Preparation and Sample Mounting

  - First, hold a 25 millimeter hollow glass pyrolysis tube with specialized tweezers and insert the necessary amount of quartz wool into the tube with pointed tweezers [1]. Authors, what is the necessary amount of quartz wool? The wool is not weight or else. It is a bit of wool inserted and is estimated through operating experience. It will be very clear upon filming 2.1 and 2.2.

    - MED: Talent picks up a glass pyrolysis tube with the tweezers and CU: inserts quartz wool into the tube with tweezers.

  - Insert two steel sticks at each side of the pyrolysis tube and compress the wool into a 1 to 2 millimeter thick stopper [1]. The stopper must be positioned at the lower third of the pyrolysis tube to achieve adequate heating during the pyrolysis process [2].

    - MED: Talent inserts the steel sticks into the pyrolysis tube and compresses the wool in the tube.

    - CU: Close-up of pyrolysis tube to show position of stopper after compressing the wool.

  - Ignite a gas burner and bake the pyrolysis tube and filling for 2 to 3 seconds from each side to remove contaminants [1].

    - MED: Talent ignites a gas burner and bakes the pyrolysis tube.

  - Next, shake the tattoo ink bottles vigorously for 1 minute by hand to ensure homogeneity [1]. Additionally, they can be placed in an ultrasonic water bath for 1 minute [2-TXT]. Authors, will the inks be mixed manually and using the ultrasonic bath OR just manually mixed in the video?Just manually in the video.

    - MED: Talent shakes the tattoo ink bottles.

    - MED: Talent places the tattoo ink bottles in an ultrasonic bath and turns the bath on. TEXT: Some inks require prolonged homogenization.

  - Dip a 2 microliter microcapillary tip in the ink and aspirate about 1 microliter of ink by filling half of the capillary [1].

    - CU: Close-up shot of the ink bottle as talent dips a microcapillary tip in it and fills the capillary with the ink.

  - Insert the capillary into the pyrolysis tube and stain the quartz wool stopper with the ink [1]. A clear color staining must be visible without adding too much ink to the sample [2].

    - MED: Talent inserts the capillary into the pyrolysis tube and stains the quartz wool stopper with the ink.

    - CU: Close-up shot of the pyrolysis tube to show the stained quartz wool stopper.

  - Using the specialized tweezers, attach the prepared pyrolysis tube to a steel transport adapter for the automated injection unit [1].

    - MED: Talent picks up the pyrolysis tube with the tweezers and attaches it to a steel transport adapter.

  - Check that the pyrolysis tube is perfectly vertical and does not fall off during shaking [1].

    - CU: Close-up shot of pyrolysis tube in steel transport adapter as talent checks it position.

  - Place the transport adapter in the tray of the automated injection unit at the desired position for pyrolysis-GC-MS analysis [1].

    - MED: Talent places the transport adapter in the tray of the automated injection unit.

- Data Evaluation Approaches

  - For data evaluation of volatile compounds, start the GC-MS analysis/MS library searching software and open the chromatogram of the desorption run [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Select commercial libraries by clicking on Spectrum and Select Library. Then, load the library of interest [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Select integration parameters and perform a library search by clicking on Spectrum and Library Search Report.

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - For pyrolysis data evaluation, mark the whole chromatogram in the GC-MS evaluation software with the right mouse button pressed down to obtain an average mass spectra, or AMS [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - In order to establish a self-made library, click on Spectrum, Edit Library select the library of interest followed by Add New Entry and fill in all information of interest [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Generate an AMS of the investigated ink pyrogram and use the library search for comparison to the self-made AMS library. Exclude masses from column bleed or other column noises [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - To identify non-volatile compounds with specialized pyrogram evaluation software, build a folder on your computer with all pyrogram entries that should serve as a library, for example pigment pyrogram library for pigment identification or pyrograms of an original ink to compare it to putative counterfeit products [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Load the unknown pyrogram in the tab Library Search by clicking on Browse [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Next, load the library folder and select only MS matching and RT matching in the Search options, since the overall abundance will vary compared to pyrogram of reference pigments [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Click on advanced in the Search options. Select an RT window width of interest, Area Threshold of 0.1% and allow multiple matching. In the Mass Spec Options, select the parameter use only peaks with specified MS spectra and use a fit threshold of 850 [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Load the file “Pigment_search_spectra.spf” or click on Add to save specified MS spectra from each pyrogram of reference pigments or polymers from the library in the advanced search options [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page. TEXT: 3–5 spectra of most abundant peaks. Video Editor: Overlay should appear at mention of “specified MS spectra”.

  - Press OK to return to the main window. Then, click on Search to start the comparison [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - If needed, go to Chromatogram Match tab, select a compound, click right followed by Send to NIST to forward the spectra to the MS library software and identify the compound [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - For manual data evaluation for non-volatile compounds, start the GC-MS analysis/MS library searching software and open the chromatogram of the pyrolysis run [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Select commercial and pyrolysis libraries by clicking on Spectrum and Select Library. Load all libraries of interest [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Integrate the pyrogram in the GC-MS evaluation software and consider all peaks with an area not less than 0.2 percent of the total area [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Next, start the library search by clicking on Spectrum and Library Search Report [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page.

  - Manually compare all library matches to specific pigment and polymer decomposition products in the text protocol or fragments stated in literature [1].

    - SCREEN: To be provided by the authors. Authors, please upload this screen capture to your project page. TEXT: See Supplementary Table 1. Video Editor: Overlay should appear at mention of “the text protocol”.

OPTIONAL – Critical Step Statement:

- An OPTIONAL brief statement may be submitted for further elaboration of the best way to perform the required technique for the single most critical step of this procedure.

- If there is no single critical step, then there is no need to fill out this statement.

- This will be an interview style shot interjected after the relevant step within the Protocol section of the video.

- This statement is limited to 30 words or less.

- Please indicate the full name of the Author who will give this statement and the step of the protocol to which the statement pertains using the step numbers from the Protocol section (above).

# Fill in the details below based on the instructions above for the “Critical Step Statement”

# Author name, Step            :           (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

Section – Results

# The Results section is restricted to 200 words of narrative. Please read through the results as presented to make sure that it accurately represents your findings. If you would like to highlight other data, please revise this section accordingly, keeping in mind the word count restriction. Please note that we cannot include narrative without an accompanying visual.

- Results: Analysis of Multi-component Tattoo Inks by Two-Step Pyrolysis-Gas Chromatography Mass Spectrometry

# Authors, please feel free to change the results title if you had something more fitting in mind.

  - Well produced inks with highly pure ingredients and a limited number of components result in chromatograms easy to interpret with standard libraries, since most peaks can be identified [1]. But even in the high-quality inks, non-declared ingredients such as propylene glycol are often found in addition to the declared glycerol [2].

    - LAB MEDIA: Figure 2

    - LAB MEDIA: Figure 3 – Video Editor: Show figure A and highlight propylene glycol and glycerol peaks when mentioned.

  - Inks containing multiple ingredients and impurities will result in a pyrogram that is difficult to interpret [1]. Most peaks occurring in the second run may not be baseline separated from each other, making identification difficult [2]. Some substances might also result in peaks below the threshold set during data evaluation [3].

    - LAB MEDIA: Figure 3 – Video Editor: Show both figures A and B and highlight figure B.

    - LAB MEDIA: Figure 3 – Video Editor: Zoom into figure B and highlight overlapping peaks under 1 and 10/11.

    - LAB MEDIA: Figure 3 – Video Editor: Continue showing figure B and highlight small peaks in spectra between 17 and 32 minutes.

  - A solution to interpreting complex data might be a step-wise approach using 400, 600, and 800 degrees Celsius in consecutive pyrolysis steps for the same sample [1].

    - LAB MEDIA: Figure 4 – Video Editor: Highlight each spectrum when 400, 600 and 800 degrees Celsius is mentioned in the voiceover.

  - Some pigment decomposition products may descend from multiple pigments [1].

    - LAB MEDIA: Supplementary Table 1 – Video Editor: Show as much of the table as possible.

  - A positive result for counterfeit product identification is demonstrated for three “lemon inks” purchased from different vendors [1]. Using a forward match factor above 0.9, the chromatogram from the first desorption run and the pyrogram from the second run of the original ink were compared against three independent acquisitions of the original ink and the two counterfeit products using pyrogram evaluation software [2].

    - LAB MEDIA: Figure 5 – Video Editor: Show the spectra only.

    - LAB MEDIA: Figure 5 – Video Editor: Highlight figure A at mention of “the chromatogram from the first desorption run and the pyrogram from the second run of the original ink” and highlight figures B and C at mention of “the two counterfeit products”.

Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

- Below are questions for statements that can be used to further emphasize the significance of your protocol. At least one statement is required.

- Each statement is limited to 30 words.

- Answer the questions in full sentences, as you will be expected to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will give each Conclusion Interview statement.

# What is most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section this advice correlates to.

  - Author Name: ____ (Step: __) (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Andreas Luch: Pyrolysis is useful to identify multiple components with just one analytical method. Quantification of specific substances can be performed afterwards with more specialized methods.

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Author Name: ____ (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

# Are any of the reagents or instruments hazardous? If so, please use this interview statement to remind viewers of what precautions they should take.

  - Author Name: ___(Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

Thank you for following the instructions and addressing our questions. We will incorporate your answers/suggestions and send you the finalized script before your shoot. You will also receive detailed shoot preparation instructions in the email accompanying the finalized script.