Thank you for the decision on our manuscript ‘Pan-lyssavirus real-time RT-PCR for rabies diagnosis.’

We would like to thank you for your time to provide constructive and helpful comments. We have updated the manuscript accordingly. Below we have taken each comment individually and replied. As the document has had major revision the changes requested, where possible, have been listed as line numbers in this response. Figures and tables have been updated to satisfy reviewer comments.

Kind regards, Denise Marston

Editorial comments:
Changes to be made by the author(s) regarding the manuscript:
1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues. The JoVE editor will not copy-edit your manuscript and any errors in the submitted revision may be present in the published version. Complete
2. Please print and sign the attached Author License Agreement - UK. Please then scan and upload the signed ALA with the manuscript files to your Editorial Manager account. This has been sent by email but has also been uploaded with this amended version for completeness. 
3. Title: Please remove commercial language (SYBR). There is already precedence within JoVE using SYBR: Nicholson P, Rawiwan P, Surachetpong W. Detection of Tilapia Lake Virus Using Conventional RT-PCR and SYBR Green RT-qPCR. Indeed many papers have been published in this manner. However, as per request (6) the use of SYBR has been reduced and we have removed from the title and key words.
4. Keywords: Please remove commercial language (SYBR). Complete
5. Short Abstract: Please add a brief description of the protocol. Complete (lines 21-24)
6. JoVE policy states that the video narrative is objective and not biased towards a particular product featured in the video. The goal of this policy is to focus on the science rather than to present a technique as an advertisement for a specific item. To this end, we ask that you please reduce the number of instances of "SYBR" within your text. The term may be introduced but please use it infrequently and when directly relevant. Otherwise, please refer to the term using generic language. See point 3 above
7. Lines 89-138: Please change the section title to “PROTOCOL”. Complete
8. Protocol: Please revise it to be a numbered list following the JoVE Instructions for Authors; step 1 followed by 1.1, followed by 1.1.1, etc. Please refrain from using bullets, dashes, or indentations.
9. Please revise the Protocol to contain only action items that direct the reader to do something (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “NOTE.” Please include all safety procedures and use of hoods, etc. However, notes should be used sparingly and actions should be described in the imperative tense wherever possible. Please move the discussion about the protocol to the Discussion. Complete
10. Everything in the protocol (except for the introductory ethics statement) should be in a numbered step (in the imperative tense and with no more than 4 sentences), numbered header, or a “NOTE”. Please move the introductory paragraphs of the protocol to the Introduction, Results, or Discussion (as appropriate) or break into steps. Complete
11. Please note that the protocol text will be used to generate the script for the video and must contain everything that you would like shown in the video. Please add more details to your protocol steps. There should be enough detail in each step to supplement the actions seen in the video so that viewers can easily replicate the protocol. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action. Complete
12. Please provide test samples, materials and experimental settings to be used in the protocol. We cannot film a generalized protocol; we need specific settings of a specific experiment. Complete
13. Figure 1: Please include a title for the figure in the figure legend. Complete
14. Figure 1 and Figure 2: Please use the micro symbol µ instead of u (i.e., µg instead of ug) and abbreviate liters to L (L, mL, µL) to avoid confusion (i.e., µL instead of ul). Please include a space between all numerical values and their corresponding units (1 µg, 100 ng, 10 ng, etc.). Complete
15. Please upload each Table individually to your Editorial Manager account as an .xlsx file. Complete
16. Table 1: Please remove commercial language (SYBR®, Kit Bio-Rad iTaq™ Universal SYBR™ Green One-Step RT-PCR kit. Cat No: 172-5150). Please include a space between all numerical values and their corresponding units (20 µM). Complete
17. Table 3: Please change the time unit “sec” to “s”. Please include a space between all numerical values and their corresponding units (50 °C, 10 s, etc.). Complete
18. Please reference Table 5 in the manuscript. Complete
19. JoVE articles are focused on the methods and the protocol, thus the discussion should be similarly focused. Please revise the Discussion to explicitly cover the following in detail in 3-6 paragraphs with citations:
a) Critical steps within the protocol
b) Any modifications and troubleshooting of the technique
c) Any limitations of the technique
d) The significance with respect to existing methods
e) Any future applications of the technique
Complete
20. References: Please do not abbreviate journal titles. Apologies, as the Endnote file was downloaded from the JoVE website the specific details of the formatting were assumed to be correct. The issue appears to be how Endnote holds the citations in the library that caused this issue. To amend this took time but all journal titles are now in full.
21. Table of Materials: Please remove trademark (™) and registered (®) symbols. Please ensure that it has information on all relevant supplies, reagents, equipment and software used, especially those mentioned in the Protocol. Please sort the items in alphabetical order according to the name of material/equipment. Complete

Reviewers' comments:


Reviewer #1:
Manuscript Summary:
The manuscript describes a diagnostic method for rabies, SYBR Green real time RT-PCR, which can detected all 16 recognized lyssavirus species in one step. This method has been accepted by OIE as a primary method, and adopted in many diagnostic laboratories. For readers to learn and adopt this method in their own laboratories more details should be added in the paper. In this regard following comments have been made.

Major Concerns:
1. Regarding controls: What are the positive and negative controls as well as the standard/parameters to prepare them? A section in the discussion has been added to describe how the positive control is prepared and address these concerns (Lines 335-359).

What are the standard/parameters to prepare internal control ß-actin? Is there a range of CT value specialized to define positive control amplification? The amplification plot of ß-actin should be included in the results.

The lyssavirus positive control is also used as the ß-actin positive control as the RNA is prepared from lyssavirus positive mouse brain. Table 5 has been added to include inter-run analysis of the positive control with mean and SD values to aid the reader. The amplification and dissociation plots are included in Fig 1 (C and D).

2. As mentioned in Abstract, this assay has been adopted in many diagnostic laboratories. It would be helpful if authors can extend some information to show how the assay was validated in detection of clinical samples or in different laboratories. Also the authors claimed that the assay can detect all lyssaviruses, including the most divergent viruses IKOV, WCBV and LLEBV, but in Result section no data show the detection of WCBV. In addition, the capacity of this method to detect newly identified TWBLV and KBLV should be discussed.

Addition details of clinical sample data and other laboratories is provided (Lines 284-294). Until recently, we only had data from plasmids containing WCBV sequence, therefore the data was omitted as all other samples were clinically derived brain tissue. However, we now have WCBV grown in mice, therefore brain material could be utilised and compared. This data has now been added to table 5. Comment regarding capability of detecting TWBLV and KBLV has been added (Lines 295-301).

3. In Materials section authors should provide brief description of sample types of rabies virus and other 15 lyssaviruses used to establish and validate the method.

Table 6 provides information regarding the different lyssavirus samples. An extra column has been added with the virus ID (according to our laboratory) which will aid the reader to identify the specific virus strain.

4. The amplification plot of detecting all 16 lyssaviruses (similar to Fig 1A) should be provided in Results. This plot can also be used to accompany Table 5.
The amplification plots are all from multiple dilution series runs. Therefore a selection of 4 (chosen to represent the whole dataset) have been used to create an extra figure (figure 3), including the dissociation curves. In addition, to link figures together, Figure 2 has been updated to present data from the same viruses.

5. Table 5: The sensitivity (lowest positive dilution) varied largely in distinct lyssaviruses from 103-109. What is the initial titer of each sample? Can these be converted to TCID50 or genome copy?

The RNA was prepared from brain material and normalised to 1 µg/µL. Therefore they cannot be converted into TCID50 or genome copy. Sentences have been added to discuss this (Lines 302-305).

Minor Concerns:
6. In Fig 1A: Is the indicator above NTC 10 pg (not ng)? Yes the legend has the indicator above the NTC as 10 pg this has been made clearer.
7. Line 71. The name of two lyssavirus species identified in animals rather than bats should be given. Complete 
8. "Samples should be tested in at least duplicate (ideally triplicate)." Is repeated (lines 93 and 107). The one at line 93 is redundant, should be deleted. Due to editorial changes these sections have been moved to the discussion, and only one remains.
9. Figure 1A. Change "ug" to "µg". Complete 


Reviewer #2:
Manuscript Summary:
In their manuscript entitled "Pan-lyssavirus SYBR Green real time RT-PCR for rabies diagnosis", the authors described the protocol of a real time RT-PCR for the detection of lyssavirus. The manuscript is well written and understandable. However, I think some practical details can be added. In addition, I have questions about the version of the protocol compared to the version already published recently the OIE Manual terrestrial 2018, which can be confusing. Details of these comments are indicated below, with other remarks/comments.

Major Concerns:
The protocol is detailed but probably more practical and basic information could be added to complete this protocol. For example, the version found in the OIE Manual terrestrial 2018 cited by the authors (and which must also have been written by these same authors) is more detailed, even if it seems to have been almost copy/paste for this JoVE manuscript (without some of these details). As examples, the authors did not indicate that the enzyme mix should kept on ice, and that the remaining reagents can be thawed at room temperature for point 2 in line 130, or that in line 125 and line 130 that to keep all reagents on ice during the preparation of the reaction master mix. Authors should also indicate the number of supplementary samples, such as in the Manual "Allow for pipetting variation by preparing at least two extra reaction mixes".

These points have been added as ‘NOTES’ in the protocol as per request of editors.

I notice some discrepancies between the protocol described recently in the OIE Manual terrestrial 2018 and the protocol presented by the authors, which is confusing for me and probably for the other readers. For example, the volume of RNA used is 1 μl here (line 133), but 2 μl of test RNA (where possible at concentration of 0.5-1 μg/μl for extracted samples) in the Manual. Another example is the volume of the primer N165 (-146): 0.6 μl per sample in the protocol of the manuscript and 1 μl in the OIE Manual terrestrial.

The discrepancies noticed are due to the changes that the kit manufacturer made which altered the final volume of the master mix. We validated the amount of RNA from 2 µL to 1 µL when this happened, but the change wasn’t reflected in the OIE manual protocol. Either volume can be used but 1 µL is preferred.  A note has been added at this step. The primer amount should have been 0.6 in the OIE manual so this protocol is correct.

- Line 32 and line 84: The authors referred to the WCBV, but I was not able to see the results with this species in the Table 5. Was this important species (in terms of genetic divergence) forgotten, or just not tested?

Until recently, we only had data from plasmids containing WCBV sequence, therefore the data was omitted as all other samples were clinically derived brain tissue. However, we now have WCBV amplified in mice, therefore brain material could be utilised and compared. This data has now been added to table 6.
- Line 119: Can the authors comments about the limitation of use of the cellular internal endogenous gene beta-actin and CSF or even some saliva samples (indicated line 102)? Indeed, by experience, we not that this internal control is quite often negative with acellular CSF (with low number of cells), and with some saliva samples. This issue can be solved with the use of an heterologous internal control, directly introduced in the sample for example.

Table 4 has a note added to it to suggest an external control to be used in such samples if the ß-actin is negative, also a section has been added in the text (lines 349-359).

- Lines 126-127: What are the evidences that the master mixes can be prepared and stored frozen up to one year?

Apologies this was data obtained for our probe based assay which used individual reagents in the master mix (dNTPs, MgCl2 etc). The kit used here has all the reagents on one reaction mix (apart from the enzyme) so it is not necessary to make large master mixes and freeze in batches. Reference to his has been removed.

- Lines 149-151 and lines 181-184: Can the authors add more precision/details about this critical step? For example, it is known that the Tm value will be different from the Tm of the positive control used, depending of the base composition of the amplicon. What are the range of Tm values acceptable for the authors to consider a sample as positive, compared to the control? In particular line 159: what do the authors mean with "equivalent to the positive control"? Or what does a "melt curve which is different to the positive control" (line 183)? Also, the authors present in Figure 1 a single and clear peak, but we know that sometime we can have a principal peak and other secondary peaks. What are the interpretation of the authors when we faced to multiple peaks, especially if one of them match with the one of the positive control?

These points have been addressed in the ‘results section’ including adding ranges of acceptable Tm results to aid the reader (lines 214-232). In the discussion it is emphasised that the range needs to be calibrated according to each laboratory setting (lines 312-315).

- Line 163-164: For my part, this looks more like the limit of detection rather than the sensitivity, even if the first one is a part of the analytical sensitivity (OIE, Principles and methods of validation of diagnostic assays for infectious diseases). So I suggest the authors to modify, adapt or justified the term used here. Here the auhors mainly look to the efficiency of the real-time RT-PCR. Also they claim that the sensitivity for all the three lyssavirus is comparable. How it is justified, measured?

The results section has been modified in view of the reviewer’s comments
- Line 166: Where are the results with WCBV? Until recently, we only had data from plasmids containing WCBV sequence, therefore the data was omitted as all other samples were clinically derived brain tissue. However, we now have WCBV grown in mice, therefore brain material could be utilised and compared. This data has now been added to table 6.

- Line 168: It will be nice to have the list of the different rhabdoviruses tested (in supplementary material for example). References to two papers investigating the same virus panel have been cited (lines 244-245).
- Table 4: For the two conditions with repeated experiments, what are the decision after them, if the results are still the same? Notes have been added to the table and this is discussed in the results section (Lines 353-359).
- Table 5: This table is difficult to interpret. Indeed, the dilution factor will depend of the initial concentration of the tested sample. But we do not have any information about them (primary samples, virus culture?). We do not have even any reference of the strain used (GenBank number if available or reference of the strain). I suggest to the authors to modify/complete this table to give to the readers the necessary information to interpret these results.

Table 5 (which is now table 6) has been modified to add the reference for each virus and the results section updated to clarify the samples are from brain material, which have been standardised to 1 µg/µL total RNA (Lines 302-304). The Tm for each virus has also been added.

Minor Concerns:
- Line 26: Develop "OIE" Complete
- Line 94: I do not have access to the Excel spreadsheet. Uploaded in this revision
- Line 164: This is not "specificity" but rather "sensitivity". Please correct. Complete
- Line 167: No capital letter for "Rhabdoviruses" Complete
- Line 169: Develop "EURL" Complete
- Line 186: There is both specificity and sensitivity. Updated
- Line 190: Can the authors provide some references to associate with "ante-mortem samples"? Examples have been given
- Figure 1A: Precise the legend (RNA?, what is NTC, what does +ve mean)? Complete
- Figure 1B: There is no legend. Added
- Figure 2: Why the authors do not indicate the parameters of the amplification curve for each samples (slope, R2, etc.)? The figure has been altered to show data used in Figure 3, and R2 calculated. It should be noted the data is not used as controls to standardise results, rather to indicate the robustness of the assay across the genus. The data is discussed (lines 233-239).

- Table 2: Indicate the position of the primers on a reference sequence (both for lyssavirus and beta-actin). These are provided in a previous publication, which has been cited. Complete