[image: image2.jpg]ﬁ HELMHOLTZ

ZENTRUM BERLIN

fur Materialien und Energie



[image: image3.emf] 


[image: image4.png]HZB Helmholtz

Zentrum Berlin

Helmholtz-Zentrum Berlin fiir Materialien und Energie GmbH | Hahn-Meitner-Platz 1 | 14109 Berlin




[image: image5.png]HZB Helmholtz

Zentrum Berlin





	Dr. Manfred S. Weiss

	Macromolecular Crystallography (HZB-MX)

	Helmholtz-Zentrum Berlin

	Albert-Einstein-Straße 15

	D-12489 Berlin, Germany

	Tel  +49-30-8062-13149

	Fax +49-30-8062-14975 USERADDRESS   \* MERGEFORMAT 

	

	msweiss@helmholtz-berlin.de

	

	Berlin, Feb 22, 2019

	


	To the

	Editorial Manager of JoVE

	Nam Nguyen

	JoVE

	


Ref.: Revised Manuscript Submission
Dear Editor(s) of JoVE,

Thank you very much for your comments on our manuscript on a new sample holder for macromolecular crystallography. In the following we would like to answer all points raised by the referees and by you. We hope that you find our comments and the concomitant changes we applied to our text and the figures satisfactory and that our manuscript is now suitable for publication in JoVE. Please note that the line numbers may be different depending on whether you view the clean manuscript or the manuscript with tracked changes. In this letter, we refer to the line numbers of the clean manuscript.
Comments from the Editor:
1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues. The JoVE editor will not copy-edit your manuscript and any errors in the submitted revision may be present in the published version.

We have tried our very best to thoroughly proofread our manuscript for spelling and grammar errors.
2. Please revise the table of the essential supplies, reagents, and equipment. The table should include the name, company, and catalog number of all relevant materials in separate columns in an .xls/.xlsx file. Please sort the Materials Table alphabetically by the name of the material.

The revised Materials Table now includes all essential supplies, reagents and used equipment and is sorted alphabetically by the name of the material. When using trade names of individual materials, the manufacturer reference number is also given. If a given material is available from more than one vendor, a comment was added to the respective line.
3. Please provide more dimensions of the sample holder. Can the STL file be provided? If someone were to repeat the protocol, how would they make the holder or where can they obtain it?

In order to provide more information on the dimensions of the sample holder, we have redone Figure 1 with scale bars added. I am very much afraid to say that we will not be able to provide the associated STL files due to the pending patent applications. The sample holder has already been made commercially available under the trade names XtalTool and XtalTool HT by the company Jena Bioscience (Jena, Germany), respectively. Sample holders can be either purchased directly from Jena Bioscience or via one of their distributors. Please refer to the Jena Bioscience website: https://www.jenabioscience.com/crystallography-cryo-em/data-collection. Note: If the provision of the STL file is a mandatory requirement for publication on your side, we will have to withdraw the manuscript. Please advise!
4. JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials and Reagents. For example: CrystalDirect, etc.

We have edited the manuscript and removed all commercial language including all trademark symbols, registered symbols, etc. All relevant information was added to the Materials Table including references. We would like to point out, however, that in your example the term CrystalDirect was already used for the system before it became commercially available. If we use other terms for it, this sound very artificial and it will be utterly confusing to the readers.
5. Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action.

We have revised our protocol and adapted it to answer the “how” questions in all necessary steps. References have been added to steps for further information (lines 198, 206 and 230). We have also redone some of our figures. For instance, the setting up of crystallization drops is now illustrated with a renewed Figure 2 to allow a better visual illustration how this is step needs to be performed.

6. 2.1: How is the surface cleaned and made dust free?

Detailed information is now included within the protocol. From line 138 onwards, the text reads now as follows: “Create a clean and dust-free surface using a damp lint-free cloth. Take one sample holder from its box and gently place it yellow foil facing up on the cleaned surface to avoid damage or unwanted puncture of the backside COC foil. Introduced scratches may disturb later on the crystal observation.” Information about the recommended material was also added to the Materials Table.

7. 2.2: What is the desired volume? It is best to provide a specific example here instead of generalized protocols. As the representative results used is lysozyme, please provide these details throughout the protocol.

We have added more detailed information throughout the protocol. To elucidate the handling, we included the new panel B of Figure 2 and added a new table summarizing the experimental details for crystallization (Table 1) and data collection (Table 2). 
8. What is the reservoir? Concentrations and volumes?

Detailed information about the used reservoir composition has now been given in the respective paragraph starting from line 253. In addition, the new Table 1 provides summarized details. 
9. What are the incubation conditions?

The protocol was adapted to provide detailed information regarding the recommended volumes. The paragraph in the representative results section describing the experimental procedure in detail was significantly revised and many more details regarding the incubation conditions are now provided (line 253 and following). 

10. Please provide all volumes and concentrations used throughout.

In order to not be too repetitive throughout the text, we have added a paragraph in the representative results section, which gives the requested information. Experimental details used in the reported case are provided in this section. We have also included Table 1 in order to provide precise information on the experimental setup. 

11. Please be more neutral and discuss some limitations of the sample holder in the discussion.
We have added a new paragraph on the limitations of the sample holder at the end of the discussion section. This new paragraph reads as follows: Limitations. The sample holder’s geometry permits unobstructed diffraction data collection by the rotation method over a total rotation range of 160°. This is sufficient so that complete diffraction data sets can be obtained for most crystal systems. In cases where this is not possible, data from more than crystal need to be merged together. When crystals are gown together, it may be possible to adjust the size of the incident X-ray beam so that only parts of individual crystals are exposed. In extreme cases, one may need to resort to a data collection strategy similar to the MeshAndCollect approach (Zander et al., 2015). In summary, while there are some limitations associated with the sample holders, these can easily be overcome in most cases. Of course, it is always possible that situations are encountered, in which none of this is possible. In such cases, one may need to resort to other crystal mounting methods.
Comments from referee 1:

Minor Concerns: Use of commercial words/language like "easy", "not really a problem".
We have corrected the protocol accordingly.

Comments from referee 2:

Manuscript Summary:
In this article, the author devised a simple device that allowed data collection screening in situ without the need for direct manipulation of the fragile protein crystals. Also, as background signal from buffer diffraction is a problem, this new invention could drastically decrease buffer content of the crystal for data collection. The manuscript is well-written and explains the idea and method clearly. The beauty of this design is that the device, or in other words, a substitute of the widely used glass cover slide, is compatible with the widely used 24-well Linbro style plate. The use of two layers of membranes, namely a self-healable COC membrane and a liquid permeable polyimide membrane allowed sealing for crystal growth and buffer exchange or drain with ease. The design itself is smart and should increase the efficiency of the screening process.
Major Concerns: However, the downside of the design is that, even though this gadget seems suitable for screening crystals without direct manipulations of the crystals, it does not allow for single crystal being isolated and harvested for data collection if there're a bunch of crystals appeared in a single drop.

This is indeed true. We have also stated this in the manuscript on page 11, lines 431-440. It is clear that there is no one solution to all possible problems. Avoiding manipulations thus means leaving the crystals alone. However, if optimized crystal growth will not lead to individual or well-separated crystals, there is still the chance that by tuning the size of the incident X-ray beam only one single crystal will be exposed at a time. We have also stated this explicitly in a new paragraph on limitations in the Discussion section.
What’s more, currently, we routinely used robot to set up 96-well sitting drop plate for screening crystal growth conditions and most of the time we only use 24 well plate for optimization of crystal growth conditions, and if single crystal do occur, then this device should for sure avoid direct manipulation of the crystal. But most of the time, multiple crystals are seen in a single drop. And if a crystal is seen that diffract x-ray well and further data collection is desired, then there’s a problem that other crystals around will hamper data collection at an angle parallel to the membrane. For crystals like those formed by lysozyme, even if there’re multiple crystals in a single drop, it should not be a problem because each crystal diffracts quite well and one can harvest a full data set by combining the data collected from multiple crystals. However, in circumstances like some particular crystal diffract well while others diffract poorly, it is always desired to isolate and harvest the single crystal for diffraction screening and for data collection. 

This comment is a continuation of the previous comment. The here described sample holders are designed for optimization experiments of crystal growth conditions, which were identified from large screening campaigns. This means that initial crystallization conditions are already known from experiments such as the one described by the referee. By varying the ratio between mother liquor and protein solution, the number of expected nucleation sites may be tuned. Other than in the screening plate, the sample holder does not only allow to test but also collect diffraction data from each crystal individually. Diffraction data collection might be hampered due to the parallel orientation of crystalline material on the membrane. However, the sample holder allows the usage of a kappa goniometer in order to minimize the chance of illuminating multiple crystals at a given time. We further like to refer to Zander et al. MeshAndCollect: an automated multi-crystal data-collection workflow for synchrotron macromolecular crystallography beamlines. Acta Cryst. D71, 2328–2343. This reference is listed in the manuscript as reference 35.
What's more, for crystal grown in a cluster or overlay 3 dimensionally, it is still a problem even with current design, limiting its possible applications.

Again, this comment is a continuation of the two previous comments. In the case of crystals grown in clusters, the addition of additional mother liquor from the back of the crystal growth support might help to relocate some clusters and allow data collection on individual crystals (or parts of them). As already mentioned the design reported here is not magic. There will always be cases, when one has to resort to other approaches.
Comments from referee 3:

Manuscript Summary:
The manuscript submitted by Feiler et al. describes a new sample holder for macromolecular crystallography. In this approach, crystals are grown on a polyimide foil attached to a 3D-printed plastic holder using conventional vapor-diffusion method. The holder is then affixed to a magnetic base for subsequent X-ray analysis at either room-temperature or under a nitrogen stream. The system can be conveniently used for crystal treatments such as ligand soaking, solvent removal and cryo-protection. The principles of the method are very similar to the CrystaldirectTM method by Cipriani et al. but applied to a popular crystallization plate type used for manual crystal optimization. Like CrystaldirectTM, crystal manipulation (typically using a "fishing" loop) is avoided but unlike CrystaldirectTM, the presence of the plastic holder frames limits the range of oscillation for crystallographic data collection, an inconvenience that is not clearly mentioned in the manuscript. The method is of interest for crystallographers and has the potential to expand serial crystallography approaches. The manuscript is worth publishing but would need to address the following comments prior to publication.
Minor Concerns:

l. 98. Refer to Table 1 of Aller et al. 2015 Methods in Molecular Biology and l. 100. Add CrystaldirectTM reference here (citation 27).

We have added the two references at this point.
l. 103. The authors present the disadvantages of the CrystaldirectTM approach, namely the need for opening the crystallization chamber and the potential difficulty of isolating single crystals. Both aspects are however still present in the method described in the present manuscript. The authors may want to tune this sentence differently.
In response to this comment, we have deleted the sentence the respective sentence from the text.
l. 106. Replace "our aim was to develop" with "we aimed at developing" and l. 111. Replace "we would like to describe … and to demonstrate" with "we describe…and demonstrate"
The text has been changed in order to avoid the use of the first person.
l. 121. Figure 1. Add a scale bar.
The new and redone Figure 1 contains now a scale bar in all panels.
Figure 1. Describe the purpose of the 2 little pins protruding out. Does it puncture the foil to hold it in place?

The protruding pins are used during the manufacturing process to align the polyimide foil. An explanatory sentence was added to the first section the description of the sample holder, page 3, lines 139-140.
Figure 1. Highlight the breaking point for 1B and 1C.
The designated break points are now highlighted by red arrows, which have been added to Figure 1, panel B. We have further induced the new Figure 6 in order to illustrate the breaking of the sample holder for data collection. 

l. 123. Add the thickness of the foil and comment on the absorption at common energies used in macromolecular crystallography.
The thickness of the foil and comments on absorption have been added on page 3, line 131. 

l. 126. Does the yellow color affect the polarized light or UV light?
Polarized light is not affected by the yellow color. Figure 4 was taken on a transmission microscope using a polarizer in place. UV light we have not tested, since we have no such facilities available in our laboratory. We have stated this on page 7, lines 258-259 and page 8, lines 309-312 of the manuscript.
l. 146. It would be good to have a dedicated figure illustrating this point.
Unfortunately, we are unable to address this point since we do not have access to an imaging robot or an imaging hotel in our laboratory. 

l. 168. The use of a finger to cover the hole sounds a bit rudimental. Please explain in more details.
We agree with this comment of the referee, but sometimes it is the simplest solutions, which are the best. We have added a few more details on this to the protocol. 

l. 173. Replace "4.1.5" with "4.1.3"
Thank you for pointing this out. Done.

l. 192. Description used on point 6.1.1 (l. 213) reads better.
We have followed the suggestion and introduced the respective changes.

l. 198. What does "correct positioning" means (a figure may illustrate this better)? What about icing? Please describe the limitation in oscillation range.
In order to illustrate the term “correct positioning” we have included indicating arrows in Figure 2. We would also like to refer to the protocol text to the figure panel 2A for better understanding. We have collected data after the sample holder was left for 1 hour on the goniometer and no significant icing was observed. This is now stated in the text on page 10, lines 389-393. A paragraph regarding limits on the rotation angle has been added on page 10, lines 394-397.

l. 205. Refer to Fig 1 and indicate the breaking points.
We have revised Figure 1 and included that point.

l. 215. Is the holder attached with glue? Describe what a "pre-prepared" magnetic base is.
Since the sample holder is a device for single use, we have designed the attachment to the base so that it works without glue. The sample holder is clamped onto the magnetic base by inserting it all the way to the indicated marker. We have revised Figure 1C by highlighting the marker, up to which the sample holders need to be inserted into the base, to explain this in detail. Further, the new Figure 6 illustrates the preparation procedure.

l. 220. How stable are both the sample holder and the film in LN2?
Both are stable in liquid nitrogen. The stability of the sensitive polyimide film has been tested at 77 K previously by Yano and Yamaoka (1995). Cryogenic properties of polymers. Progress in Polymer Science. 20, 585–613. The appropriate reference has been added to our list of references as reference no. 34.
l. 221. The transfer into vials only works for case described on Fig 1B and C (not A). Please amend.
This has been clarified on page 6, line 241-243.

l. 229. What is the recommend volume for the drops?
We have added recommended drop volumes and drop numbers for all types of sample holders in the manuscript, page 7, line 249-252.
l. 279. The use of the polarizer could be illustrated in a figure as well.
Figure 4 was taken using a polarizer. We have now also included the possible usage of a polarizer in the figure description.

l. 242. Add beam size, flux and dose.
The requested information has been added to the text on page 7, line 265.
l. 243. It would be good in this paragraph to add the merging statistics if enough data were collected.
We have now included a new Table 1 that provides the merging statistics for a data set collected using the sample holder. 

L. 253. Add thickness of the foil.
The thickness of the foil was added to the general description. 

l. 277. Replace “Observation of crystal growth” with “Picture of hen egg-white crystals”
The requested change has been made.

Figure 5. Add resolution rings. Indicate where the background from the foil is.
Figure 5 was revised and includes now distinct resolution rings. 

Comments from referee 4:

Manuscript Summary:
This paper describes an interesting approach using a newly designed device to handle crystals grown in drops, highly decreasing their manipulation before X-ray diffraction data acquisition. The device was conveniently prepared to be used with Linbro plates, which are still highly used for protein crystallization experiments, despite many other plates have been made commercially available. The device has also been made commercially available by Jena, showing its putative commercial potential. There are some issues which could be improved/changed for a better understanding of the reader.
Major Concerns: 

Whereas I agree that the use of capillaries became impractical for cryoprotection many years ago, at present, protein crystals grown in capillaries by the counterdiffusion technique can be readily prepared in the lab, sent to a synchrotron source in a dewar, and robotically mounted for data acquisition under a nitrogen steam, totally avoiding crystal manipulation; furthermore, ligand soaking and cryoprotection can also be achieved by counterdiffusion once the crystals have been prepared just by placing a cut capillary inside a mother liquor containing the compound of interest (Gavira et al. (2002). Acta Cryst D58:1147-54; Martinez-Rodriguez et al. (2008). Acta Cryst F64, 50-53). This information could be added in the introduction section, where the authors describe the different methodologies describing the problems of protein crystal handling.

We have added the method and the associated references to the introduction section of the manuscript (page 2, lines 75-77).
One of the major drawbacks of the device might be a high number of crystals in the drop (e.g., Fig 4). Whereas this issue could be minimized by setting up different drops reducing nucleation, I think this issue should be mentioned somewhere in the paper (maybe a footnote), just in case crystal improvement is not feasible; this issue would difficult data acquisition if proper centring is not allowed due to the limitations of positioning the goniometer, even when we know that the corresponding datasets could be treated afterwards to discriminate between different crystals. Thus, a short explanation for putative users of the holders might help to understand how to overcome some minor issues on the handling at the X-Ray source. 

We would like to thank the reviewer for this remark and we have included a paragraph explaining the limitations of the device and how they can be overcome. The sample holder is designed for optimization of crystallization conditions. It allows multiple drops to be set. Tuning the ratio of protein to mother liquor may optimize the number of nuclei to grow fewer, but larger crystals. If this is not applicable to the specific case, using kappa angles on the goniometer can increase chances of proper centering and data collection in challenging cases. This is now all mentioned and discussed on page 10, lines 394-399.
I also suggest the authors to consider the gelling of the crystallizing drops as they will fix the crystal for further measurements. 

While we agree that gelling might be an option for immobilization of crystals; we do believe that it is not really necessary, since the crystals will hold on to the sample holder by mere surface tension. Since gelling might also interfere with ligand soaking or cryo protection, we would rather refrain from adding this as a recommendation.
I miss some statistics of crystals collected using the classical "fishing out" method and crystals measured using the holder presented in this work would reinforce the holder benefits.
We have now included some data processing and merging statistics of a collected data set using the sample holder (see Table 2). 

However, the main concern is not about the science but about the protocol itself. I think that the protocol needs a large revision to be understandable and useful for readers, mostly students. For instance, I have missed some images or extra information showing the kind of cannula used to pierce the foil, the fine paper wick, or how to reseal the foil after manipulation.

We have included a new Figure 2 to illustrate how to use a cannula and the fine paper wick and we have tried to be as informative as possible in the protocol, in particular for novice users.
I am sure it is quite straightforward, but since I could not see the holders to make an idea, it would help readers to know what kind of materials to be used.

The full list of materials used in now in the revised Materials Table.

A short sentence (a movie?) describing how the holders with crystals can be easily sent/transported to a synchrotron source might be added.

Thanks a lot to the reviewer for mentioning this. This is exactly the reason why we are trying to publish this work in JoVE.
In section 4 needs a picture or illustration on how and where to perforate the COC film.

A new panel for illustration on how to prepare the sample holder for ligand soaking and removal of mother liquor was added to Figure 2.

Section 4.2 needs a fully rewritten. If the solution is located between the polyimide and the COC films, will the ligand migrate through the polyimide film to the crystal sample? Please explain it clearly.

Yes, the reviewer is exactly right in this. We have thoroughly reviewed section 4.2 and clarified it in various places. Further, the new panels in Figure 2 should help to understand the procedure further.

In 4.2.4., do you mean that just by touching the COC film it will auto-seal it? Please explain it better.

Yes, this is indeed the case. We are now provided a more extensive explanation in the protocol.

Section 5 & 6: how is the holder fix to the magnetic cap? How does it looks like? 

Figure 6 was added to elucidate, how the holder is fixed on the magnetic cap. The procedure was also added to the text, page 5, line 209 and is illustrated in the new figure 6.
Is there any problem with the humidity control system and the polyimide foil? A picture exemplifying the set-up could be of great help.

The employed polyimide does not cause any problems with the humidity controlled sample environment. The material itself is rather inert to humidity and is therefore not prone to swelling are absorb a significant amount of humidity. Taptimthong et al. have tested this previously with slightly thicker material of 25 µm. The result was that the total humidity absorption was negligible (Taptimthong et al. (2014). Flexible Magnetic Writing / Reading System: Polyimide Film as Flexible Substrate. Procedia Technology 15, 230–237.). Within the humidity controlled sample environment, its porous structure allows to keep an equal humidity level on either side. We now have included new Figure 6 to exemplify the workflow.

With best regards,

[image: image1.wmf] 


Manfred Weiss

[image: image6.jpg]Helmholtz-Zentrum Berlin fiir Materialien und Energie GmbH | Mitglied der Helmholtz-Gemeinschaft Deutscher Forschungszentren e. V. | www.helmholtz-berlin.de

Vorsitzender des Aufsichtsrats: Prof. Dr. Dr. h.c. mult. Joachim Treusch | Stellvertretende Vorsitzende: Dr. Beatrix Vierkorn-Rudolph
Geschaftsfiihrer: Prof. Dr. Anke Rita Kaysser-Pyzalla | Dr. Ulrich Breuer | Sitz der Gesellschaft: Berlin | Handelsregister: AG Charlottenburg | 89 HRB 5583









