# Submission ID #: 59746

# Scriptwriter Name: Brigid Stadinski

# Project Page Link: http://www.jove.com/files_upload.php?src=18222578

Title: 4D Printed Bifurcated Stents with Kirigami-Inspired Structures

# Authors and Affiliations:

Dahong Kim1, Taeyoung Kim1, Yong-Gu Lee1

1School of Mechanical Engineering, Gwangju Institution of Science and Technology, Gwangju, Republic of Korea

Authors: Please ensure that all authors’ names are spelled correctly and that the affiliations listed here are correct. This is how your names and affiliations will appear in your video.

# Corresponding Author:

Yong-Gu Lee

lygu@gist.ac.kr

# Email addresses for Co-authors:

Dahong Kim	(kdhong65@gist.ac.kr)

Taeyoung Kim	(idealty90@gmail.com)

# PLEASE READ THE INSTRUCTIONS IN THE GRAY BOXES CAREFULLY AND USE TRACK CHANGES WHILE MAKING ANY EDITS TO THE DOCUMENT.

# This document has several sections on separate pages, so take care to view each page.

Author Questionnaire:

Authors, please fill out the unanswered questions below.

1. Microscopy: Does your protocol involve video microscopy, such as filming a complex dissection or microinjection technique? (Y/N)  No

Can you record movies/images using your own microscope camera? (Y/N) No

If no, JoVE will need to record the microscope images using our scope kit (through a camera port or one of the oculars). Please list the make and model of your microscope.

2. Does your protocol include software usage? (Y/N) Yes.

If yes, we will need you to record using screen recording software to capture the steps. If you use a Mac, QuickTime X also has the ability to record the steps.

3. Which steps from the protocol section below are the most important for viewers to see? Please list 4-6 individual steps using the step numbers listed in this document. This information is important to prepare your Videographer for your shoot. (You do not need to include steps that will be screen captured. Please do not list entire sections.)

Authors, please answer this question with the steps listed here in the Protocol section below for use by the videographer.

4. What is the single most difficult aspect of this procedure and what do you do to ensure success? Please list 1-2 individual steps using the step numbers listed in this document. (Please do not list entire sections.)

Authors, please answer this question with the steps listed here in the Protocol section below for use by the videographer.

5. Will the filming need to take place in multiple locations? (Y/N) Yes

If yes, how far apart are the locations? The locations are in 500m in our campus. 10 minutes on foot, 5 minutes by car.

Section - Introduction

Videographer: Interviewee Headshots are required. Take a headshot for each interviewee.

Authors, these headshots will be used for the JoVE Dedicated Author Webpage. Here is one example if you wish to take a look.

- REQUIRED Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

- The total introduction length (i.e., Required and Optional Interview Statements) cannot exceed 150 words.

- Restrict the length of each statement to no more than 30 words.

- Please answer the questions below in full sentences to highlight the significance of your protocol. You will be expected to memorize and deliver these sentences as spoken interview statements during filming.

- Indicate the full name of each author who will give each statement. If only one author is giving the REQUIRED statements, the same author may speak both statements.

# Why is your protocol significant? OR What key questions can this method help answer?

  - Yong-Gu Lee: We show how to 3D print using shape memory polymers. Shape memory polymers are interesting that it can be programmed to change shape when certain conditions are met. 3D printing can produce shapes by stacking layers after layers. Allowing us to make any shape that is imaginable. Together, we show how we can produce stents that can be shrunk to a compact form when being delivered and recover the most complex shape at the right moment. Unfortunately, great care needs to be exercised to use shape memory polymers in 3D printing. We answer how anyone can do this at their own laboratory.

# What is the main advantage of this technique?

  - Dahong Kim: We can make stents that is custom tailored to each patient. This is possible because our method can produce stents using 3D printers and anything that can be designed can be exactly manufactured. Delivering stents through narrow and long tubes can be challenging as stents are much bigger than the diameter of human vessels. We show programmed shape changing mechanism that maintains its compact form when traveling through the passages and return to its intended shape at the target site. By the use of our method, it is possible to design custom tailored stents for each patient. Furthermore, we demonstrate that bifurcated stents can be inserted into the y-shape blood vessels with single operation, compared to the multiple operations needed in the conventional operations.

OPTIONAL Interview Statements: (Said by you on camera)  - All interview statements may be edited for length and clarity.

- If you would like to have additional speakers, the following OPTIONAL questions may be answered to provide additional introductory information about your protocol.

- These statements must be spoken by different authors than those who gave the Required Interview Statements, and are limited to one statement per additional author.

- The length of each statement is restricted to no more than 30 words and contributes to the total introduction length, which cannot exceed 150 words.

- Indicate the full name of each author who will give each OPTIONAL statement.

# Do the implications of this technique extend toward the therapy (or diagnosis) of a particular disease, disability, or challenge? How so?

  - Dahong Kim: Our method can be used to treat coronary artery diseases, blood vessel blocked by plaques. Our method can also open may passageways such as bile ducts, which are tubes that carry bile to and from digestive organs, bronchi, which are small airways in the lungs and ureters, which are tubes that carry urine from the kidneys to the bladder

# Are there any specific areas of research that this method could provide insight into?

# Can this method be applied to any other systems?

  - Dahong Kim: 4D printing and Kirigami can be used where objects need to be compacted when delivered and return to its larger shape on site. There are many applications that need this in areas such as medicine and military.

# How would you expect an individual who has never performed this technique to struggle?

# Do you have any advice to offer to somebody who is trying this technique for the first time?

  - Dahong Kim: It is important to find optimal printing conditions when making the stent. Also it is recommended to avoid or minimize the generation of supporters, as they will impose additional postprocessing burdens.

# Why is visual demonstration of this method critical?

  - Dahong Kim: Our method requires the understanding of time-based morphing. This can be most effectively understood through visual demonstrations.

# Introduction of Demonstrator: (Said by you on camera)

- Please use this statement ONLY if any of the individuals who will be demonstrating the procedure on camera have not given a required or optional Introduction interview statement already.

- Include the full name(s) of the person(s) demonstrating the experiment followed by their title (e.g., technician, post doc, grad student, clinician, etc.)

- Also indicate the full name of the author who will introduce the demonstrator(s).

  - Dahong Kim: All Demonstrations of the procedures will be done by Dahong Kim from my laboratory.

    - Interview style: Author saying the above

    - The named technician, post doc, student looks up from workbench or desk or microscope and acknowledges the camera.

Section - Protocol

- Read through the entire protocol carefully to understand what you will need on the filming day and prepare accordingly.

- The two-digit numbers (e.g. 2.1., 2.2.) represent the “steps” of you protocol and will be read by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the “shots” that our videographer will capture at your lab.

- To ensure that your protocol can be filmed in a single work day, the protocol is restricted to 30 steps and/or 60 shots.

- It is critical for a smooth and organized shoot that all materials and work spaces are prepared and labeled (if applicable) in advance.

- Any specimens/samples that require long or overnight incubation steps should be prepared in advance. (i.e. day 0 sample preparation will be filmed on the day of the shoot; day 1 samples should be prepared the day before the shoot so their processing can be filmed on the day of the shoot/after their overnight culture/treatment/etc.)

- Blood Vessel Mock-Up Design and Fabrication by Molding

  - Set the diameter of the proximal main vessel to 25 millimeters. Then, set the diameters of the distal main vessel and the side branch equal to 22 millimeters. Set the total length of the vessels equal to 140 millimeters [1].

    - SCREEN: To be provided by the authors - Screen capture movie of the 3D printer software as talent sets the diameter of the proximal main vessel to 25 millimeters. Then, talent sets the diameters of the distal main vessel and the side branch equal to 22 millimeters. Talent sets the total length of the vessels equal to 140 millimeters. Authors, please upload this screen capture to your project page.

  - Next, set the length of the proximal main vessel to 65 millimeters, the distal main vessel to 75 millimeters, and the side branch to 65 millimeters [1].

    - SCREEN: To be provided by the authors - Screen capture movie of the 3D printer software as talent sets the length of the proximal main vessel to 65 millimeters, the distal main vessel to 75 millimeters, and the side branch to 65 millimeters. Authors, please upload this screen capture to your project page.

# I think the part of 2.1.1 “Talent sets the total length of the vessels equal to 140 millimeters.”can be moved 2.2.1

  - Print the computer model of the branched vessel by using a fused deposition modeling 3D printer. Use a polycarbonate filament [1]. Authors, is this printing process filmable? If not, I can simply re-phrase this step. Please re-phrase this step.

    - CU: 3D printer printing the branched vessel.

  - Create a box-shaped container that will house the 3D printed part. Set the container dimensions to 110 by 105 by 70 millimeters and use an acrylic plate. –Authors, is this printed as well? If so, would you prefer to show the screen capture footage here or footage of the container being printed? Container is not printed by 3D printer. It is made cutting and gluing acrylic plate. But I attached the screen capture to indicate dimension of the acrylic box using 3D CAD modeling.

  - With the 3D printed branched vessel placed at the center of the box, gently pour the silicone inside the container to minimize bubble formation [1]. Dry the liquid silicone and harden it for 36 to 48 hours [2].

    - CU: Box with branched vessel in the center as talent gently pours the silicone inside he container to minimize bubble formation.

    - MED: Talent leaves the liquid silicone to harden. –Authors, how do you dry the liquid silicone? Do you leave it at room temperature to harden? Is any other action performed here? Yes, just put it at room temperature. But it takes very long time to dry. Probably, I’m ready to make silicone mock-up in advance before filming.

  - Once hardened, remove solidified silicone from the container and cut it in half to remove the 3D printed part [1]. Rejoin the divided silicone at the cut plane. The resulting joined body is the blood vessel mockup [2].

    - CU: Solidified silicone as talent cuts it in half to remove the 3D printed part.

    - CU: Divided silicone as talent rejoins it at the cut plane.

- Design of the Branched Stent Based on Kirigami and 3D Printing with Shape Memory Polymer Filaments

  - Design the trunk of the bifurcated stent following wavy patterns similar to conventional stents. Set the diameter of the trunk to 22 millimeters and the length of the trunk to 38 millimeters [1].

    - SCREEN: To be provided by the authors - Screen capture movie of the 3D printer software as talent sets the diameter of the trunk to 22 millimeters and the length of the trunk to 38 millimeters. Authors, please upload this screen capture to your project page.

  - Design the bifurcated branches to be a cylinder. Set the diameter of the branch to 18 millimeters and the length of the branch to 34 millimeters. Set the total length of the stent to 72 millimeters [1].

    - SCREEN: To be provided by the authors - Screen capture movie of the 3D printer software as talent designs the bifurcated branches to be a cylinder. Talent sets the diameter of the branch to 18 millimeters and the length of the branch to 34 millimeters. Talent sets the total length of the stent to 72 millimeters. Authors, please upload this screen capture to your project page. I uploaded both 3.1 and 3.2 at once. Is it okay?

  - Print the bifurcated stent in a fused deposition modeling 3D printer using a shape memory polymer filament. The major composition of this filament is polyurethane [1].

    - MED: Talent places the SMP filament into the 3D printer. – Authors, is this an appropriate action to show here? If not, could you suggest another shot? I think it can be used the scene of inserting the SMP filament into the 3D printer.

  - Use slicing software for model slicing and to control the settings of the 3D printer. Set the extruder temperature to 230 degrees Celsius and the temperature of the printer bed to room temperature. Set the layer height to 0.1 millimeters to minimize the staircase effect [1].

    - SCREEN: To be provided by the authors - Screen capture movie of the slicing software as talent sets the extruder temperature to 230 degrees Celsius and the temperature of the printer bed to room temperature. Talent sets the layer height to 0.1 millimeters to minimize the staircase effect. Authors, please upload this screen capture to your project page.

  - Set the printing speed to 3600 millimeters per minute. Set the amount of interior fill percentage to 80%. Include the supporter formation during printing, which is needed because the structure is hollow in the interior [1].

    - SCREEN: To be provided by the authors - Screen capture movie of the 3D printing software as talent sets the printing speed to 3600 millimeters per minute. Talent sets the amount of interior fill percentage to 80%. Talent includes the supporter formation during printing. Authors, please upload this screen capture to your project page.

- Smoothing out the Surface

  - Smooth out the surface because rough surfaces can damage the vessels by abrasion [1].

    - MED or CU: Talent lifts the printed device out of the 3D printer.

  - Remove the supporters using cutters. The supporters are attached at the interior of the stent. When removing the stents, exercise extreme caution to avoid tearing the stents [1].

    - CU: Stent as talent removes the supporters using cutters.

  - Rub the surface against sandpaper to remove the layer lines, striations, or blemishes on the printed surface [1]. Repeated polishing may be needed where the supporters are removed by the cutters [2].

    - CU: Stent as talent rubs the surface against sandpaper.

    - ECU: Area where the supporters were removed as talent polishes there.

  - Paint the surface using a spray in a well-ventilated location, and wear a personal mask. Clean, sand, and dry the surface. Protect from overspraying by applying thin layers of repeated paints. Use black paints to enhance the contrast between the silicone vessel mockup and the stent. –Authors, is this process filmable? Yes, it can be filmed.

- Deforming the Bifurcated Stent and Insertion of the Bifurcated Stent into the Vessels

  - Place the bifurcated stents in warm water such that the temperature is above the glass transition temperature [1]. When the stent becomes softened, push one half of the branch against the other half. Nest one half within the other half [2].

    - MED: Talent places the bifurcated stents in warm water.

    - CU: Softened stent as talent pushes once half of the branch against the other half and nests one half within the other half.

  - Now, fold the two branches into a single cylinder so that it can travel through the main vessel [1]. Perform the same nesting process to the other branch [2]. Subsequently, the two halves of the cylinders are closed into one [3].

    - CU: Stent as talent folds the two branches into a single cylinder so that it can travel through the main vessel.

    - CU: Stent as talent performs the nesting procedure to the other branch.

    - CU: Stent as talent finishes it up, closing the two halves of the cylinders into one.

  - Immerse the silicone vessel mockup inside the tank filled with warm water [1-TXT]. Orient the mockup such that the main vessel is above and the branches are below [2].

    - CU: Tank with warm water as talent immerses the silicone vessel mockup there. TEXT: 55-60 °C

    - CU or ECU: Mock-up as talent orients it so that the main vessel is above and the branches are below.

  - Now, insert the folded bifurcated stent into the opening of the silicone vessel mockup [1]. Orient the folded bifurcated stent such that its branches are towards the opening [2].

    - CU: Folded bifurcated stent as talent inserts it into the opening of the silicone vessel mockup.

    - CU: Folded bifurcated stent as talent orients it such that its branches are towards the opening.

  - The folded bifurcated stent will start to expand, and the lower branches will divide such that each branch will slide towards its mating pathway from the bifurcation core of the Y-shaped vessels [1].

    - CU or ECU: Folded bifurcated stent as it starts to expand

OPTIONAL – Critical Step Statement:

- An OPTIONAL brief statement may be submitted for further elaboration of the best way to perform the required technique for the single most critical step of this procedure.

- If there is no single critical step, then there is no need to fill out this statement.

- This will be an interview style shot interjected after the relevant step within the Protocol section of the video.

- This statement is limited to 30 words or less.

- Please indicate the full name of the Author who will give this statement and the step of the protocol to which the statement pertains using the step numbers from the Protocol section (above).

# Fill in the details below based on the instructions above for the “Critical Step Statement”

# Author name, Step            :           (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

Section – Results

# The Results section is restricted to 200 words of narrative. Please read through the results as presented to make sure that it accurately represents your findings. If you would like to highlight other data, please revise this section accordingly, keeping in mind the word count restriction. Please note that we cannot include narrative without an accompanying visual.

- Results: The Original and Deformed State of the Bifurcated Stent - Authors, feel free to change this title if you add more to the results section (please just keep it to <2 lines of text)

  - The stent uses a kirigami structure to allow the bifurcated stent to fold into a compact cylindrical tube, which is very suitable for sliding through the narrow pathways of blood vessels [1]. - Authors, please indicate how the narrators should pronounce “kirigami” in the narration of the video. It is pronounced [|kɪrɪ|ɡɑ:mɪ].

    - Figure 11.TIF - Authors: If you have flattened this figure, please submit it in its original file format as unflattened, layered images so that different components can be emphasized in the video. Video editors, please emphasize the right picture as this point is narrated. I have unflattened form of the stents picture. I uploaded.

  - The shape memory polymer allows the folded structure to return to its original shape when the temperature reaches the glass transition temperature. The original shape closely matches the branched vessels [1].

    - Figure 11.TIF - Video editors, please emphasize the left picture as this point is narrated.

# Authors, if you would like to add more to the results, perhaps you make SCREEN capture movies looking at the final printed structure, possibly showing it rotating or zooming into the discussed aspects. If you decide to do this, please send me the movies and the corresponding descriptions (keep in mind there is a 200 word limit for the entire Results section)

Section - Conclusion

- Conclusion Interview Statements: (Said by you on camera) - All interview statements may be edited for length and clarity.

- Below are questions for statements that can be used to further emphasize the significance of your protocol. At least one statement is required.

- Each statement is limited to 30 words.

- Answer the questions in full sentences, as you will be expected to memorize and deliver the sentences as spoken interview statements during filming.

- Indicate the full name of the author who will give each Conclusion Interview statement.

# What is most important thing to remember when attempting this procedure? Please indicate the steps (e.g., 2.4., 2.5.) in the Protocol section this advice correlates to.

  - Dahong Kim: The temperature of the printer bed was set at room temperature. We experienced unwanted deformation of the structure when the printer bed temperature was set higher. (Step: 3.4.1) (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

# Following this procedure, what other methods can be performed? What questions would these additional methods answer?

  - Dahong Kim: It can be replaced by a polymer that has a shape memory effect. However, an optimal temperature of memorizing is different depending on the used material. (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

# After its development, did this technique pave the way for researchers to explore new questions within a specific scientific field? If so, how?

  - Author Name: (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

# Are any of the reagents or instruments hazardous? If so, please use this interview statement to remind viewers of what precautions they should take.

  - Dahong Kim: When experimenters paint the surface using spray, don’t forget to wear gloves and mask. (Write your answer here in the form of a spoken statement. Don’t forget to replace “Author Name” with the name of the person who will be speaking the statement on camera)

Thank you for following the instructions and addressing our questions. We will incorporate your answers/suggestions and send you the finalized script before your shoot. You will also receive detailed shoot preparation instructions in the email accompanying the finalized script.