TITLE:

Isokinetic robotic device to improve the test-retest and inter-rater reliability for stretch reflex measurements in stroke patients with spasticity

AUTHORS & AFFILIATIONS:

Minki Sin1, Won-Seok Kim2, Kyujin Cho3, Nam-Jong Paik2

1Department of Medical Assistant Robot, Daegu Research Center, Korea Institute of Machinery & Materials, Daegu, Republic of Korea

2Department of Rehabilitation Medicine, Seoul National University College of Medicine and Seoul National University Bundang Hospital, Seongnam, Republic of Korea

3School of Mechanical and Aerospace Engineering, Seoul National University/IAMD, Seoul, Republic of Korea

Minki Sin and Won-Seok Kim equally contributed to this study and preparation of the manuscript.

EMAIL ADDRESSES FOR AUTHORS:

Minki Sin (mksin@kimm.re.kr)

Won-Seok Kim (wondol77@gmail.com)

Kyujin Cho (kjcho@snu.ac.kr)

Nam-Jong Paik (njpaik@snu.ac.kr)

Corresponding Author:

Nam-Jong Paik (njpaik@snu.ac.kr)

Tel: 82-31-787-7731

KEYWORDS:

Stroke, muscle spasticity, stretch reflex, isokinetic, reliability, quantification, electromyography, torque

SHORT ABSTRACT

Using our developed robotic isokinetic device with electromyography (EMG) measurements, we proved that isokinetic motion itself can improve the inter-rater reliability for the angle of catch measurements in stroke patients with mild elbow flexor spasticity.

LONG ABSTRACT

Measuring spasticity is important in treatment planning and determining the efficacy after treatment. However, the current tool used in clinical settings has been shown to be limiting inter-rater reliability. One of the factors for poor interrater reliability is the variability of passive motion while measuring the angle of catch (AoC) measurements. Therefore, an isokinetic device has been proposed to standardize the manual joint motion; however, the benefits of isokinetic motion for AoC measurements has not been tested in a standardized manner. Here, we investigated whether the isokinetic motion itself can improve the inter-rater reliability for AoC measurement. For this purpose, we developed a robotic isokinetic device combined with surface electromyography (EMG); and two conditions – manual and isokinetic motions – were compared with the standardized method to measure the angle and subjective feeling of catch. We showed that in 17 stroke patients with mild elbow flexor spasticity, isokinetic motion improved the intraclass correlation coefficient (ICC) for the inter-rater reliability of AoC measurements to 0.890 (95% confidence interval (CI): 0.685-0.961) by the EMG criteria and 0.931 (95% CI: 0.791-0.978) by the torque criteria, from 0.788 (95% CI: 0.493-0.920) by manual motion. In conclusion, the isokinetic motion itself can improve the inter-rater reliability of AoC measurements in stroke patients with mild spasticity. Given that our system may provide greater standardized angle measurement and catch of feeling, it may be a good option for the evaluation of spasticity in the clinical setting.

INTRODUCTION:

Spasticity after stroke is common. It has been shown to induce complications, including pain and contractures, resulting in reduced quality of life1-3. Measurement of spasticity is important to properly plan the course of treatment and to determine the efficacy of treatment. Commonly used tools in the clinical setting are the modified Ashworth scale (MAS)4, which is a nominal measurement system for resistance to passive movement, and the modified Tardieu scale (MTS), which measures the angle of catch (AoC), representing the velocity-dependent characteristic of spasticity5. However, these measurement tools have been shown to have limited inter-rater reliability6,7, requiring the same rater to perform these tests to maintain satisfactory reliability8.

Three factors have been shown induce the variability in AoC during MTS measurement: (1) errors from angle measurements by a goniometry; (2) variability of manually moved joint motion profile between raters; and (3) variability in sensing the catch between raters9. We developed a new isokinetic robotic device with torque sensors; this device was applied to stroke patients with a mild elbow flexor spasticity with surface electromyography (EMG) measurements10.

Our hypothesis was that the standardization of elbow joint motion could improve the inter-rater reliability for AoC measurements elicited by the elbow flexor stretch reflex. To prove this, a robotic device was developed to control assessment motion and measure the rater's motion quantitatively, and used it to compared the reliability for AoC, measured by surface EMG and torque, between the isokinetic passive and manual fast elbow extension. In addition, relation between the assessment motion consistency and the AoC reliability were investigated in order to identify the effect of the assessment motion difference on reliability. The overall configuration and sequence of the experiment are shown in Fig.1.

PROTOCOL:

- Experimental setup

  - Patient Recruitment

Note: All of the procedures were reviewed and approved by the Seoul National University Bundang Hospital Institutional Review Board. These subjects were inpatients or outpatients with stroke from four rehabilitation hospitals in the region.

    - Perform the screening process using the following inclusion criteria.

Note: The criteria are designed to screen patients who are able to participate in the experiment and to regulate factors affecting the result.

(1) Upper extremity hemiparesis due to stroke; (2) over the age of 20 years; (3) mild elbow joint spasticity of MAS 1-2; (4) no previous disease affecting the function of the hemiparetic arm, except for stroke; (5) free from hemodynamic instability; (6) no severe elbow contracture; (7) possibility for the shoulder to be abducted 90 degrees and the forearm to be in the neutral position without any joint pain; (8) normal cognitive, language, visuospatial or attention ability to follow experimental procedures.

    - Recruit subjects who were provided a detailed explanation of the entire study and the expected clinical issues; their consent was obtained prior to inclusion.

    - Demographics and baseline characteristics of the recruited subjects are shown in Table 1.

  - Experimental System

Note: A customized robotic device was developed and used to produce standardized motion and measure the quantitative data, simultaneously. The robotic system consists of a robotic part, a control system and measurement units. The overall configuration is shown in Fig. 2.

    - Robotic part

      - Basically, the robotic part is a 1-DoF planar robot consisting of a motor and a forearm manipulandum. It has three other components for adjusting the robot height and installing the device to various desks. The overall composition is shown in Fig. 2(a).

      - The forearm manipulandum consists of an elbow joint connected to the motor, a linear slider with a fixation block to adjust the overall length and two cuff units for fixating the forearm and hand (as shown in Fig. 3). The elbow joint has a rotating plate and a thrust bearing to prevent chafing during the experiment, and the cuff units were curved similar to that of a human forearm and were made using a 3D printer (3Dison+, Lokit, Korea). The handle unit is designed to accommodate both left and right handed individuals, making it available to any subjects.

      - The motor (EC-60, Maxon, Switzerland) uses a low gear ratio of 51:1 to have backdrivable characteristics and possible to produce nominal speed of 315 °/s and a continuous torque of 42.33Nm.

      - A lab jack unit is attached to the bottom of the motor to adjust the height of the motor unit. The height of the robot can be adjusted to the sitting height of various subjects.

      - A fixation arm for mounting the device to the desk is placed on the front part of the robot. The fixation arm is movable up and down through a linear shaft and has clamps for securing to the desk.

      - Casters with stopper are placed on the bottom of the robot, making the robot movable and grounded during the experiment.

    - Control system

      - The control component part consists of PC, real-time processor (sbRIO-9632, National Instruments, USA), and motor driver (DC Whistle, Elmo, Israel). The detailed control architecture block diagram is shown in Fig. 4.

      - The PC provides a graphic user interface (GUI) to control the experiment mode (maximum ROM measurement, Isokinetic MTS, and manual MTS measurement modes) and store robot motion data. The GUI was made using LabVIEW (National Instruments, USA). It contains a control panel and a monitoring panel (Fig. 5). Details on the GUI configuration are included in the Appendix.

      - The robot control algorithm was implemented by a real-time processor. The control algorithm consists of three control loops. The first loop is a data input/output loop that runs at 1 MHz from the FPGA module of the sbRIO. The second is a robot motion control loop that runs at 1 kHz from the real-time VI level. The last is a data communication loop that runs at 250 Hz. This loop transmits robot data (time, angle, torque and trigger signal for matching with EMG data).

      - The real-time processor has two communication modules: NI-9237 and NI-9853 (National Instruments, USA). The NI-9237 is an analog input device for receiving torque sensor data, and the NI-9853 is a CAN communication module for communicating with the motor driver.

    - Measurement units

      - A torque sensor (TRS-1K, Transducer Techniques) was mounted between the manipulandum and the motor to measure the reaction force. The torque data is transferred to the real-time processor via NI-9237. The NI-9237 has its own passband, stopband, and alias-free bandwidth filter. The filtered data enters the FPGA module and is processed again at 100 Hz with a low-pass filter to remove noises.

      - The joint angle is measured by an encoder (HEDL 9140, Maxon, Switzerland ) attached to the motor. The angle data is transferred to the real-time processor via the motor driver.

      - An eight-channel surface EMG device (WEMG-8, Laxtha, Korea) was used to measure muscle activity. The EMG data was collected at a sampling rate of 1024Hz, and initially processed with a bandpass filter (20-450 Hz) and a notch filter (60 Hz). The measured EMG data is transferred to the PC directly.

- Experiment Setup

Note: Two raters were participated in this experiment. The first rater was a physiatrist with more than 6 years of experience in the rehabilitation and the second rater was an occupational therapist with more than 3 years of experience in stroke rehabilitation.

  - Initial posture setting

    - Place the patient in a chair with his/her back in a straight posture.

    - Secure both sides of the shoulder and abdomen with seat belts to keep the shoulder position stable throughout the experiment.

    - Place the subject’s hemiparetic arm lightly on the robot manipulandum without fastening the strap.

    - Unfasten the fixation block of the linear slider so that the cuff can be moved freely on the slider and allow the subject’s hemiparetic arm to be placed on the robot manipulandum without fastening the straps.

    - Adjust the height of the robot using the lab jack until the patient’s shoulder is abducted 90 degrees. Confirm the abduction angle using a goniometer.

    - Make the subject hold the handle and fasten the hand to the handle with straps and align the rotation axis of the robot and the anatomical axis of the elbow joint.

    - Flex and extend the elbow joint so that the position of the cuff can be readjusted naturally in an optimal position without generating resistance during the elbow movement. Then, fasten the fixation block to fix the position of the cuff and fasten the straps of the forearm cuff.

    - Attach the surface EMG electrodes on the biceps brachii muscle in the hemiparetic arm.

  - Passive ROM measurement

Note: Passive ROM is used as a boundary ROM in the following experiments to prevent problems caused by movement outside the patient’s operation range.

    - Input the patient's hemiparetic side information into the program GUI (right or left).

    - Set the elbow 90 degrees flexed using a goniometer. Press ’90 deg set’ button on the GUI panel. This process matches the angle recognized by the robot with the actual human joint angle.

    - Press the ‘Finish set’ button on the GUI to switch the robot to the actuating state.

    - Click the buttons on the ‘Motor run’ panel on the left side of the GUI in order from top to bottom.

    - Turn on the ‘Angle set’ button and set the speed to 1 ̊/s. Then click the ‘run’ button. The robot will extend the elbow slowly at 1 ̊/s from 90 degrees flexed posture until the reaction torque reaches a certain threshold level or extends by 170 degrees.

Note: In this experiment, the torque threshold was set to 0.6 Nm. This value is determined experimentally via a pilot study.

    - The maximum extended angle is automatically stored as the maximum ROM.

    - Change the speed to -1 ̊/s and click the ‘run’ button again. The robot flexes the elbow slowly until the reaction torque reaches the threshold level.

    - The maximum flexed angle is automatically stored as the minimum ROM.

- MTS measurement

Note: The time required for each step is shown in Fig. 1. Total time taken by one subject to perform all the experiment was about 50 minutes (including the experiment setup step); but most of the time was spent resting to maintain consistency of fatigue.

  - Inertia effect compensation

Note: Theoretically, there was no inertia effect during the isokinetic motion. However, there can be an inertia effect at the beginning of the motion. The inertial force should be compensated to measure only the reaction force generated by a stretch reflex. Since the magnitude of the inertial force is different for each subject, a preliminary test for the inertial force compensation should be performed before the actual MTS measurement. Example result is shown in Fig. 6.

    - Click the ‘Back’ button on the control panel. The robot will flex the elbow to minimum angle posture (maximally flexed posture)

    - Set the speed as 150°/s and turn on the ‘Inertia test’ button and click the ‘Run’ button. The robot will apply a short perturbation of 5-degrees to the patient at a rate of 150°/s. The peak torque and period value of each trial are automatically stacked and displayed on the GUI panel.

    - Repeat steps 3.1.2 and 3.1.3 two more times. Determine a proper peak torque value and a period value from the measured data and enter the value on the program GUI.

    - The compensation torque profile (τcomp) is automatically generated based on equation (1). Where a represents the determined amplitude, and λ represents the period.

τcomp=a2∙(1-cos2π∙tλ)     (t≤λ)-0.15∙a2∙(1-cos2π∙tλ)(λ<t≤2λ)          0            (t>2λ)  	(1)

Note: The shape of the inertial torque is modeled as a raised cosine shape to reduce the calculation load. The compensation torque, which was designed for two periods due to the inertia effect, almost disappears after the second period. The amplitude of the second period was designed to be 15% of the first period.

  - Familiarization step

    - Prior to the actual experiment, three training operations are carried out to familiarize the patient with sudden movements.

    - Click the ‘Back’ button on the panel. The robot will flex the elbow to the minimum angle posture.

    - Click the ‘Run’ button after informing the subject. The robot will extend the patient’s elbow at a rate of 150°/s, until the angle reaches to maximum angle or the reaction torque reaches the threshold level.

    - Repeat steps 3.2.2 and 3.2.3 two more times and take a 5-minute rest before starting the actual test.

  - Isokinetic MTS measurement

Note: The isokinetic MTS measurement was designed to implement an ideal MTS measurement condition. The robot produces accurate constant velocity motion at a predetermined speed (150°/s) until reaching the maximum ROM or until reaching a certain threshold of the reaction torque. The maximum ROM value is determined in step 2.2 and the torque threshold value was determined as 0.6 via previous pilot studies to be sufficient for detecting stretch reflexes.

    - Click the ‘Back’ button to flex the elbow to the minimum angle posture.

    - Click the ‘Run’ button without informing the subject. The robot will extend the patient’s elbow at a rate of 150°/s until the angle reaches the maximum angle or the reaction torque reaches a certain threshold level. Time, angle, reaction torque, and trigger signal data are stored during the test.

    - Take a 2-min break between sets and repeat steps from 3.3.1 to 3.3.3 two more times.

    - Take a 5-minute rest after performing 3 sets.

  - Manual MTS measurement

Note: The manual MTS measurement was designed to simulate the MTS measurement normally performed at the actual medical sites. To compare the result of the isokinetic MTS, the robotic device was used only as a quantitative measuring tool that removed the measurement error, and actual measurement operation was performed by a human rater. For this purpose, the robot only drives to compensate the friction of the robot itself. Detail of the friction removal is in the appendix.

    - Click the ‘Back’ button to flex the elbow to the minimum angle posture.

    - Click the ‘Free run’ button; robot operation will change to manual operation mode.

    - Hold the handle of the manipulandum and stretch the subject’s arm. During the operation, the rater tries to generate a constant speed of 150 °/s.

    - Turn off the ‘Free run’ mode and take a 2-min break.

    - Repeat steps from 3.4.1 to 3.3.4 two more times

  - Repeat MTS measurement

    - Take a 10-min rest after finishing the whole experiment with the first rater.

    - Change the rater (to second rater) and repeat the experiment steps 3.3 and 3.4.

- Quantifying the AoC

Note: AoC was determined based on two data: EMG and torque. AoC was determined by a manual analysis due to the noisy characteristics of the EMG data and given the variability of individual characteristics. The AoC selection was carried out by a third rater, who was blind to the order of raters.

  - Isokinetic MTS experiment data analysis

    - AoC evaluation using EMG data

Note: Generally, AoC was determined as the angle at which the maximum peak value of the EMG occurs. However, a stretch reflex duration was different for each patient; thus, using the EMG maximum peak point as an AoC was expected to have a low reliability. The time difference may not be large; however, the AoC error can be significant due to the fast assessment speed of the MTS method. Therefore, the angle at the start of EMG upsurge point was selected as the AoC.

      - Process the raw EMG data using the root-mean-square (RMS) to smooth the data and amplify it 50 times.

      - Synchronize the EMG data and the robot angle data using the trigger signals of each data set.

Note: In our system, the EMG data was measured by an independent device, unlike other data; therefore, the reference time may be different. The EMG device has a trigger interrupt marking function, which gets the trigger signal from a real-time processor at the onset of MTS assessment.

      - Determine the AoC manually as the starting point of the RMS EMG upsurge. The example is shown in Fig. 7.

Note: The RMS EMG of less than 0.1 was ignored because it appeared frequently even without the stretch reflex. Thus, a clear upsurge point at the start of the peak was selected as the AoC.

    - AoC evaluation using the torque data

Note: Muscles have a passive mechanical characteristic that acts like a spring-damper system. Even if the muscle does not exert any force, the reaction force can increase as the muscles stretch. Because the intensity of the passive mechanical property and stretch reflex varies from patient to patient, it is difficult to identify the catch using only the absolute value of the reaction force. Instead, in this study, the catch was determined by changing the passive property due to spasticity rather than the absolute value of the reaction force. The change of the passive property is manually determined by the change in the slope of the regression line of the reaction torque.

      - Draw one regression line from the point where the trigger signal goes up and draw another regression line from the point where the trigger signal goes down.

      - Compare the slopes of the two regression lines. If the gradients of two regression lines show a significant difference, AoC can be determined at the intersection of two regression lines. The example is shown in Fig. 8.

  - Manual MTS experiment data analysis

Note: In the case of manual MTS it is difficult to separate the force exerted by the subject and the force applied by the rater using only one torque sensor. Therefore, in the case of manual MTS, only an AoC analysis using EMG data was performed without performing AoC analysis using the torque data.

    - AoC Evaluation using EMG data

- Note: Basically, the method to determine the AoC evaluation using EMG is the same with the isokinetic MTS case.

      - Process the raw EMG data using the RMS method to smooth the data, and amplify it 50 times.

      - Synchronize the EMG data and the robot angle data using the trigger signals of each data set.

      - Determine the AoC manually as the starting point of the RMS EMG upsurge. An example is shown in Fig. 9.

- Data Analysis

  - Normalized Assessment Motion Index (NAMI)

Note: The AoC of MTS can be affected by various motion factors, such as assessment speed, acceleration, and etc. Therefore, the assessment motion should be as isokinetic as possible. The NAMI was proposed to evaluate the ideality of the assessment motion. The proposed index is a non-dimensional index that can be used to evaluate the consistency of the assessment motion assigned to the subjects in each trial.

    - Calculate the ROM, maximum velocity and assessment time from each assessment trial.

Note: The angle was measured by the encoder; thus, the calculated velocity was noisy. Therefore, the maximum velocity was determined as the maximum velocity of the trend line, not the peak point.

    - Calculate the NAMI value for each trial during the whole experiment using the following equation (3):

NAMI=θmax- θminωmax∙∆t→degdeg/s∙s  (3)

      - The θmax and θmin represent a maximum and minimum angles, respectively, measured during the experiment, ωmax is the maximum assessment speed and ∆t is the total time spent in one assessment. Fig. 10 shows the example of each variable.

Note: The proposed index gave a score close to 1 if the assessment motion was close to complete isokinetic, and a score close to 0 if the velocity of the motion was inconsistent.

  - Statistical analysis

Note: All statistical analyses were performed using the PASW statistical package (SPSS version 18.0). The intraclass correlation coefficient (ICC) method was used to identify the test-retest reliability and the inter-rater reliability. Only results from the second and third tests were used to calculate the ICC.

    - To verify the test-retest reliability, calculate the ICC from the measured AoC data and the NAMI result.

    - To verify the inter-rater reliability, calculate the ICC from the average of AoC and the NAMI data.

    - Calculate the p-value of the AoC results using paired sample t-tests to evaluate the differences between each rater or each assessment trials.

Note: P-values of less than 0.05 were considered as statistically significant.

    - Calculate the Pearson correlation coefficient between AoC based on EMG criteria and torque criteria to verify the correlation between the two methods.

REPRESENTATIVE RESULTS:

The reliability is divided into four grades according to the ICC value: extremely excellent (>0.90), excellent (0.75<ICC≤0.90), fair to good (0.40<ICC≤0.75) and poor (<0.40). The standard error of measurements (SEM) was calculated to determine the error component of the variance. The smallest detectable difference (SDD) was calculated from the SEM of test-retest data.

Normalized Assessment Motion Index (NAMI)

The NAMI score during an isokinetic motion was always one, which means that the isokinetic device always generated a uniform constant input velocity. However, the test-retest reliability of the NAMI during a manual motion was poor for both rater 1 (ICC [95% CI] = -0.035 [-0.495-0.441]) and rater 2 (ICC [95%CI] = 0.438 [-0.038-0.752]). Moreover, the inter-rater reliability of the NAMI during manual motion was also poor (ICC [95% CI] = 0.148 [-0.344-0.576]).

Conversely, the results of the two human raters showed almost equal averaged NAMI value each other (0.68 and 0.67 for each rater). The consistency error of the two human raters was larger than that of the isokinetic device, showing a large difference between the two raters. These results indicate that an assessment motion by a human rater is lacking in the isokinetic characteristics and the motion is inconsistent depending on the subject.

Test-retest reliability

Table 2 shows the test-retest reliability for the AoC results in three conditions (isokinetic-EMG, isokinetic-torque and manual-EMG). The test-retest reliability for manual MTS was excellent (ICC = 0.804 and 0.840). However, the isokinetic MTS measurement improved the test-retest reliability to the extremely excellent grade on both the EMG and torque criteria (Table 2)

Inter-rater reliability

Table 3 show the inter-rater reliability for the AoC measurement performance in three conditions. The ICC of the inter-rater reliability of the manual MTS was 0.788, which was near the lower limit of the excellent grade. The isokinetic MTS improved the inter-rater reliability to the ICC of 0.890 based on the EMG data and to the ICC of 0.931 based on the torque data.

Correlations and consistency of timing of AoC between the EMG and torque criteria

The two AoC results calculated from the EMG data and the torque data during the isokinetic MTS show a significantly high correlation in both rater 1 (Pearson correlation coefficient=0.937, p<0.001) and 2 (Pearson correlation coefficient=0.957, p<0.001). Moreover, the timing of AoC between the two results was highly consistent with an ICC of 1 (p<0.001).

FIGURE AND TABLE LEGENDS:

Figure 1. Experiment flow chart.

Figure 2. Isokinetic MTS test Robot. (a) Configuration of the isokinetic robot device, (b) Inside configuration of the device. The control system includes real-time processor and motor driver. Fig 2. (b) is from Sin et al.10

Figure 3. Composition of the manipulandum. Two cuffs for wrist and forearm are connected to the linear slider through fixation block making the position of the cuff adjustable. Handle and hand strap are switchable left and right.

Figure 4. Control system Configuration. The right three blocks show the hierarchy of the control system and the arrows show the data flow between each unit.

Figure 5. Graphic User Interface (GUI) used for the experiment. The left side is the controller panel, which contains the various buttons or numeric controls required for robot control and the right side is a monitoring panel that shows the angle, interaction torque and trigger signal in real time.

Figure 6. Example of inertia effect compensation. The green line indicates the raw torque; the blue dotted line indicates the inertial force model; and the red line indicates the inertial torque compensation result. This figure is from Sin et al.10

Figure 7. Example of AoC evaluation using EMG data (isokinetic MTS case). RMS EMG value less than 0.1 is regarded as a normal. Select the starting point of the clear EMG upsurge point and the angle value at that time is determined as AoC. This figure is from Sin et al.10

Figure 8. Example of AoC evaluation using torque data (isokinetic MTS case). Draw two lines connecting the torque of the assessment starting point and the end point with an arbitrary torque data, respectively. Find the point where the two lines become the regression line of the torque data before and after the selected point. If there is a significant difference between the gradient of two regression line, it is judged that a stretch reflex occurs at this point. This figure is from Sin et al.10

Figure 9. Example of AoC evaluation using EMG data (Manual MTS case). Same as the isokinetic case (Fig. 7), the AoC is determined as the angle when clear upsurge of the EMG happened. This figure is from Sin et al.10

Figure 10. Variables for the Normalized Assessment Motion Index (NAMI). Intuitively, the NAMI value is the ratio of the area under the velocity graph to the area of the gray box. The more isokinetic movements show a value close to 1. This figure is from Sin et al.10

Table 1. Subjects demographics and baseline characteristics

Table 2. Test-retest reliability results for the angle of catch measured with isokinetic robotic devices and robotic devices with manual motion. This table is from Sin et al.10 SEM: Standard error of measurement, SDD: Smallest detectable difference, ICC: Intraclass correlation coefficient, EMG: Electromyography

Table 3. Inter-rater reliability results for the angle of catch measured with isokinetic robotic devices and robotic devices with manual motion. This table is from Sin et al.10 SEM: Standard error of measurement, ICC: Intraclass correlation coefficient, EMG: Electromyography

DISCUSSION:

In this study, we tried to standardize the MTS measurement using a robotic isokinetic device.

We investigated how the consistency of assessment motion affects the results of MTS measurement.

The NAMI value was proposed to represent the degree of variability in the assessment motion. As expected, unlike the isokinetic motion method with no variability, the manual method showed variability between tests and between raters, resulting in poor reliability, which is consistent with the results from previous studies7,8. The result on the reliability for AoC measurement show that the isokinetic motion itself can increase the interrater reliability, compared with the manual motion. Although, there have been concerns regarding the less stretch reflex provocation by the isokinetic motion11,12, subjects in our study with a mild elbow flexor spasticity (MAS 1, 1+, 2) showed consistent stretch reflexes measured by the surface EMG during isokinetic motion. This demonstrates that an isokinetic device can be used to measure AoC reliably, even in patients with mild elbow spasticity. AoC was also calculated by the torque criteria in this study. Interestingly, AoC measured by using both the EMG and torque criteria showed a high correlation, while the torque criteria alone showed a higher inter-rater reliability, which is consistent with the results provided by Lynn et al.13. Therefore, the spasticity evaluation using the torque criteria is expected to be a better mehod with respect to reliability and convenience.

However, this is new approach for quantifying the MTS measurement still has some issues and limitations. First, the posture during AoC measurements in this study was different from the conventional MTS measurement14. The conventional MTS was made in the absence of shoulder abduction; in contrast, in this study, measurement was performed with the shoulder abducted 90 degrees. However, the purpose of this study was to verify the effects of consistency of the assessment motion on the AoC reliability. The posture used in this experiment makes it easy to measure AoC using the torque data by eliminating the influence of forearm weight, which is difficult to measure separately. Therefore, this experiment provides a perspective on how the assessment motion affects the reliability of AoC measurement. Second, the AoC measurement using both the torque and EMG criteria was performed subjectively. However, this was conducted by a third rater who was blind to the subject information and the order of raters, to mimize potential bias. Third, the increase of the reaction torque due to passive mechanical properties was unexpected when designing the experiment initially. We expected the reaction torque to be mainly caused by stretch reflex; however, in patients with mild spasticity, many cases found that the reaction torque caused by passive stiffness was dominant. Therefore, the AoC was through post-experiment data analysis rather than real-time identification. Finally, there was relaxation of the elbow flexor during repetitive passive stretching. The experiment was designed to incorporate sufficient resting time to prevent fatigue throughout the experiment; no subject complained of fatigue. However, it is hard to prevent relaxation of the muscle due to repetitive passive stretching. To reduce this impact, the experiment was designed to randomize the order of raters, and the result showed no significant relaxation phenomenon between the two raters.

The goal of our study was to improve evaluation method that rely on the subjective sense of the rater to more objective and quantitative standards. Current results show possibility to increase assessment reliability by using a robotic device. Although, the total method performed in this study is only half automated method because the AoC evaluation was done by a person. It is expected that the further study will enable real-time spasticity evaluation with high reliability and objectivity.

ACKNOWLEDGMENTS:

This study was supported by the Seoul National University Bundang Hospital Research Fund (14- 2014 - 035) and Korea and National Research Foundation of Korea (NRF) Grant funded by the Korean Government (A100249).

DISCLOSURES:

All authors declare no conflict of interest.

REFERENCES:

- Sommerfeld, D. K., Gripenstedt, U. & Welmer, A.-K. Spasticity after stroke: An overview of prevalence, test instruments, and treatments. American Journal of Physical Medicine & Rehabilitation. 91 (9), 814–820 (2012).

- Sommerfeld, D. K., Eek, E. U.-B., Svensson, A.-K., Holmqvist, L. W. & von Arbin, M. H. Spasticity after Stroke: Its Occurrence and Association with Motor Impairments and Activity Limitations. Stroke. 35 (1), 134–139 (2004).

- Lundström, E., Terént, A. & Borg, J. Prevalence of disabling spasticity 1 year after first-ever stroke. European Journal of Neurology. 15 (6), 533–539 (2008).

- Ashford, S. & Turner-Stokes, L. Systematic Review of Upper-limb Function Measurement Methods in Botulinum Toxin Intervention for Focal Spasticity. Physiotherapy Research International. 18 (3), 178–189 (2013).

- Patrick, E. & Ada, L. The Tardieu Scale differentiates contracture from spasticity whereas the Ashworth Scale is confounded by it. Clinical Rehabilitation. 20 (2), 173–182 (2006).

- Li, F., Wu, Y. & Li, X. Test-retest reliability and inter-rater reliability of the Modified Tardieu Scale and the Modified Ashworth Scale in hemiplegic patients with stroke. European journal of physical and rehabilitation medicine. 50 (1), 9–15 (2014).

- Mehrholz, J., Wagner, K., et al. Reliability of the Modified Tardieu Scale and the Modified Ashworth Scale in adult patients with severe brain injury: a comparison study. Clinical Rehabilitation. 19 (7), 751–759 (2005).

- Ansari, N. N., Naghdi, S., Hasson, S., Azarsa, M. H. & Azarnia, S. The Modified Tardieu Scale for the measurement of elbow flexor spasticity in adult patients with hemiplegia. Brain Injury. 22 (13-14), 1007–1012 (2008).

- van den Noort, J. C., Scholtes, V. A. & Harlaar, J. Evaluation of clinical spasticity assessment in Cerebral palsy using inertial sensors. Gait & Posture. 30 (2), 138–143 (2009).

- Sin, M., Kim, W.-S., Cho, K., Cho, S. & Paik, N.-J. Improving the test-retest and inter-rater reliability for stretch reflex measurements using an isokinetic device in stroke patients with mild to moderate elbow spasticity. Journal of Electromyography and Kinesiology. 39 (1), 120–127 (2018).

- Grippo, A., Carrai, R., et al. Biomechanical and electromyographic assessment of spastic hypertonus in motor complete traumatic spinal cord-injured individuals. Spinal Cord. 49 (1), 142–148 (2011).

- Rabita, G., Dupont, L., Thevenon, A., Lensel-Corbeil, G., Pérot, C. & Vanvelcenaher, J. Differences in kinematic parameters and plantarflexor reflex responses between manual (Ashworth) and isokinetic mobilisations in spasticity assessment. Clinical Neurophysiology.  116 (1), 93–100 (2005).

- Lynn, B.-O., Erwin, A., et al. Comprehensive quantification of the spastic catch in children with cerebral palsy. Research in Developmental Disabilities. 34 (1), 386–396 (2013).

- Boyd, R. N. & Graham, H. K. Objective measurement of clinical findings in the use of botulinum toxin type A for the management of children with cerebral palsy. European Journal of Neurology. 6 (1), 23–35 (1999).